[image: image14.png]


 

                          Temperature Probe Assembly

	COMPANY OPERATING PROCEDURES


0210410 Paediatric Rectal / Oesophageal

VM3/COP/34.10.2
	Date:
7-Oct-03
	Revision date: 31-Jan-07
	Issue:   3




[image: image1.wmf]
  [image: image2.jpg]


                                                                          [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. NEED UPDATED!!!!!!  SHOULD BE 3.5MM JACK not 6.5mm – FOR USE WITH DOLPHIN 2150 ONLY (D Lamb)
	Parts

	Qty
	Description
	Part No.

	
	
	

	1
	
[image: image4.wmf] Thermister sub-assembly
	0231120

	1
	
[image: image5.wmf] Connector
	

	1
	
[image: image6.wmf]                Casing
	0231001

	1
	
[image: image7.wmf]              Strain Relief
	0030618

	1
	
[image: image8.wmf]                       Cable Grip
	0030500

	1
	(6 x 30 mm Clear heat shrink
	0032331

	1
	(6 x 40mm Black heat shrink
	0032321

	2
	(9 x 15mm Clear heat shrink
	0032332

	
	
	


ASSEMBLY OPERATIONS


[image: image9.wmf]Red

Black

Cable grip

Strain relief

Casing

Connector


1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above from the stores.

3.
Feed (6 x 30 mm clear heat shrink, casing, strain relief, and cable grip, (9 x 40mm Black heat shrink and both (9 x 15mm pieces of heat shrink over cable end.

4.
Trim Red and black wires to 8mm and strip 2mm off the ends.

5.
Solder red wire to the leg of the connector (see diagram) and heat the (9 x 15mm clear heat shrink over it to cover naked wire and bolt.

[image: image10.jpg]



6.
Solder the black wire to the surface and heat (6 x 40mm black heat shrink to the end.

7. 
Clamp the grip and cable grip around the cable. 

8.
Screw the casing onto the connector.

Labelling

1.
Labels to be attached:


· 1 x CE Label

· 1 x Serial no. Label

· 1x White ‘do not throw away’ dumbbell label with the following label attached to the reverse side:

[image: image14.png]

[image: image11.wmf]37204

CE

0086

Ø6 x 30 clear heat shrink

Serial No.

temperature probe

A150260708

Neonatal rectal/esophageal

White dumbell 

labell


Testing

1.
Using the test equipment check to see if the probe works. Reading for mono should show between 2 and 3 k(.

[image: image12.jpg]


      [image: image13.jpg]



Packaging and Storage

1.
Visually check all labels are attached properly.

2.
Using a twist tie (bunny clip) wrap the cable and place in a plastic bag. Attach a label on the bag showing all necessary details. Also seal the bag using a packed and tested sticker.


0210410


Paediatric rectal/oesophageal temperature probe. 


AI50260807





0210410


Paediatric rectal/oesophageal temperature probe. 


AI50260807














Page 1 of 3

_1090932673.dwg

_1094293263.dwg

_1094293562.dwg

_1090932719.dwg

_1090837400.dwg

_1090837439.dwg

_1090837323.dwg

_1090837367.dwg

