[image: image1.png]Vandagraph
>ensor Technologies






Specification Datasheet - 8010020
Divesoft s.r.o.
type 22D Oxygen Sensor
[image: image2.jpg]



v1.3
21st January 2021
	Specifications


	VST part number
	8010020

	OEM model number
	Type 22D Oxygen Sensor, P/N 3092

	Output
	9 – 13 mV

	Output connector
	3 pin Molex

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Range
	0 – 100% oxygen

	Pressure range
	600 to 2500 mbar

	Repeatability
	<1% volume oxygen at constant temperature and pressure

	Linearity error
	0.1 to 1.0 bar < 3% relative
0.1 to 2.5 bar < 5% relative
(at ppO2 > 1.0 bar after 1 hour equilibration time)

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25 ºC after 36 seconds

	Cross interference
	< 0.5% volume oxygen response to 5% CO2 balance N2

	Response time
	< 5 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 50 °C, 5% relative error

	Nominal sensor life
	≥ 500,000 % oxygen hours

	Long term output drift
	< 1% volume oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	≥ 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes

	Conformal coating on PCB
	Yes

	Pressure equalization - PCB
	Yes 

	Pressure equalization - output cap
	None

	Weight
	25 g

	Flow diverter
	No

	Packaging
	Gas barrier bag - transparent

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25 °C dry ambient air.



Page 2 of 4
	Drawings/images


[image: image3.jpg]@145
|
T
~ i
|
|
| w0
| 3 -
T
|
\
X
! M16 x 1 (5/8-24 UNEF)
@28
1 _—Locking Laich
(-)f%/ TN
h==a
$29

%\N MOLEX CONNECTOR



                                                                                
Image 

[image: image4.jpg]



[image: image5.png]DIVESOFT Oxygen Sensor type 22D

#3092

ivesoft. ™M SIN - D123456
wowdivesottcom /N (98] C€ 002 NI




[image: image6.png]N\
DIVESOFT

Halkova 2495 #3002
413 01 Roudnice nad Labem www divesoft.com
Czech Republ info@divesoft.com
+420 416 857 057

Oxygen Sensor type 22D

Fore use with Divesoft Liberty CCR

(ﬂ 2020-02 [l Ce
T A % .

‘This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do ot puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor
may require 30 minutes to reach signal stability after installation.

J





Page 3 of 4
	Labelling


Oxygen sensor label: 90.0mm x 15.0mm
White vinyl label, text & graphics – black
[image: image7.jpg]“nsor type 227
SIN D10105%
A

|

I




Gas barrier bag label: 63.5mm x 46.6mm
White label, text & graphics – black
[image: image8.jpg]\ygen Sensor type 220

(vj] 1055
Carees NI





Serial number prefix: D 
Page 4 of 4
Vandagraph Sensor Technologies Ltd.
15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom
Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com
www.vandagraphst.com

