Specification
Range	0- 100% Oxygen
Accuracy :	+/- 1% of full scale
Response Time	90% step change < 10sec
Resolution :	+/- 0.1 %
Battery Type :	MN1500 LR6 1.5v
Battery Life :	12 months (typical)
Sensor Type :	R-17VAN (Galvanic)
Sensor Life in air	Expected 36-48 months (10 months in 100%O2)
Sensor output	7.5mV - 13.5 mV in air
Dimensions	60mm x 120mm x 35mm
Weight	198 gms including battery & sensor
Storage Temp	0o-50o C (recommended 10-30C)
Operating Temp	0o-40 oC
Case :	Splash proof
Waterproofing	IP65 NEMA 4
ISO Standards	ISO EN 9001:2000
EMC CE	ISO EN 60601-1-2
RoHs	Complies with RoHS
WEEE	Complies with WEEE

      Specification subject to change
Accessory Parts list
		R-17VAN Micro Fuel Cell
		DINKIT Restricter Kit
		A-268 Tee Adaptor
		B-50057 Flow- Divertor
		DM22M10 Male adapter
		BS111 Viton O Ring
		DIN22F Restrictor,
		Quick-Ox gas sampling kit
		V mount clamp
		V mount adapter
		Lanyard
		Waterproof box
9530002		Battery/Sensor cover seal
9520001		TEK-OX sensor extension
	Trouble shooting
	Symptom		Possible Cause
	No Display		Not switched on Battery exhausted / Batteries in wrong way
	Batt Lo symbol		Replace low battery
	Zero Reading		Sensor exhausted Sensor disconnected Rotate sensor whilst in the instrument. There may be a deposit on the Jack plug or the sensor Jack socket
	Not calibrating		Sensor nearly exhausted Check sensor in 100% then Check sensor again in Air
	Reading drifts		Sensor nearly exhausted; in flowing gas Temperature changing wind blowing on divertor
	Inaccurate reading		Move away from RF source, Boat VHF radio Do not use flowing gas use On/Off /On/Off system:Condensation on sensor face Remove condensate by shaking

Vandagraph Ltd 15, Station Road, Crosshills,
Keighley West Yorkshire, BD20 7DT. United Kingdom.
Tel 01535 635900
        Fax 01535 635582
        www.vandagraph.co.uk
        sales@vandagraph.co.uk
        technical@vandagraph.co.uk

VN202 mkII
Oxygen Analyser
User Manual

These instructions should be read before
 using the

Mixed gas diving should only be undertaken by divers trained by a recognised training organisation.
Copyright 2007 Vandagraph Ltd.
. All worldwide rights reserved.

Contents / Please check you have the following items: / 1 VN202 / 2. Quick-Ox or DINKIT

The Vn202mkII
has been designed as one of the next generation of diving Oxygen analysers following onfrom the sucessful VN202 Suitable for the worldwide hostile diving environment it is easy to use with either hand,:and has very large digits.
It uses the proven R-17VAN oxygen sensor (as used in the VN202) with integral temperature compensation
It is powered by two internationally available AA alkaline batteries with an estimated life of 12-18 months. The instrument is available in two versions:
Yellow On/Off button automatic switch off.
Green On/Off button non automatic switch off The instrument is sealed for water resistance (IP65) and designed to be drop resistant. The battery is housed in a separate compartment from the sealed electronics and can be changed without the use of a tool. Ensure batteries are fitted correctly. See internal labels
Full comprehensive user manual is available as a download from

 www.Vandagraph.co,uk/manual.
    Just register your purchase.

 How to use the VN202 mkII
1.Leave Flow divertor on;
Wave through the Air

2. Adjust Calibration

3. Add The Quick-Ox

4,Turn on the Gas to a
    gentle Flow

5.Mate the Quick-Ox
    to the cylinder outlet

 6.When the reading stops rising turn off cylinder
If in doubt repeat!

Effects of Altitude & Humidity
At very high altitudes some sensors (with high outputs) may not calibrate down to the level required. In this instance total pressure, ambient and altitude must be taken into consideration
Feet	Metres	Pressure mb	Cal
-1000	-305	1.03	21.5
0	0	1	20.9
1000	305	0.97	20.2
2000	610	0.94	19.4
3000	915	0.92	18.7
4000	1220	0,89	18.1
5000	1220	0.89	17.4
6000	1830	0.84	16.8
8000	2440	0.79	15.5
10,000	3050	0.74	14.4
12,000	3660	0.69	13.3

Although not substantial, humidity can effect the maximum accuracy that can be obtained. This accounts for the difference observed between ambient air calibration and calibration with dry gas from a cylinder
	RH40%	RH60%	RH80%	RH100%
0oC/32oF	20.9%	20.8%	20.8%	20.8%
10oC/50oF	20.8%	20.7%	20.7%	20.6%
20oC/70oF	20.7%	20.6%	20.5%	20.4%
30oC/90oF	20.5%	20.3%	20.1%	19.9%
40oC/100oF	20.4%	20.1%	19.8%	19.5%

Using the VN202 mkII with a DINKIT
      Follow instructions 1-6 as with Quick-Ox. See information on DINKIT