[image: image1.png]


[image: image2.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid





                       

Vn202 Mkii  Oxygen Analyser CE file.


Risk Assessment  EN ISO 14971:2000 Annex D : Possible hazards with medical devices.

	Ref.
	Hazard.


	Related part / Component posing risk.
	Sev

of

Haz.
	Like

of

Haz.
	Risk.
	Solution.
	Document referenced.
	Sev

of

Haz.
	Like

0f

Haz.
	Risk.

	D.2
	Energy hazards and contributory factors

	D.2.1
	Electricity
	Mains powered
	3
	3
	9
	Battery powered

Designed to ISO 60601-1 & MDD requirements
	(G) Technical manual

(E) Risk analysis

ISO 60601-1
	1
	1
	1

	D.2.2
	Heat
	If battery polarity reversed

No direct heat
	3
	1
	3
	Protection diode added in circuit. 

Warning in user manual
	(G1) Technical manual

(Y.Z) Design file

(F5) User manual
	1
	1
	1

	D.2.3
	Mechanical force
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.4
	Ionising radiation
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.5
	Non ionising radiation
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.6
	Moving parts
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.7
	Unintended motion
	 
	2
	3
	6
	 Mounting screw
	 User Manual
	1
	2
	2

	D.2.8
	Suspended masses
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.9
	Patient support failure
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.10
	Pressure (vessel rupture)
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.11
	Acoustic pressure
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.12
	Vibration
	 
	2
	3
	6
	 
	(F5) User manual / inserts 

(E) Risk analysis 
	1
	2
	2

	D.2.13
	Magnetic fields 

(eg. MRI)
	Instrument  
	4
	4
	16
	Warning in user manual
	(F5) User manual

(E) Risk analysis 
	2
	1
	2

	D.3
	Biological hazards and contributory factors

	D.3.1
	Bio-contamination
	 
	1
	2
	2
	 
	(F) User manual / inserts

(E) Risk analysis  
	1
	1
	1

	D.3.2
	Bio-incompatibility
	N/A
	1
	1
	1
	
	(T2) Manufacturers data
	1
	1
	1

	D.3.3
	Incorrect formulation (chemical composition)
	N/A
	1
	1
	1
	
	(T2) Manufacturers data
	1
	1
	1

	D.3.4
	Toxicity
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.5
	Allergenicity
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.6
	Mutagenicity
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.7
	Oncogenicity
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.8
	Carcinogenicity
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.9
	Re and/or cross infection
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.10
	Pyrogenicity
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.11
	Inability to maintain hygienic standards
	Instrument / cable / sensor
	4
	2
	8
	 
	Manufacturers data

(F5) User manual

(E) Risk analysis  
	1
	2
	2

	D.3.12
	Degradation
	Instrument / cable through normal wear & tear
	1
	1
	1
	Care instructions given in the user manual
	(F) User manual
	1
	1
	1

	D.4
	Environmental hazards and contributory factors

	D.4.1
	Electromagnetic fields
	Instrument / cable / sensor
	2
	2
	4
	Circuit design & layout
	(D) EMC report

(Y16) Design reviews

(E) Risk analysis  
	1
	1
	1

	D.4.2
	Susceptibility to electromagnetic interference
	Instrument / cable / sensor
	2
	2
	4
	Circuit design & layout
	(D) EMC report

(Y16) Design reviews

(E) Risk  
	1
	1
	1

	D.4.3
	Emissions of electromagnetic interference 
	Instrument / cable / sensor
	2
	2
	4
	Circuit design & layout
	(D) EMC report

(Y16) Design reviews

(E) Risk analysis  
	1
	1
	1

	D.4.4
	Inadequate supply of power 
	Battery
	2
	3
	6
	Circuit design

Battery low alarm

Description in user manual
	(Y16) Design reviews

(F5) User manual
	1
	2
	2

	D.4.5
	Inadequate supply of coolant
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.6
	Storage / operation outside prescribed environmental conditions
	Instrument / cable / sensor
	1
	1
	1
	User manual / labelling
	(F5) User manual

(F7) Labels

(E) Risk analysis  
	1
	1
	1

	D.4.7
	Incompatibility with other devices with which the product is intended to be used
	 
	3
	2
	6
	 
	(F7) Labels

(F5) User manual
	1
	1
	1

	D.4.8
	Accidental mechanical damage
	Instrument / cable 
	1
	1
	1
	Impact resistant case used

Robust cable used
	(T2) Manufacturers data
	1
	1
	1

	D.4.9
	Contamination due to waste products and/or device disposal
	Instrument / cable / sensor
	4
	2
	8
	 
	(T2) Manufacturers data

(F5) User manual / inserts

(E) Risk  
	1
	1
	1

	D.5
	Hazards resulting from incorrect output of energy and substances

	D.5.1
	Electricity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.2
	Radiation
	Radiated EM interference
	1
	1
	1
	Circuit design & layout
	(D) EMC report

(Y.Z) Design file
	1
	1
	1

	D.5.3
	Volume
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.4
	Pressure
	N/A
	1
	1
	1
	Atmospheric Pressure
	User Manual
	1
	1
	1

	D.5.5
	Supply of medical gases
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.6
	Supply of anaesthetic agents
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6
	Hazards related to the use of the medical device and contributory factors

	D.6.1
	Inadequate labelling
	Instrument
	3
	1
	1
	Product easy to use

User manual / insert

Labels
	(F5) User manual & inserts

(F7) Labels
	1
	1
	1

	D.6.2
	Inadequate operating instructions
	User manual & Insert
	3
	1
	1
	Product easy to use

User manual & insert
	(F5) User manual & inserts
	1
	1
	1

	D.6.3
	Inadequate specification of accessories
	User manual & insert
	4
	1
	1
	Product easy to use

User manual & insert
	(F5) User manual & inserts
	1
	1
	1

	D.6.4
	Inadequate specification of pre-use checks
	User manual & insert
	3
	1
	1
	Product easy to use

User manual & inserts
	(F5) User manual & inserts
	1
	1
	1

	D.6.5
	Over-complicated operating instructions
	User manual & insert
	2
	1
	1
	Product easy to use

User manual & inserts
	(F5) User manual & inserts
	1
	1
	1

	D.6.6
	Inadequate specification of service and maintenance
	User manual & inserts

Technical manual
	1
	1
	1
	No service required except external cleaning of case & cable and battery replacement
	(F5) User manual & inserts

(G1) Technical manual
	1
	1
	1

	D.6.7
	Use by unskilled / untrained personnel
	Instrument / cable / sensor
	1
	1
	1
	Product easy to use

User manual & Insert
	(F5) User manual & inserts

(E) Risk analysis  
	1
	1
	1

	D.6.8
	Reasonable foreseeable misuse
	Instrument / cable / sensor
	1
	1
	1
	Product easy to use

User manual & Insert
	(F5) User manual & inserts

(E) Risk analysis  
	1
	1
	1

	D.6.9
	Insufficient warning of side effects
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.10
	Inadequate warnings of hazards likely with re-use of single use devices
	 
	4
	2
	8
	 
	(F5) User manual & inserts

(E) Risk analysis  
	1
	1
	1

	D.6.11
	Incorrect measurement and other metrological aspects
	
	1
	1
	1
	 
	(F5) User manual & inserts
	1
	1
	1

	D.6.12
	Misrepresentation of results
	Instrument
	1
	1
	1
	Product easy to use 

Description / photos in user manual
	(F5) User manual & inserts
	1
	1
	1

	D.6.13
	Incompatibility with consumables / accessories / other devices
	 
	4
	3
	12
	1. 
	(F5) User manual & inserts

(E) Risk analysis (mask sensor & universal sensor)
	1
	1
	1

	D.6.14
	Sharp edges or points
	
	1
	1
	1
	
	
	1
	1
	1

	D.7
	Inappropriate, inadequate or overcomplicated user interface ( man/machine communication)

	D.7.1
	Mistakes & judgement errors
	Instrument / cable / sensor
	1
	1
	1
	Product easy to use

User manual & insert 
	(F)5 User manual & inserts

(E) Risk analysis 
	1
	1
	1

	D.7.2
	Lapses and cognitive recall errors
	Instrument / cable / sensor
	1
	1
	1
	Product easy to use

User manual & insert
	(F5) User manual & inserts

(E) Risk analysis  
	1
	1
	1

	D.7.3
	Slips & blunders (mental or physical)
	Instrument / cable / sensor
	1
	1
	1
	Product easy to use

User manual & Insert
	(F5) User manual & inserts

(E) Risk analysis  
	1
	1
	1

	D.7.4
	Violation or abbreviation of instructions,  procedures etc
	Instrument / cable / sensor
	1
	1
	1
	Product easy to use

User manual & Insert
	(F5) User manual & inserts

(E) Risk analysis  
	1
	1
	1

	D.7.5
	Complex or confusing control system
	Instrument
	1
	1
	1
	Description / photos  in user manual
	(F5) User manual
	1
	1
	1

	D.7.6
	Ambiguous or unclear device state
	Instrument
	1
	1
	1
	Instrument specific LCD

Description / photos in user manual
	(F5) User manual
	1
	1
	1

	D.7.7
	Ambiguous or unclear presentation of settings, measurement, or other information
	Instrument
	3
	3
	9
	Instrument specific LCD

Description / photos in user manual
	(F5) User manual
	1
	1
	1

	D.7.8
	Misrepresentation of results
	Instrument
	1
	1
	1
	Breath rate only

Description / photos in user manual
	(F5) User manual
	1
	1
	1

	D.7.9
	Insufficient visibility, audibility or tactility
	 
	1
	3
	3
	  Large LCD.
	(Y16) Design review
	1
	2
	2

	D.7.10
	Poor mapping of controls to action or of displayed information to actual state
	Instrument
	1
	1
	1
	 
	(F5) User manual
	1
	1
	1

	D.7.11
	Controversial modes or mappings as compared to existing equipment
	Instrument 
	1
	1
	1
	Easy to use
	
	1
	1
	1

	D.8
	Hazards arising from function failure, maintenance and ageing and contributory factors

	D.8.1
	Erroneous data transfer
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8.2
	Lack of, or inadequate specification for maintenance including post maintenance functional tests
	Instrument / cable
	1
	1
	1
	No service required except external cleaning of case & cable, and battery replacement. Return to manufacturer for repair.
	(F5) User manual & inserts

(G1) Maintenance manual
	1
	1
	1

	D.8.4
	Inadequate maintenance
	Instrument / cable
	1
	1
	1
	No service required except external cleaning of case & cable, and battery replacement. Return to manufacturer for repair.
	(F5) User manual & inserts

(G1) Maintenance manual
	1
	1
	1

	D.8.5
	Lack of adequate determination of end of device life
	Instrument / cable / sensor
	1
	1
	1
	Oxygen Sensor depletion
	(E) Risk analysis  
	1
	1
	1

	D.8.6
	Loss of electrical integrity
	N/A
	1
	1
	1
	
	(F5) User manual & inserts
	1
	1
	1

	D.8.7
	Loss of mechanical integrity
	Instrument / cable / sensor
	1
	1
	1
	
	(F5) User manual & inserts

(E) Risk  


	1
	1
	1

	D.8.8
	Inadequate packaging (contamination and / or deterioration of the device)
	Instrument / cable / sensors
	2
	2
	4
	Specifically designed packaging 
	(M) Packaging

(M) Packaging manufacturers info

(E) Risk  
	1
	1
	1

	D.8.9
	Re-use and/or improper re-use
	Instrument / cable / sensors
	1
	1
	1
	 .
	(F5) User manual & inserts

(E) Risk analysis  
	1
	1
	1

	D.8.10
	Deterioration in function (gradual occlusion of fluid / gas path or change in resistance to flow, electrical conductivity) as a result of repeated use
	Instrument / cable
	1
	1
	1
	Failure of circuitry / cable likely to be obvious to user.  Return instrument to manufacturer for repair.
	(F5) User manual & inserts
	1
	1
	1



VN202 mkii Annexe D.doc        24/07/2012      Page 5 of 6
Explanation   1 = Insignificant   2 =  Significant    3  = Tolerable    4  =  Catastrophic     Risk  calculated as product   1 - 16


