	Vandagraph Sensor Technologies

	COMPANY OPERATING PROCEDURE

	VM3/COP/41.00

Sensor Pressure test with Narked @ 90

	Date 25/01/12
	Revision Date     25/01/12
	Issue 1

	



.


1. There are two ways to test the sensors under pressure.

Using 100% Oxygen to 3-4 Bar and      Air to 10 Bar

Each method stresses the sensor beyond its specification.

100% Oxygen has been chosen as it is easier to work a lower pressures and 
prevent leaks.

Past tests have shown that within the accuracy levels achieved with 10 Bar

air
there is no measurable difference in response

2. Change the pot head to accommodate the sensors to be tested

3. Check Linearity 
100% Dry Oxygen  Note reading
Dry Air Note reading
Ratio  100/20.9 = 4.78   If the ration is incorrect there is no point in testing further. A quick check should give a value within +/ 2%  eg,4.69 – 4.88
Care should be taken to ensure the sensor is stable before readings are taken, It takes time for oxygen to completely enter the sensor and to be evacuated from the sensor.    

4. This test can be recorded on the Pico system below or on the Narked at 90 DVM's

5. Switch display on and attach the cable from display to top of Pot head head. This checks the pot head continuity. The display DPM should read.

6. Disconnect the pot head cable and screw the pot head onto the pot. It should be leak proof.

7. Displays can be switched off at this point

8. Connect the Narked at 90 connector into the High Resolution Display Data Logger.

9. Connect the other end into the Pot Head connector.
NB this is now adapted so that the Pico and the DVM,s can be viewed at the same time 

10. Plug in the Pressure Transducer

11. Add the Remobulator and switch on by pressing the blue switch for 2 seconds

12. Press the Yellow switch to read ABS BAR

13. Press the Blue switch and the digits should read around 1 Bar depending on atmospheric pressure

14. The Green switch changes the calibration for Oxygen/Air/or Nitrox.

15. Switch on the computer and ensure the High Resolution Display Data Logger is connected.

16. Using the remote display for clarity choose the Icon Pico Recorder which runs Pico Log

17. Set up Pico Log for all tests

18. File Change File name

19. File:   New settings    ( leave on default)

20. Choose sample speed This depends on the test.
1/sec gives good resolution but has a lot of data
1/min or multiple minutes for long-time tests

21. Choose stop to enable test to be complete 
I.e determine intermittent tests or continuous
e.g 1/sec  500 samples is 8 min

22. Converter details 
Converter type ADC20ADC24     OK

23. Set up channels

24. Highlight Channel and set Type and serial number of sensors e.g VST2200 0100234

25. Conversion time set to 100ms

26. Range has to be calculated
Sensor in Air @ 1 Bar is approx 10mV
Sensor in Air @ 10 Bar is approx 100mV
Sensor in Oxygen @1 Bar is approx 50mV
Sensor in Oxygen @4 Bar is approx 200mV
Sufficient to set at 315mV
Up to three sensors can be used together so 3 channels are allocated

27. Channel 4 set as Pressure transducer
Conversion time 100mS
Range 39mV

28. OK

29. PLW Record should display the sensors and the Pressure readings

30. The test box has been modified.








        Hold    Vent            Vent control

Gas In



        Gas out


 

31. Oxygen Inject Injects gas into the Pot

32. The Hold Vent switch Allows gas to leave the Pot. This is very sudden especially at high Bar pressures.

33. The hold vent control can be adjusted using a needle valve.
To let gas out of the POT the Vent control is closed
The Vent Hold switch is opened then the Vent control is slowly opened to let the pressure drop. The Hold vent switch can be closed at any time to stop the gas leaking whilst the Vent control can be adjusted to speed up or slow down the gas escape. For fine control the the two can be used to trap gas between them so it can be finely controlled. 

34. Use the notebook to record experiment. eg. below
Read 35- 48 before proceeding.

35. Connect Air source to the gas feed in on th Narked @90 Test box.

36. Vent Control Open

37. Vent/Hold switch to Vent.

38. Open Air cylinder

39. Inject small amounts of gas ,which is dry) keeping the pressure reading on the Remobulator to 1Bar.( atmospheric Pressure.)

40. The gas is cold so it must be injected slowly and the sensor allowed to stabilise.

41. Press the Red Button to record and record for 10 seconds.
Press the pause to stop the recording.

42. If satisfied with the recording close the Vent/Hold switch and close the Vent control.

43. Use the inject button to gently raise the pressure in the pot to say 2Bar.

44. If the reading is over 2 Bar 

45. Open then close the Vent/Hold switch. This allows a very small portion of gas passed the Vent/Hold switch which can be vented.

46. If the overshoot is great the Vent Control valve can be opened slowly with the Vent/Hold switch open. 

47. The Vent control can then be used to slowly reduce the pressure .Once near the target switch the Vent?Hold switch to Hold.
NB If the overshoot is large the excess Oxygen in the sensors need to equalise allowing the Oxygen to come out of the sensors. This can take time.

48. At High Bar settings everything happens much faster and there is a possibility of small leaks. The double vent systen does reduce this substantially but care should be taken to ensure no leaks. 

49. The readings on the computer should be observed as they are more accurate that the Remobbulator 

50. Standard 1- 3 sensor tests Air or Oxygen

51. To record the data ;choose the master file  required which is read only and immediately save under the new name I.e company initials and the SRN

52. The serial numbers of the sensors should be added by editing the settings.

53. When ready select record ( Red square)

54. Readings should start to be recorded.

55. For standard reports only record 10 readings.

56. Press pause to pause recording 
NB The first reading is always incorrect so the reporting system tries to take the middle sample only

57. Repeat for a range of Bar pressures  10 second samples.

58. After final reading press stop ( Black square)

59. Save file

60. Samples required


	Pressure to be set
	Test

	1Bar
	Air Tests / Oxygen Tests

	1.5 bar
	Oxygen Tests

	2.0 Bar
	Air Tests / Oxygen Tests

	2.5 bar
	Oxygen Tests

	3.bar
	Air Tests / Oxygen Tests

	3.5 bar
	Oxygen Tests

	4.0 Bar
	Air Tests / Oxygen Tests

	5.0 Bar
	Air Tests 

	6.0 Bar
	Air Tests 

	7.0 Bar
	Air Tests 

	8.0 Bar
	Air Tests 

	9.bar
	Air Tests 


61. If Oxygen is being used the maximum pressure need only be 4 Bar.

62. To test ratio use 1 Bar Air Then 1 Bar Oxygen. This will give the basic O2/Air Ratio 4.78.

63. The data plotted out  should only be for a single gas.

64. The data can be viewed in the Pico log either in chart or graphical format.

65. The data is stored in Documents so should be copied as soon as possible onto the T Drive system.

66. T:/John Lamb/Vandagraph/Vandagraph VST110710/VST 


Spreadsheet and reports


67. The files are stored locally in Documents and remotely in T:/John Lamb/Vandagraph/Vandagraph VST110710/VST Sensor returns Reports

68. To create a report or to analise the data Load Pico Log Recorder on any computer

69. No ADC is attached therefore OK

70. Load required file. See 67

71. View Spreadsheet in Pico log

72. No ADC is attached therefore OK

73. Highlight 1st line and track to last value by holding SHIFT

74. Copy to clipboard ( 2nd icon at top yellow page)

75. Open Spreadsheet Master air or Master Oxygen Oxygen both are read only so should be saved immediately as the initials of the company and the SRN

76. Save

77. Unicode OK

78. The first sheet name is Raw and contains all data

79. Sheet 2 is Calculate and has the calculation and graphs. 

80. Incorrect dat can be removed on this page

81. The fist reading on every sample is incorrect so we select readings 10 seconds apart starting at 8 seconds and process the results into a report

82. Sheet 3   Sensor 1 report

83. Sheet 4   Sensor 2 report

84. Sheet 5   Sensor 3 report

85. Comments can be added on Sheets  3-5

Pot








VM3COP41.00 Narked @ 90 Tests equipment updates.doc                  Page 5 of 5      25. Jan. 2012


