	VANDAGRAPH
  COMPANY OPERATING PROCEDURE

	Final Test Procedure Vandagraph VN202 mkII 

VM3/COP/40.17

	Date: 28 Jun 2007
	Revision Date: 10-Oct-12
	Issue 1

	Only to be attempted by trained personel



	Parts Required

	Quantity
	Description
	Part No.
	

	1
	Oxycal
	
	


1. Connect Oxycal to VN202mkII

2. Switch Oxycal to 100% and set VN202 mkII calibration to read 100% +/- 0.2

3. Switch Oxycal to 21% VN202 mkII skould read 20.9 +/- 0.2 

4. Switch Oxycal to 100% . Turn calibration until 88.8 or all the segments have been visible.

5. Autoswitch off versions should switch off after about 10 minutes

VM3COP40.12   VN202 mkII Oxycal set up.doc       10/10/2012      Page 1 of 1

