WGA-01

WGA-01 is used in the following instruments:

· Weld Wide Ar-Gone

· Sumner Ar-Gone

Parts Required:

· R-33S1

· 2 x OEM Label 

· 1 x Sensor label

· 1 x Barcode label

· Sensors Bag

Conversion Procedures:

1) Change the polarity of the R-33S1 sensor at the Molex connector.

2) Provide Derek with the barcodes of the R-33S1 sensors. Inform Derek or Cathy of the Model (WGA-01) and the company name to be on the sensor (Weld Wide or Sumner). Derek can then convert the sensor on Intrastats and print 2 labels for each sensor using the Citizen Printer.

3) Remove original Teledyne label and replace with new OEM label.

4) Add the second label to the sensor bag, add the sensor and heat seal, Add final barcode label.

