	COMPANY OPERATING PROCEDURES


Teledyne R17MED (PN 0110017), R17A (PN 0110117), 

R17AH (PN 0110137), R17VAN (PN 0110217),
O2 Sensor Test Procedure
VM3/COP/37.35

Test & Equipment Required
Functional test required using R17 tester and R17 test adapter, (digital multimeter and 3.5mm jack plug adapter).

Output specification as in document: (VM3COP37.00, column (i)).

Method.
1. Open the packaging and remove the O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte. 

3. Connect the O2 sensor to the R17 tester, select 200mV range (marked with arrows) and wait for a few seconds for the meter reading to stabilise.

4. Check for an output  See Table 1

5. If no, or very low output, or sensor is returned as faulty leave the O2 sensor exposed to air for 1 Hours Minimum.

6.  If the O2 sensor fails the output specification test, it should be left unsealed and boxed, placed in a docket and put on the goods in desk with a short note explaining why it failed, ready to be booked in with a SRN.

7. Disconnect the O2 sensor from the R17 tester and label with Viamed stickers if appropriate.

8. Reseal in serial numbered packet, integrity seal the edge of the packet with initialed and dated Viamed sticker and re-box. 

Table 1
	Part Number
	Output Range

	0110017
	10.5mV +/- 3 mV

	0110117
	10.5mV +/- 3 mV

	0110137
	Between 10.5mV / 13.5 mV

	0110217
	10.5mV +/- 3 mV


17/12/2012             1 of 1

