VM3COP38.10

Procedure to convert AX300 oxygen analysers and

MX300 oxygen monitors 

from analogue output to RS232 output.

Conversion of analyser/monitor part:

0111230 AX300: analogue output & 3.5 digit display to 0111232
or

0111233 AX300: analogue output & 2.5 digit display to 0111234
or

0111235 MX300: analogue output & 3.5 digit display to 0111237
or

0111238 MX300: analogue output & 2.5 digit display to 0111239
Parts list:

a) Analyser/monitor as above

b) 0131245 RS-232 transceiver

c) 0131246 100μF Tantalum capacitor, +/- 20%, 10V 

If utilizing base models 0111230 and 0111235 check to see if display modification also needs to be carried out (converting from 3.5 digit to 2.5 digit).

Procedure:

1) Remove link JP3 and reposition to JP7

2) Install RS2323 transceiver (part number 0131245) to position U6.

3) Remove existing capacitor C14 (22μF) and replace with 100μf capacitor (part number 0131246).

4) Test for stable RS232 output, ref. VM3COP38.09
5) Confirm  required display format and carry out change procedure if necessary.

6) Remark PCB reference on the white square, Assy C73884- (
7) Carry out standard QA test.

Cross reference table:

	P/N
	Description
	Output
	Digit display format
	PCB ref.

	
	
	Analogue
	RS232
	3 1/2
	2 1/2
	

	0111230
	AX300 Oxygen Analyzer
	(
	
	(
	
	E

	0111232
	AX300 Oxygen Analyzer
	
	(
	(
	
	G

	0111233
	AX300 Oxygen Analyzer
	(
	
	
	(
	F

	0111234
	AX300 Oxygen Analyzer
	
	(
	
	(
	H

	0111235
	MX300 Oxygen Monitor
	(
	
	(
	
	A

	0111237
	MX300 Oxygen Monitor
	
	(
	(
	
	C

	0111238
	MX300 Oxygen Monitor
	(
	
	
	(
	B

	0111239
	MX300 Oxygen Monitor
	
	(
	
	(
	D


v1.1

7th March 2013
