VM3COP38.11

Procedure to convert AX300 oxygen analysers and 

MX300 oxygen monitors 

from 3.5 digit display to 2.5 digit display format.

Conversion of analyser/monitor part:

0111230 AX300: analogue output & 3.5 digit display to 0111233
or

0111232 AX300: RS-232 output & 3.5 digit display to 0111234
or

0111235 MX300: analogue output & 3.5 digit display to 0111238
or

0111237 MX300: RS-232 output & 3.5 digit display to 0111239
Parts list:

a) Analyser/monitor as above

b) PCB link 

If utilizing base models 0111230 and 0111235 check to see if RS-232 output modification also needs to be carried out.

Procedure:

1) Install link at JP6.

2) Confirm  required output format and carry out change procedure if necessary.

3) Remark PCB reference on the white square, Assy C73884- (
4) Carry out standard QA test.

Cross reference table:

	P/N
	Description
	Output
	Digit display format
	PCB ref.

	
	
	Analogue
	RS232
	3 1/2
	2 1/2
	

	0111230
	AX300 Oxygen Analyzer
	(
	
	(
	
	E

	0111232
	AX300 Oxygen Analyzer
	
	(
	(
	
	G

	0111233
	AX300 Oxygen Analyzer
	(
	
	
	(
	F

	0111234
	AX300 Oxygen Analyzer
	
	(
	
	(
	H

	0111235
	MX300 Oxygen Monitor
	(
	
	(
	
	A

	0111237
	MX300 Oxygen Monitor
	
	(
	(
	
	C

	0111238
	MX300 Oxygen Monitor
	(
	
	
	(
	B

	0111239
	MX300 Oxygen Monitor
	
	(
	
	(
	D


v1.1

7th March 2013
