	     [image: image3.jpg]



	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley -West Yorkshire BD20 7DT - United Kingdom
Tel: +44 1535 634542                 Fax: +44 1535 635582

Email: info@viamed.co.uk  Website: www.viamed.co.uk



[image: image1.png]



Specification Datasheet - 0110040

[image: image2.jpg]


Viamed

R-30V Oxygen Sensor
	      Specifications


	Part number
	0110040

	Model
	R-30V

	Output
	10 mV – 12 mV 

	Output connector/connection
	Modular (phone) socket 6P4C

	Input connector/connection
	Threaded

	Input O ring
	Yes

	Range
	0 – 100% oxygen

	Accuracy
	< 1% volume oxygen when calibrated in 100% oxygen

	Influence of pressure
	Output proportional to change in oxygen partial pressure

	Linearity error
	< 3% relative

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 5 minutes

	Cross interference
	Meets requirements of EN ISO 21647

	Response time (t90)
	< 12 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% relative after fall from 1m

	Operating temperature range
	0 – 50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC thermistor

	Temperature compensation error
	+20 to +40 ºC: 3% relative

    0 to +50 ºC: 8% relative

	Nominal sensor life
	> 1,000,000 % oxygen hours

	Long term output drift
	< 1% volume oxygen per month, typically < -15% over sensor lifetime.

	Recommended load
	> 10 K(

	Shelf life
	24 months

	Weight
	50 g

	Flow diverter
	Not applicable

	Standards
	Meets with requirements of EN ISO 21647. Designed and manufactured according to EN ISO 9001:2008 & EN 13485:2007

	Packaging
	Blister/Blister card

	Warranty period
	15 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.



v1.0






[image: image3.jpg]