	VANDAGRAPH SENSOR TECHNOLOGIES 


  COMPANY OPERATING PROCEDURE

	Oxygen Sensor manufacture

	Modification of R-22DEM Sensors 

VM3/COP/40.14

	Date: 28 Jun 2007
	Revision Date: 25-Apr-13
	Issue 1

	Only to be carried out by trained personel





	Parts Required

	Quantity
	Description
	Part No.
	

	1
	R-22D Sensor (any version will work R-22D is preferred R-22A is lowest cost and has a grid)
	
	

	
	Serial number Labels
	
	

	2
	Serial number Label for Outer bag
	
	

	1
	Polythene Bag (outer)
	
	

	1
	PCB Blank or
	9012004
	

	
	PCB pre-wired with resistors
	9015005
	

	2
	4K7 ohm surface mounted resistors or 
	
	

	2
	5K ohm surface mounted resistors
	
	

	1
	Cable Red/White wires + low profile 2 way Molex or
	9070120
	

	1
	Crimp housing low profile 
	9081047
	

	1
	Crimp pins Low profile
	9071046
	

	1+1
	Red & White/Black wires
	
	


[image: image1.jpg]


[image: image2.jpg]



1. Open the packaging and remove the R-22D 


Sensor from the Gas barrier bag

2. Check the O2 sensor for damage and signs of 

leaking electrolyte.

3. Remove the O2 sensor specification leaflet and 

discard.  

4. [image: image3.jpg]


[image: image4.jpg]


Remove excess outer case unto level of PCB. 

VM3COP40.11

5. Using a Flow divertor hold the sensor in a vice

6. Remove the Molex carefully. The PCB is non   

conformal coated so the seal must be broken on 

the sensor carefully 
[image: image5.jpg]



1. W[image: image6.jpg]


here possible use Pt number  9015005   pre-wired PCB with 4K7 or 5K6 resistors and molex connected  The pads should be cleaned with a PCB rubber
If not PCB needs to be  loaded see page 2

7. Test for output 4.0 mV- 6.5mV.

8. Wire so that the leads lie naturally across the PCB

9. Cut the three pins as close to the solder joint as possible.

10. Remove the existing label and add R-22DEM label to sensor  

11. Ensure the O2 sensor is booked out in the stock books.

12. Put the sensor into a gas barrier bag 

13. Seal as near to the edge as possible

14. Add label

15. Add the instruction leaflet, and place the whole into the polythene bag, 





Notes for PCB construction

16. [image: image7.jpg]


[image: image8.jpg])




[image: image9.jpg]


The PCB should be cleaned with a PCB rubber to remove oxidation.

17. The Red & White or Red & Black wires should be cut to 5 cms and stripped with the tool. Insert the wire until the end touches the Red stop. Close the tool and the cable will be stripped. If not rotate the wire and repeat.

18. Tin the bare ends.

19. Inserts should be crimped onto the wires using the special tool 

20. The 4K7 or 5K6 resistors should be soldered onto the PCB using the Microscope set up. Only to be carried out by trained SMT engineer. Use two resistors of the same value on the same PCB

7112.doc       25/04/2013      Page 2 of 2

