Mixcheck Basic fault finding


Power    Verify power supply rails are +/- 5VDC +/- 0.2    (JP1-JP2)   &  (TP1 & TP2)




O2 Display





Helium reads 00.0









Use Calibration procedure Check He 0%   100%   and Trimix    approx 17% he   22% oxygen . If incorrect change Software.

Power on with Batteries





No





Yes





Change LCD2





No





Votage @ TP4





Check SW1 & SW2





No





Yes





No





Yes





Power On with


Mains adapter





5v at JP1





No





OK





U2





Jack Skt P1





Low Voltage Indicator Failure 





Check Voltages across


R1D





Display reads 00.0


Check Sensor output 7-13mV in air


& 


Connections








Sensor OK





OK





No





Change sensor R33S1





Check 


Output on R7





Yes





Voltage on Pin 7 LCD2





Yes





No





No





Check switch SW1





Check wiring  or Change


R7





Change switch SW1





Check Resistance on J3      	 Pins 1-3 =             


                       		Pins 2-3 =


		     		 Pins 1-2 =





Values approximate values with fault normally open circuit 





Check wiring 





Change He sensor





Yes





Check Voltages  on J3 to gnd  Pin 2 =    


			    	                     Pin 3 =                  Values approximate vary with R21





No





Check connector J3 J4 & wiring loom 





Yes Ok





Replace connector or wiring loom





No





Check TP3 = 





Yes





Check CPU





Check Xtal V3





Check PWM U4 pin 15 & U3 pin 14





Check Pin 7 He LCD 1





Yes





Check Pin 19 CPU 


or U4 pin10





Yes





Yes





Change LCD 1





Yes





See separate procedure for changing LCD








Mixcheck Fault finding.doc                                               Page 2       09/05/13


