Specification

Range     
0- 100% Oxygen 

Accuracy    : 
+/- 1% of full scale

Response Time  
90% step change  < 10sec

Resolution    :
+/- 0.1 %

Battery Type   :
MN1500 LR6 1.5v

Battery Life   :  
12 months (typical)

Sensor Type   : 
R-17VAN (Galvanic)

Sensor Life   in air     


Expected 36-48 months (10 months in 100%O2) 

Sensor output
7.5mV - 13.5 mV in air 

Dimensions  
60mm x 120mm x 35mm 

Weight 
198 gms including battery & sensor 

Storage Temp
0o-50o C (recommended 10-30C) 

Operating Temp
0o-40 oC  

Case     : 
Splash proof 

Waterproofing
IP65 NEMA 4

ISO Standards
ISO EN 9001:2000

EMC   CE
ISO EN 60601-1-2 

RoHs
Complies with  RoHS 

WEEE
Complies with WEEE

      Specification subject to change

Accessory Parts list


9710018
R-17VAN Micro Fuel Cell 


9730215
DINKIT Restricter Kit 


9711001
A-268   Tee Adaptor  


9711002
B-50057 Flow- Divertor


9711004
DM22M10 Male adapter 


9711003
BS111   Viton O Ring


9711006
DIN22F   Restrictor, 


9703210
Quick-Ox gas sampling kit


9713022
Lanyard


9713040
Waterproof box


9530002
Battery/Sensor cover seal


9520001
TEK-OX sensor  extension



Trouble shooting


Symptom
Possible Cause


No Display
Not switched on Battery exhausted

Batteries in wrong way


Batt Lo symbol
Replace low battery


Zero Reading
Sensor exhausted Sensor disconnected Rotate sensor whilst in the instrument. There may be a deposit on the Jack plug or the sensor Jack socket


Not calibrating
Sensor nearly exhausted Check sensor in 100% then  Check sensor  again in Air


Reading drifts
Sensor nearly exhausted; in flowing gas Temperature changingwind blowing on divertor


Inaccurate reading
Move away from RF source, Boat VHF radio Do not use flowing gas use On/Off /On/Off system:Condensation on sensor face; remove condensate by shaking

[image: image1.jpg]



Vandagraph Ltd   15, Station Road,  Crosshills,

Keighley  West Yorkshire, BD20 7DT.  United Kingdom.

Tel 01535 635900 

                              Fax 01535 635582

                              www.vandagraph.co.uk

                              sales@vandagraph.co.uk

                              technical@vandagraph.co.uk

TEK-OX

Oxygen Analyser 

User Manual

These  instructions should be read before

 using the 

TEK-OX

[image: image2.jpg]



Mixed gas diving should only be undertaken by divers trained by a recognised training organisation.

Copyright 2007 Vandagraph Ltd. 

.  All worldwide rights reserved.     

Contents

Please check you have the following items:

      1 TEX-OX      or

      2. Quick-Ox or DINKIT

The TEK-OX 

has been designed for the worldwide hostile diving environment. It is easy to use with either hand,:and has very large digits.

It uses the proven R-17VAN oxygen sensor  (as

 used in the VN202) with integral temperature compensation  

It is powered by two internationally available AA alkaline batteries with an estimated life of 12-18 months. The instrument is available in two versions:

Yellow On/Off button automatic switch off.

Green On/Off button non automatic switch off

The instrument is sealed for water resistance (IP65) and designed to be drop resistant.

The battery and sensor are housed in a separate c[image: image3.jpg]


ompartment from the sealed electronics and available by removing the 3 screws in the                                                base. Ensure batteries are fitted correctly. See internal labels

Full comprehensive user manual is available as a download from www.Vandagraph.co,uk/manual. 

    Just register your purchase.

 How to use the TEK-OX

1.Le[image: image4.jpg]


ave Flow divertor on; Wave through the Air

[image: image5.jpg]


2.  Adjust   Calibration

[image: image6.jpg]



3. Add The Quick-Ox

[image: image7.jpg]



4,Turn on the Gas to a 

    gentle Flow 

[image: image8.jpg]


5.Mate the Quick-Ox

    to the cylinder outlet 

6.Whe[image: image9.jpg]TEK-0X



n the reading

stops  rising turn 

off cylinder

Effects of Altitude & Humidity

At very high altitudes some sensors  (with low outputs) may not calibrate up to the level required. In this instance total pressure, ambient and altitude must be taken into consideration

Feet
Metres
Pressure mb
Cal

-1000
-305
1.03
21.5

0
0
1
20.9

1000
305
0.97
20.2

2000
610
0.94
19.4

3000
915
0.92
18.7

4000
1220
0,89
18.1

5000
1220
0.89
17.4

6000
1830
0.84
16.8

8000
2440
0.79
15.5

10,000
3050
0.74
14.4

12,000
3660
0.69
13.3

Although not substantial, humidity can effect the maximum accuracy that can be obtained. This accounts for the difference observed between ambient air calibration and calibration with dry gas from a cylinder    


RH40%
RH60%
RH80%
RH100%

0oC/32oF
20.9%
20.8%
20.8%
20.8%

10oC/50oF
20.8%
20.7%
20.7%
20.6%

20oC/70oF
20.7%
20.6%
20.5%
20.4%

30oC/90oF
20.5%
20.3%
20.1%
19.9%

40oC/100oF
20.4%
20.1%
19.8%
19.5%

Using the TEK-OX with a DINKIT

      Follow instructions 1-6 as with Quick-Ox[image: image10.jpg]



User insert Eng 080131.doc     01/02/08


