Serial Number	Comments External	Comments Internal
0093	TEK-OX label on front not straight / Case marked as though it had been gripped in a vice without soft padding / Calibration control back to front	No front panel retaining screws / Lug broken on front panel / No intermediate window.
0094	Calibration control back to front	No front panel retaining screws / Lug broken on front panel / No intermediate window.
0095	Calibration control back to front	No front panel retaining screws / No intermediate window.
0096	Calibration control back to front	No front panel retaining screws / Lug broken on front panel / No intermediate window.
0097	Calibration control back to front	No front panel retaining screws / No intermediate window.
	Calibration control back to front
	Calibration control back to front
	Calibration control back to front
	Calibration control back to front
	Calibration control back to front
	Calibration control back to front
	Calibration control back to front
	Calibration control back to front
	Calibration control back to front

Notes:

Serial no of PCB is useful to have and reference against the instrument serial number1. The Calibration controls are back to front. The instrument should calibrate clockwise.
2. Two of the front TEK OX labels are not straight.
3. You have serial numbers on the PCB. It would be very useful to have these noted against the instrument serial number for traceability purposes.
As I need to re-wire all the calibration controls I need to get inside all 8 of the instruments so I will let you have a comprehensive report on my findings in a few days
As to the scratches I have found some myself
I have also found so far
A missing intermediate window. This will fail the static test
Missing screws needed to retain the front panel
A broken lug on the front panel These two will effect the Waterproof rating
Have you read my procedure and QA list. All these points are addressed as I have done exactly the same when I was manufacturing the first.
We need a completed QA form with each on confirming QA has been carried out correctly.
Otherwise construction looks good .