Possible problems with Blu-Ox


Problem

After unit switches off automatically when switched on again E1 can appear

Hardware 

Changing LCD can remove problem

Occurs with both batches of PCB

Disabling Auto switch off does not clear problem

5 PCB's showed E1 ; removed from chassis then re-assembled 3 days later . Initially worked then showed E1

Software 

Re-Programme does not clear problem

S/W 3.98  Rechecks the reference while E1 is being displayed

Also has 

S/W 3.99   Quick auto shut off.  Does not clear problem. Problems hows after a time delay.

S/W 4.01 Clean screen  Does not clear problem

CP  A #152 change to 40 Works

S/W 402 

5 units tested 

E1 appears then is reset

Sometimes appears 

R-8 removed E1 flashes 

R-7 in parallel up with 86K ohm.

No effect until unit switched off then on E1 appears after 4.03.

Linearity not effected

Lo Batt not effected.

