# Certificate of Calibration

## Certificate number: 1689598886

## V1000 Foetal Heart Simulator
Part Number: 1410000
Serial number: PR03488A11

This certificate reports recorded values for the instrument specified above.

The output rates were measured in accordance with procedure: VM3COP55.01.

Equipment used:	Sonicade FM6 foetal monitor (CE051)

	TTi TF930 3GHz counter (CE185) - calibration traceable to UKAS calibration standards.

V1000 / BPM setting	Period values stated in milliseconds				Result / PASS/FAIL
	Target	Lower limit	Upper limit	Actual
60	1000.00	990.00	1010.00	995.08	PASS
90	666.70	660.03	673.37	666.96	PASS
120	500.00	495.00	505.00	501.02	PASS
150	400.00	396.00	404.00	400.92	PASS
180	333.30	329.97	336.63	333.28	PASS
210	285.70	282.84	288.25	286.40	PASS

Name:	Phil Crossley

Date:	17th July 2023

Signed:	______________________