VANDAGRAPH

O2 Analyser

VN 202 - Repairs

VM3/COP/40.01

	Date: 23 Sep. 03
	Revision Date:16-Dec-09
	Issue: 1



	Test and Equipment Required


1.0
Visual and functional test required – EQUIPMENT REQUIRED – Microcal or sensor to check the cable and calibration potentiometer, voltmeter to check for battery drain and voltage of battery. 

	Goods-In


1.0 Book the goods into the system; the procedure is the same as Viamed Except that the complete packaging is forwarded to the Vandagraph office.

2.0
File a copy of the work sheet with the original customer correspondence on Repairs file and send a copy of both back with the repaired item.

	Method


1.0 Check the case and the cable for damage.

2.0 If the battery is in – check the battery condition, if not up to standard replace with a new battery.

3.0 Check the condition of the battery connections.

4.0 Check the L.C.D. for obvious faults. 

5.0 Check the “Offset” at  “-3 to +3 (Maximum), if the reading is greater than “1” then reset to optimum.

6.0 Add Sensor Check calibration in Air – this should be reading “20.9” steady. Check in 100% Oxygen, check calibration returns to “21%” in Air.

7.0 Check the cable Analyser end for intermittent faults, Check the cable at the Sensor end for intermittent faults.

8.0 Rotate the sensor and check for dead point. 

9.0 Check the PCB for any water damage or dry solder joints.

10.0 Check for battery drain.

11.0 Check auto – switch off.

i)
Yellow button switch-off automatic after 2-3 minutes



ii)
Green button push for “ON” push for “OFF”

12.0 Check the labels and replace where necessary.

13.0 Clean the case with a soft damp cloth, clean or replace the carrying case if necessary. 

14.0 Once complete remove battery ready for despatch/storage

15.0 Fill in the work sheet completely.

VANDAGRAPH repair modified.doc           Page 1 of 1

