
Oxygen Analysers display Partial pressure so if the Analyser is used in Altitude diving care must be exercised to understand and compensate for the changes in partial pressure that occur at Altitude.

At high altitude normal calibration may not be possible due the output restrictions of the sensor. Oxygen analysers should be designed  not to calibrate if the output is outside the specified limits of the sensor. The correct value for the altitude must be set on calibration. The pressure varies with altitude and the table below is a guide.

To set a calibration at 3,000  feet ( approx 1,000 mtrs) the normal pressure of 1Bar will only be 0.92 Bar.

The normal calibration 20.9 must be multiplied by 0.92 = 19.2%

At 3,000 ft  calibrate to 19.2%

In practice Latitude and atmospheric conditions also effect the pressure.

Latitude effects can be ignored  within the accuracy of normal monitoring

For maximum accuracy the exact atmospheric pressure at the altitude the measurement is to be taken should be used.

As a rule of thumb a 5 mb pressure change will change the calibration by 0.1%

	Altitude correction chart at 1Bar

	Altitude
	Pressure
	Calibration

	Feet
	Metres
	Millbar mb
	%

	-1000
	-305
	1.03
	21.7

	0
	0
	1
	20.9

	1000
	305
	0.97
	20.2

	2000
	610
	0.94
	19.4

	3000
	915
	0.92
	18.7

	4000
	1220
	0.89
	18.1

	5000
	1525
	0.86
	17.4

	6000
	1830
	0.84
	16.8

	7000
	2135
	0.81
	16.1

	8000
	2440
	0.79
	15.5

	9000
	2745
	0.76
	14.9

	10000
	3050
	0.74
	14.4

	11000
	3355
	0.72
	13.8

	12000
	3660
	0.69
	13.3

	13000
	3960
	0.67
	12.8

	14000
	4265
	0.65
	11.3

	15000
	4570
	0.63
	11.8

	16000
	4880
	0.61
	11.3

	17000
	5180
	0.59
	10.9

	18000
	5485
	0.57
	10.4

	19000
	5790
	0.55
	10.0

	20000
	6095
	0.54
	9.6


The contents of the cylinder do not change.

A mixture of 32% Oxygen at sea level is still 32% Oxygen at altitude. The best analogy is the balloon. The mixture cannot change but as the balloon rises the external pressure reduces and the balloon expands . 

Until  gas is exhaled a standard Scuba  rig can be considered a closed circuit

 A closed circuit system  cannot change the consistency of the gas being breathed.

A 32% mix will still be a 32% mix when it is breathed.

The divers problems are not concerned with the % mix used but the partial pressures of the gasses as they are breathed. 32% breathed at sea level has a PO2 = 0.32 bar

At 30 metres sea water depth  the  PO2 = 1.28 bar

32% breathed at  5,000 mtrs altitude  the PO2  = 0.16 bar

When diving the absolute pressure is  Gauge Pressure + Atmospheric Pressure

This equation does not change. The Gauge Pressure  depends on the depth and the Atmospheric Pressure depends on the Altitude above Sea Level.

An Analyser calibrated in Air at seal level would read 10.45 at 5,000 metres.

When measuring Oxygen at altitude the altitude must be known before the instrument can be calibrated. The same techniques should be applied with the outlet open to atmosphere and the measurement taken in still gas at ambient pressure. 

It is then possible to determine the PO2  in the gas delivered to the lungs whilst diving.

In order that the lungs remain working and do not continue to expand or collapse the pressure inside the lungs must be almost equal to the pressure outside of the lungs. 

32% at sea level  has a PO2  =  0.32

32% at 5,000mtrs  has a PO2  =  0.16

32% at 5,000mtrs diving to 30 mtrs has a  PO2  =    0.56Bar 

3 Bar water pressure +  0.5bar Atmospheric Pressure = 3.5 Bar.    x      0.16 bar . 

Altitude
Page 2
February 25, 2003

