	VANDAGRAPH
  COMPANY OPERATING PROCEDURE

	Tek-Ox Battery checks

	VM3/COP/42.07


Revision Date: 2-Sep-13
	
	Issue 1
	

	Only to be attempted by trained personel



[image: image1.jpg]



Procedure 
1. Use the modified Tek-Ox battery base.
This has connections to the +ve & -ve terminals of the battery in series with the instrument and can therefore measure the current drawn by the instrument.




2. [image: image2.jpg]


[image: image3.jpg]


Switch on using the mA range ( The uA offers too high an impedance
The reading should be less than 5mA and the LCD should show a reading












3. Switch the Tek-Ox off and change the range to uA. It should read less than 20uA


Tek- Ox pcb Battery checks.doc       02/09/2013      Page 2 of 2

