
Risk analysis report

	Company
	Ontecklong(Nanjing) Medical Supply Corporation Limited

	Address
	Bldg 10,Jinshangfang,No.288 South Xiongzhou Road, Liuhe District,Nanjing,210048 Jiangsu, P.R.China

	Product
	Neonatal Phototherapy Protection Mask

	Accessories
	N/A

	Standard
	EN ISO 14971:2009

	Conclusion

	It has been concluded through the process of risk analysis that this is a low risk device and any risks that existed were eliminated or reduced through safety testing, proper choice of materials and thorough instructions for use.

	Edited by
	Li Yun
	Date:
	9th June,2011

	Reviewed
	Li Nong
	Date
	9th June,2011

	Approved by
	Richard Yu
	Date
	9th June,2011


1. Devices features(According to EN ISO 14971:2009 Annex C）
	No
	Question
	conclusion

	1
	Intended use and how to use

	To protect neonatal baby’s eyes during phototherapy treatment

Cover the neonatal phototherapy protection mask on the baby’s eyes with adhesive tape.

	2
	Intended to be implanted
	N/A

	3
	Intended to contact patient or other person
	Contact the surface of baby’s skin

	4
	Materials/components used
	Elastic physical bonding fabric, spunlaced non-woven cloth

	5
	Energy to/from patient
	N/A

	6
	Substances to /from patient
	N/A

	7
	Biological materials processed
	N/A

	8
	Sterile/Intended to be sterilized
	N/A

	9
	routinely cleaned and disinfected by the user
	N/A

	10
	Intended to modify patient environment
	N/A

	11
	Measurements
	N/A

	12
	Interpretative
	N/A

	13
	use in conjunction with medicines or other medical technologies
	N/A

	14
	Unwanted outputs of energy or substances
	N/A

	15
	Susceptible to environmental influences
	Stored in dry environment

	16
	influence the environment
	N/A

	17
	Consumables/accessories associated
	N/A

	18
	Routine maintenance/calibration
	N/A

	19
	Software
	N/A

	20
	Restricted “shelf-life”:
	Three years

	21
	Delayed and/or long-term use effect
	N/A

	22
	Mechanical forces
	Friction between the patient and product 

	23
	Lifetime of the device determined
	Storage environment

	24
	Single use/re-use
	Single use

	25
	safe decommissioning or disposal
	N/A

	26
	Special training required to install or use
	N/A

	27
	How to offer information for using safely
	Instruction for use

	28
	new manufacturing processes need to be established or introduced
	N/A

	29
	successful application of the medical device critically dependent on human factors, such as user interface
	N/A

	29.1
	Design feature of User interface leads to use mistake
	N/A

	29.2
	Using in wrong environment due to attention distraction
	N/A

	29.3
	connecting parts or accessories
	N/A

	29.4
	control interface
	N/A

	29.5
	display information
	N/A

	29.6
	controlled by a menu
	N/A

	29.7
	Used by people with special need
	N/A


2. Relative Standard
· EN ISO 13485:2003/AC2009  Medical devices - Quality management systems - Requirements for regulatory purposes
· EN ISO 14971:2009 Medical Devices - Application of Risk Management to Medical Devices
· EN980:2008 Symbols for use in the labeling of medical devices

· EN 1041:2008 Information supplied by the manufacturer of medical devices 
· EN ISO 10993-1:2009/AC2010 Biological evaluation of medical devices-Evaluation and testing

· EN ISO 10993-5:2009 Biological evaluation of medical devices-Tests for in vitro cytotoxicity 
· ISO 10993-10:2002 AMD 2006 Biological evaluation of medical devices -Tests for irritation and delayed-type hypersensitivity
3. Risk Evaluation

	PROBABILITY OF OCCURRENCE
	Probability range
	DESCRIPTION

	FREQUENT
	≥ 10-3
	OCCURRING OFTEN OR REPEATEDLY

	PROBABLE
	< 10-3 and  ≥ 10-4
	REASONABLY LIKELY TO OCCUR

	OCCASIONAL
	< 10-4 and  ≥ 10-5
	IRREGULAR OCCURRENCE, INFREQUENT

	REMOTE
	< 10-5 and  ≥ 10-6
	NOT LIKELY TO OCCUR

	IMPROBABLE
	< 10-6 
	UNLIKELY TO EVER OCCUR



	PROBABILITY OF
OCCURRENCE
	SEVERITY CATEGORIES

	
	NEGLIGIBLE
	Minor
	Serious
	Critical
	CATASTROPHIC

	FREQUENT
	14
	7
	5
	3
	1

	PROBABLE
	16
	9
	7
	5
	2

	OCCASIONAL
	18
	12
	9
	6
	4

	REMOTE
	19
	15
	13
	11
	8

	IMPROBABLE
	20
	17
	15
	13
	10


	HAZARD RISK INDEX
	ACCEPTANCE CRITERIA

	1 TO 5
	UNACCEPTABLE

	6 TO 9
	UNDESIRABLE. WRITTEN & REVIEWED DECISION TO PROCEED

	10 TO 16
	ACCEPTABLE UPON COMPLETION OF QUALITY ASSURANCE/REGULATORY REVIEW

	17 TO 20
	ACCEPTABLE WITHOUT REVIEW


4. Risk analysis
	Potential Risk
	Potential Reason
	Risk Evaluation

（Before the risk controlling measure）
	Measure to alleviate Risk
	Response
	Risk Evaluation

（After the risk controlling measure）

	
	
	
	
	
	Severity
	Probability 

of occurrence
	Risk

	
	
	Severity
	Probability 

of occurrence
	Risk
	
	
	
	
	

	Infection to patient
	Too much bacterial
	3
	4
	ALARP
	Manufacturing environment controlling
	Instruction for manufacturing

Instruction for environment controlling
	3
	1
	ACC

	Infection to patient
	Cross-infection 

due to repeated use
	3
	4
	ALARP
	Single use marked in the label and instruction for use
	Label, Instruction for use
	3
	1
	ACC

	Stimulation to skin
	Improper 

raw material
	3
	4
	ALARP
	Select the qualified raw material


	List of raw material

Testing report 
 
	3
	1
	ACC

	Unable to block the blue light effectively
	Improper 

raw material
	3
	4
	ALARP
	Select the qualified raw material


	List of raw material

Testing report 


	3
	1
	ACC

	Damage to the product
	Destroyed package
	2
	4
	ALARP
	Use qualified packing material
	List of packing material

Description of package


	2
	1
	ACC

	Damage to the product
	Destroyed package
	2
	4
	ALARP
	Use qualified packing material
	List of packing material

Description of package


	2
	1
	ACC

	Unable to be used effectively


	Improper 

instruction for use
	3
	4
	ALARP
	Complete and detailed Instruction for use shall include 


	Labeling,

Instruction for use
	3
	1
	ACC

	Unable to be used effectively


	Not 

used by professionals
	3
	4
	ALARP
	People who use this product shall have nursing ability and understand the instruction for use


	Labeling,

Instruction for use
	3
	1
	ACC

	Hurt to baby’s skin
	Sharp acute angle
	2
	4
	ALARP
	Use soft material and wit design
	List of raw material

Design 
	2
	1
	ACC

	Damage to the product
	Shed-off 

due to movement
	3
	4
	ALARP
	Heat-sealing quality and strength testing


	Quality testing
	3
	1
	ACC


5. Conclusion
It has been concluded through the process of risk analysis that this is a low risk device and any risks that existed were eliminated or reduced through safety testing, proper choice of materials and thorough instructions for use.
Ontecklong(Nanjing) Medical Supply Corporation Limited�
No：CE/ATL-003�
�
�
Version：A/01�
�
Risk analysis report�
Validation date：9th June,2011�
�
 





SEVERITY


�
DESCRIPTION


�
�
CATASTROPHIC�
Results in patient death


�
�
Critical�
Results in permanent impairment or life-threatening injury�
�
Serious�
Results in injury or impairment requiring professional medical intervention�
�
Minor�
LOW RISK FAILURE NOT EXPECTED TO CONTRIBUTE TO AN INJURY�
�
NEGLIGIBLE�
INSIGNIFICANT FAILURE NOT SERIOUS ENOUGH TO CONTRIBUTE TO AN INJURY�
�
 










