	VANDAGRAPH
  COMPANY OPERATING PROCEDURE Tek-Ox repair
VM3COP40.41


Revision Date: 29-Oct-13
	
	Issue: 1
	



Introduction
There are two common problems apart from gross external damage or water /battery/ electrolyte damage.

No Display

Error message.

Error message. Procedure
Error messages normally require a pcb change.

It is better to replace the complete pcb+ display.

Remove the side panels

Remove the two screws holding the front panel to the enclosure

Lever the top locator off the locating block

Lever each side of the front plate locking pips over the enclosure body.

NB  Take care . On early models the small pips may break off. These have been strengthened in later batches ( approx s/n > 300)

Replace with new front panel and both left and Right side panels

NB A New round logo needs to be added.

No Display.  Procedure
Check the three screws in the battery cover.

If they turn, without tightening  the fault could be a torn thread and poor battery connection.

The torn thread can be repaired using a slither of black plastic ( from old Tek-Ox parts) inserted into the screw thread and glued with UHU.

The alternative is a new enclosure.

Some very early instruments had a long plastic locating pip for the battery connector. This should be shortened and flattened using a soldering iron.

Check the base connector for contamination and clean.

Check the springs.

The battery connector can be slightly levered up with a knife blade to put more pressure on the battery springs.

Battery springs should be cleaned particularly if a battery has leaked.

Gross leakage will require the replacement of the battery springs. NB the springs are different and if replaced must be Araldited back into position making sure there is a waterproof seal between the battery compartment and the electronic compartment.

The instrument can be checked by holding the bottom battery cover on and checking the instrument functions.

Some batch was incorrectly assembled and the calibration control is the opposite way to normal. This can be changed by changing the position of the wires on the switch. This is not recommended as it can cause further problems.

Tests
Test rigs are available to test the pcb. VN202 mkII and Tekox pcb's are interchangeable. The VN202 may have a pcb plug on the board for software changes. The Tek-ox does not need this and may may cause interference with the Tek-Ox enclosure . The plug can be cut off.

Check and calibrate instrument.

Tek-Ox repair VN3COP 40.41.doc       29/10/2013            Page 2 of 2

