	VANDAGRAPH
 SENSOR TECHNOLOGIES
Company OPERATING PROCEDURE

	Oxygen Sensor manufacture R-33VAN-4 Sensor

	VM3/COP/40.47


Revision Date: 11-Nov-13
	
	Issue 1
	

	Only to be attempted by trained personel

	Used in Analox 101D2 analyser



[image: image1.jpg]



This sensor is basically an R-33VAN with a cable 1.2mtrs long terminating in a 3.5mm Jack plug. The easiest method is to have an original and replace the sensor part To do this the Top is carefully removed. The sensor can be pushed out by pressing a round abject up the Case funnel. Unsolder the wires and replace the sensor. Re-assemble and test

	Parts Required for New version

	Quantity
	Description
	Part No.
	

	1
	R33S  no screw thread
	0110214
	

	1
	R-17 Top
	
	

	1
	R-33VAN Body
	9720200
	

	1
	O ring  VIT-0130 10-274
	9730201
	

	1
	Cable tie
	9711018
	

	1
	Co-ax cable 1.2  mtr long not critical .To fit grommet.
	
	

	1
	Tini Jack 3.5mm plastic cover
	9071007
	

	1
	Grommet 6mm
	9040031
	

	1
	Sensor Serial number label
	
	

	1
	Serial number Label for Outer bag
	
	

	1
	Polythene Bag (outer)
	
	

	1
	Gas Barrier bag
	
	


Warning  Use Soldering Iron set at 300C Do Not overheat the PCB on the Sensor
Method.
[image: image2.jpg]



1. Fit a grommet into the R-17 Top. 

2. Remove the molex and wires from the R-33S sensor

3. Strip the co-ax cable pass through the grommet and solder to the sensor. Screen to +ve Red Green to opposite end of the molex

4. Strip the other end of the 1.2mtr co-ax cable and solder the Green core to the centre pin and the screen to the screen.

5. Insert the sensor into the R-33Van body leaving 0.5 cms proud. For the R-17 cap to fit.

6. Add Silicone around the edge of the sensor face to prevent gas leakage around the sensor. 

7. Carefully fit the cap and glue into position.

8. Test for the correct out put 25mV +/- 2mV The 25mV can be adjusted with the pot on the top of the sensor.  

9. Relabel the sensor and test
R-33Van-4.doc         11/11/2013                       Page 1 of 1

