[image: image1.png]


 

                   O2 Sensors
	COMPANY OPERATING PROCEDURES


Teledyne R17MED (PN 0110017), R17A (PN 0110117), 

R17AH (PN 0110137), R17VAN (PN 0110217),
R17AV (PN 0110127), R17A-LV (PN 0110107),

O2 Sensor Test Procedure
VM3/COP/37.35
	Date:
  3-Feb-03
	Revision date: 23-Apr-14
	Issue:   3



Test & Equipment Required
Functional test required using R17 tester and R17 test adapter, (digital multimeter and 3.5mm jack plug adapter).

Output specification as in document: (VM3COP37.00, column (i)).

Method.
1. Open the packaging and remove the O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte. 

3. Connect the O2 sensor to the R17 tester, select 200mV range (marked with arrows) and wait for a few seconds for the meter reading to stabilise.

4. Check for an output  see VM3COP37.00 or Intrastats QA form for Values 
5. If no, or very low output, or sensor is returned as faulty leave the O2 sensor exposed to air for 1 Hour then re-test.

6.  If the O2 sensor fails scan them to return to supplier shelf (182560) they are sealed, initialed and output/fault written on bar code .
7. Disconnect the O2 sensor from the R17 tester and label with Viamed stickers if appropriate.

8. Scan the Barcode into the QA form in intrastats, and complete the form.

9. Reseal in serial numbered packet, 

             23/04/2014             1 of 1

[image: image1.png]