[image: image1.png]NOTES: UNLESS OTHERWISE SPECIFIED.
1. LABEL FER CLSTONER SPECFICATON
2. CONNEGTONS POLARITY ¢

e ) roame ()

WATING PLUG- WLEX P/
22013057 or Hor R ousTouER comvE
227012037

3. TENPERATURE COMPENSATION ERROR IS 45% OF FULL SGALE OVER THE OPERATIG.
TENPERATURE FANGE. WORST CASE TRAGKING ENFOR (WIMIN THE FIST HOUX
JFTER A MAXWUM TENPERATURE STCP) 15 £10% OF FULL SCALE. (548 SAWPLES
MUST BE BROUGHT TO ABIENT TEWPERATURE) PERGENT READOUT IS ONLY
WIHIN £1% AT CONSTANT PRESSURE (EG. A 0% INCREASE. I PRESSURE.

WAL RESULT W A 10% NGREASE IN FULL SGALE),

4. ASSEMELY TD BE PACKAGED IN A GAS BARRIER BAG

5. ASSEMELY NANUFACTURED AND. TESTED TO MPDG1575.

SPECFICATIONS:

1. OUTPUT ~ (7-13MV) IN AR 25C AT SEA LEVEL.
2 RANGE — 0-100% OXYGEN (MAX), 0~1% OXYGEN (MIN)

3, RESPONSE TIWE = LESS THAN 5 SEC, FOR 90% OF FINAL VALUE.

4. OPERATING TEMPERATURE RANGE — 0* TO 40 (32 TO 104'F)

5. AVG EXPECTED CELL LFE — 36 MONTHS IN AR AT 95'C, 50% RH

6. STORAGE TEMPERATURE — & TO 50°C (37" TO 12ZF)

7. HUMIDITY =0 TO 68% RH (NON-CONDENSING).

8. ACCURACY — WITHIN 1% OF FULL SCALE AT CONSTANT TEMPERATURE
AND PRESSURE (SEE NOTE 3)

9. WEIGHT ~ 1.2 OZ (32 GM)

10, OFFSET ~ < 0.1% 02 N < 30 SECONDS FOR A STEP CHANGE FROM

AR T0 0% 02.
1. REQUIRED LOAD 10K



MODEL: R-22VAN

[image: image2.png]O-RING M16 X 1mm THDS

260 = b=

SENSING SURFACE

R — |



PART NUMBER: 0110232

DATE: 23/10/2013

[image: image3.png]& VANDAGRAPH







