Calculating the Temperature range of the Tom thumb,

The only parts susceptible to temperature are the O rings and the Gauge,

 The Specifications of the O rings are -16C to 204C

 The Specifications of the Gauge is -20C to 60C

There fore the Storage / operating Temperature range of the Tom thumb is:

 -16C to 60C

Derek Lamb

