	Product Specification
	[image: image1.png]Vandagraph Sensor Technologies ~ VS§T2200 Oxygen Sensor

1K5 _SLaItion v50ad+CrI:>sI$ Hills M SIN  A123456
g t
et /N (98] CE oo NI







	Oxygen Sensor 

OOD103 / OOD103-1 / OOD103-1V / OOD103-2V


	

	Confidential - Only for Vandagraph!
	



	Measurement  Range:
	0-100 % oxygen

	Output in ambient air:
	9 to 13mV

	Electrical Interface:
	OOD103: 3pin (Molex 22-11-1031)

	
	OOD103-1: 3.5mm phone jack (Switchcraft type 41)

	
	OOD103-1V: Mini HF (RADIALL R114.426.000)

	
	OOD103-2V: Mini HF (RADIALL R114.186.000)

	Repeatability:
	< 1 % vol. O2 @ constant temp. and pressure

	Linearity error:
	0.1 to 1.0 bar less than 3 % relative

	
	0.1 to 2.5 bar less than 5 % relative

	Response time:
	< 5sec. to 90 % of final value

	Zero Offset Voltage:
	< 0.5% oxygen reading in 100 % nitrogen @ 

25°C after 36 seconds

	Cross Interference:
	< 0.5 % vol. O2 response to:

	
	5 % CO2 balance N2

	Influence of Humidity:
	- 0.03% rel. per % RH at 25°C

	Pressure Range
	600 to 2500 mbar

	Influence of Mechanical Shock:
	< 1% relative after a fall from 1m

	Operating Temperature:
	0 to 50°C

	Temperature Compensation:
	built-in NTC compensation

	Effect of Temperature Compensation (steady state):
	between 0 °C and +50 °C: 5 % relative error

	Operating Humidity:
	0-99% RH non-condensing

	Long Term Output Drift:
	< 1% vol oxygen per month

	
	typically < - 15 % relative over lifetime

	Storage Temperature:
	-20 to +50 °C

	Recommended Storage:
	+5 to +15 °C

	Recommended Load:
	( 10 kOhms 

	Warm-Up Time:
	< 30 minutes, after replacement of sensor

	Nominal Sensor Lifetime:
	( 500,000 % vol oxygen hours

	Weight:
	approximately 28 grams

	Warranty Period:
	15 months (including 3 months shelf life)

	Part No.:
	OOD103:

1001570 (REF 8010000 / VST2200)

	
	OOD103:

1001570 (REF 8010001 / APD 5)

	
	OOD103-1:

1001580 (REF 8010003 / VR1700)

	
	OOD103-1:

1001580 (REF 8010005 / VST1700)

	
	OOD103-1V:
1001581 (REF 8010002 / APD 11)

	
	OOD103-2V:
1001721 (REF 8010004 / JJ-CCR)

	All specifications  are applicable at standard conditions: 1013 mbar, 25°C dry ambient air


[image: image14.wmf]P/N OOM204


[image: image15.png]@145

345
41

M16 x 1 (5/8-24 UNEF)

34,5
41

M16 x 1 (5/8-24 UNEF)

Locking Latch
+

@29

@28

*)

3 PIN MOLEX CONNECTOR

PHONE JACK

M16 x 1 (5/8-24 UNEF)





[image: image16.png]ENVITEC

by Honeywell



[image: image17.png]160+ 5

1
M16x1 (5/8-24 UNEF)

@29+0,2




[image: image18.png]





[image: image19.png]



[image: image20.png]




Customized Design Specification:

· Integration of additional Zitex membrane in sensor gas diffusion way

· 2mm hole drilled in sensor cap near to the connector at sensor type OOD103-D

· Conformal coating on the sensor PCB with PLASTIK 70

· 2 x 2mm holes drilled in the sensor PCB

· None information about EnviteC on PCB top side / bottom side only information for ordering 

      process and traceability

PCB Layout:

Type:  8010000 / 8010001 / 8010003 / 8010004 / 8010005


        8010002


        Top


    Bottom


         

  Top


           Bottom



Customized Label Specification:

Type 8010000 (OOD103)

Sensor label

[image: image21.png]


  90x15
Blister-Card / bag label

[image: image2.png]- Vandagraph Sensor Technologles

e-mail
15 Station Road, Cross Hills Tel. +44 (0)1535 534900
Keighley, West Yorkshire
BD20 7DT, United Kingdom

VST2200 Oxygen Sensor

M 2010-06

Ce€
SIN A123456

0
50°C
4o
20°€

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor
may require 30 minutes to reach signal stability after installation.




 63,5x46,6

Type 8010001 (OOD103)

Sensor label

[image: image3.png]A.P. Diving Ag C€ ‘b@ Oxygen Cell

dive technology o 2010-06 | [sIN A123456
Pk el @epthing.com Te 0044 (01326 53 634 UATINIEW




 90x15
color: green Pantone 354 C, script in white on the customer block and symbols; black on the product info, arrays white with black frame, AP logo in black behind customer info to be discussed
Blister-Card / bag label

[image: image4.png]A.P. Diving

dive technology
Helston, Cornwall, UK

Email: sales@apdiving.com Tel: 0044 (0)1326 563 834
APD5 Oxygen Cell

] 2010-06 [1i] Cé€

SIN A123456

This oxygen cell contains caustic liquid, in case of leakage avoid contact with eyes and
skin. Do not puncture or staple this bag. Do not remove the oxygen cell from this bag
until ready to use. Depending upon the ambient conditions the oxygen cell may require
30 minutes to reach signal stability after installation.




 63,5x46,6

Type 8010002 (OOD103-1V)

Sensor label

[image: image5.png]A.P. Diving Ag C€ ‘b- Oxygen Cell

dive technology (1] 2010-06 | [N A123456
it ”s'aﬁ?éﬁﬁ'fj}wunﬁ_com Tel: 0044 (0}1326 563 834 IR0




 90x15

color: green Pantone 354 C, script in white on the customer info and symbols; black on the product info, arrays white with black frame, AP logo in black behind customer info to be discussed
Blister-Card / bag label

[image: image6.png]A.P. Diving

dive technology
Helston, Cornwall, UK

Email: sales@apdiving.com Tel: 0044 (0)1326 563 834
APD11 Oxygen Cell

] 2010-06 [1i] Cé€

SIN A123456

This oxygen cell contains caustic liquid, in case of leakage avoid contact with eyes and
skin. Do not puncture or staple this bag. Do not remove the oxygen cell from this bag
until ready to use. Depending upon the ambient conditions the oxygen cell may require
30 minutes to reach signal stability after installation.




 63,5x46,6

Type 8010002 [for samples only] (OOD103-1V) 

Sensor label

[image: image7.png]A.P. Diving Ag C€ ‘b- Oxygen Cell

dive technology (1] 2010-06 | [N A123456
it ”s'aﬁ?éﬁﬁ'fj}wunﬁ_com Tel: 0044 (0}1326 563 834 IR0




 90x15

color: black script on white 
Type 8010003 (OOD103-1) 

Sensor label

[image: image8.png]" VR Technolgy Ltd Class VR1700

Oxygen Sensor
1 Rd West
Holton Hes ith Industrial Est

] A123456
Pods D, AN CE LI T




 90x15

color: black script on white 
Blister-Card / bag label

[image: image9.png]=== VR Technolgy Ltd

Unit 12 Blackhill Rd West
Holton Heath Industrial Est
Poole, Dorset

BH16 6LE, UK

Class VR1700 Oxygen Sensor

| 2010-06 [1i] Cé€
SN A123456 5
_;?,szﬂ/

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor

may require 30 minutes to reach signal stability after installation.




 63,5x46,6

Type 8010004 (OOD103-2V) 

Sensor label

[image: image10.png]JJ CCR Aps R17JJ-CCR Oxygen Sensor
SIN  JJ123456
e o penmanc 1\ | 98] CE zm N T T




 90x15
color: black script on white 

Blister-Card / bag label

[image: image11.png]JJ-CCR ApS

Dyrlevvej 11
DK-4720 Presto
Denmark

R17JJ-CCR Oxygen Sensor

(M 2010-11

SIN  JJ123456

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor
may require 30 minutes to reach signal stability after installation.




 63,5x46,6

Type 8010005 (OOD103-1) 

Sensor label

[image: image12.png]Vandagraph Sensor Technologies  V§T1700 Oxygen Sensor

1K5_SLaItion v50ad+CrI:>sI$ Hills M SIN VS123456
g t
ANt QS erw R 11T ]




 90x15
color: black script on white 

Blister-Card / bag label

[image: image13.png]Vandagraph Sensor Technologles

e-mail: inf t.com

15 Station Road, Cross Hills Tel. +44 (0)1535 634900
Keighley, West Yorkshire
BD20 7DT, United Kingdom

VST1700 Oxygen Sensor

(M 2011-01 [ 1i] (€

SIN  VS123456

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor
may require 30 minutes to reach signal stability after installation.




 63,5x46,6


All dimensions in mm

















ENVITEC- WISMAR GMBH





Alter Holzhafen 18	


D-23966 Wismar





Phone:  +49-(0) 3841- 360 1


Fax:      +49-(0) 3841- 360 222


E-Mail: � HYPERLINK "mailto:info@envitec.com" ��info@envitec.com�


� HYPERLINK http://www.envitec.com ��http://www.envitec.com�














OOD103





OOD103-1








OOD103-2V








OOD103-1V








Document No.:001-05-OOD10300-01, 01/2011

