[image: image1.png]A.P. Diving A C€ ¢l,_ Oxygen Cell

dive teghnology (M) 2014-04 || SIN APD123456
i apdiing com AR AR

Email: sales@apdiving.com






Specification Datasheet - 8010022

Ambient Pressure Diving Ltd. 

APD16 Oxygen Sensor
© Vandagraph Sensor Technologies Ltd.

v1.0

16th April 2014

	Specifications


	VST part number
	8010022

	OEM model number
	APD16 Oxygen Sensor

	Output
	9 – 13 mV

	Output connector
	Radiall SMB Series 

R114.426.000 straight male receptacle for PCB, solder legs

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Range
	0 – 100% oxygen

	Pressure range
	600 to 2500 mbar

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25 ºC after 36 seconds

	Cross interference
	< 0.5% vol. O2 response to 5% CO2 balance N2

	Response time
	< 6 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 50 °C, 5% relative error

	Nominal sensor life
	≥ 500,000 % oxygen hours

	Long term output drift
	< 1% vol oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	≥ 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes

	Conformal coating on PCB
	Yes 

	Pressure equalization - PCB
	Yes

	Pressure equalization - output cap
	None

	Weight (approximate)
	28 g

	Flow diverter
	No

	Packaging
	Gas barrier bag - clear

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25 °C dry ambient air.



Page 2 of 4

	Drawings/images


[image: image3.wmf] 


[image: image4.wmf] 


[image: image5.png]Vandagraph
>ensor Technologies





Page 3 of 4

	Labelling


Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics - black

[image: image8.wmf] 


Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

[image: image2.png]A.P. Diving

dive technology
www.apdiving.com
Email: sales@apdiving.com

APD16 Oxygen Cell

(M| 2014-04 [1i]
SIN APD123456 —
i A\ A

This oxygen cell contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen cell from this
bag until ready to use. Depending upon the ambient conditions the oxygen cell may
require 30 minutes to reach signal stability after installation.

C€





 NOTE: Serial number prefix to be  APD e.g. APD123456
Page 4 of 4
� EMBED Word.Picture.8  ���





� EMBED Word.Picture.8  ���








Vandagraph Sensor Technologies Ltd.

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

www.vandagraphst.com

[image: image6.wmf] 

[image: image7.jpg]M16x 1(5/8-24 UNEF)

©

smB



_1380544990.doc
[image: image1.png]






_1458978213.doc
[image: image1.png]






