[image: image1.png]Vandagraph
>ensor Technologies






Specification Datasheet - 8010000

Vandagraph Sensor Technologies.

VST2200 Oxygen Sensor
[image: image2.jpg]@145
|
T
~ i
|
|
| w0
| 3 -
T
|
\
X
! M16 x 1 (5/8-24 UNEF)
@28
1 _—Locking Laich
(-)f%/ TN
h==a
$29

%\N MOLEX CONNECTOR




v1.2

10/11/11

	Specifications


	VST part number
	8010000

	OEM model number
	VST2200 Oxygen Sensor

	Output
	9 – 13 mV

	Output connector
	3 pin Molex

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Range
	0 – 100% oxygen

	Pressure range
	600 to 2500 mbar

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds

	Cross interference
	<0.5% vol. O2 response to 5% CO2 balance N2

	Response time
	< 6 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 40 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	+/- 5% of full scale over the operating temperature range. Worst case tracking error (within first hour after maximum temperature step) is +/- 7.5% of full scale, (gas sample must be brought to ambient temperature). 

Percent readout is only within +/-1% at constant pressure (e.g. a 10% increase in pressure will result in a 10% increase in reading).

	Nominal sensor life
	( 500,000 % oxygen hours

	Long term output drift
	< 1% vol oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	( 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes

	Conformal coating on PCB
	Yes

	Pressure equalization - PCB
	Yes 

	Pressure equalization - output cap
	None

	Weight
	25 g

	Flow diverter
	No

	Packaging
	Gas barrier bag

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.



Page 2 of 4

	Drawings/images


[image: image3.jpg]



[image: image4.jpg]



[image: image5.jpg]


[image: image6.jpg]s VST2200 Oxyger

M Vsl
N "l

| TTe—




Page 3 of 4

	Labelling


Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics - black

[image: image7.png]Vandagraph Sensor Technologies ~ VS§T2200 Oxygen Sensor

1K5 _SLaItion v50ad+CrI:>sI$ Hills M SIN  A123456
g t
et /N (98] CE oo NI





Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

[image: image8.png]- Vandagraph Sensor Technologles

e-mail
15 Station Road, Cross Hills Tel. +44 (0)1535 534900
Keighley, West Yorkshire
BD20 7DT, United Kingdom

VST2200 Oxygen Sensor

M 2010-06

Ce€
SIN A123456

0
50°C
4o
20°€

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor
may require 30 minutes to reach signal stability after installation.





Page 4 of 4
Vandagraph Sensor Technologies Ltd.

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

www.vandagraphst.com

