[image: image1.png]JJ CCR Aps R17JJ-CCR Oxygen Sensor
SIN  JJ123456
e o penmanc 1\ | 98] CE zm N T T






OEM Specification Datasheet - 8010004

JJ-CCR

R17JJ-CCR Oxygen Sensor
[image: image3.png]Vandagraph
>ensor Technologies





© Vandagraph Sensor Technologies

v1.0

15/03/11

	Specifications


	VST part number
	8010004

	OEM model number
	R17JJ-CCR Oxygen Sensor

	Output
	9 – 13 mV

	Output connector
	Radiall SMB Series R114.186.000

Right angled plug, crimp type

	Cable
	Belden RG174AU, length 15 cm

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Input casing cap
	8030101

	Sensor body casing
	8030200

	Output casing cap
	8030301

	Range
	0 – 2 bar PPO2

	Pressure range
	600 to 2500 mbar

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds

	Cross interference
	<0.5% vol. O2 response to 5% CO2 balance N2

	Response time
	< 6 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 40 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	+/- 5% of full scale over the operating temperature range. Worst case tracking error (within first hour after maximum temperature step) is +/- 7.5% of full scale, (gas sample must be brought to ambient temperature). 

Percent readout is only within +/-1% at constant pressure (e.g. a 10% increase in pressure will result in a 10% increase in reading).

	Nominal sensor life
	( 500,000 % oxygen hours

	Long term output drift
	< 1% vol oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	( 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes - Zitex®

	Conformal coating on PCB
	Yes – Kontakt Chemie Plastik 70

	Pressure equalization - PCB
	Yes – 2 x 3mm holes

	Pressure equalization - output cap
	Yes – 1 x 2mm hole

	Weight (approximate)
	31 g

	Flow diverter
	No

	Packaging
	Gas barrier bag - transparent

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.



Specifications subject to change 

Page 2 of 4

	Drawings/images


[image: image4.jpg]R1 7JJ -CCR oxygen Ser

(€ SN Jutol

o
AT 11





[image: image5.jpg]



[image: image6.jpg]



[image: image7.jpg]



[image: image8.jpg]


[image: image9.png]













All dimensions are in mm

	PCB layout



Page 3 of 4

	Labelling


Oxygen sensor label 

90.0mm x 15.0mm

White vinyl label, text & graphics - black

[image: image10.png]



Gas barrier bag label 

63.5mm x 46.6mm

White label, text & graphics – black

[image: image2.png]JJ-CCR ApS

Dyrlevvej 11
DK-4720 Presto
Denmark

R17JJ-CCR Oxygen Sensor

(M 2010-11

SIN  JJ123456

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor
may require 30 minutes to reach signal stability after installation.





NOTE: Serial number prefix to be  JJ e.g. JJ100001
Page 4 of 4
M16 x 1 (5/8-24 UNEF)





( 28 





( 29





41





35.5





(14.5





+ve           -ve











Vandagraph Sensor Technologies

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

