[image: image1.jpg]IMAGINATION
INNOVATION
: REALISATION

wwwirnorkedostrao.com

SENSOR TYPE: NaNS01 @ 2011-05

2§

C

S/N S01123456





Specification Datasheet - 8010007

Narked at 90 

NaNS01 Oxygen Sensor
[image: image3.png]Vandagraph
>ensor Technologies





v1.1

01/08/11

	Specifications


	VST part number
	8010007

	OEM model number
	NaNS01 Oxygen Sensor

	Output
	9 – 13 mV

	Output connector
	SMB

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Range
	0 – 100% Oxygen

	Pressure range
	600 to 2500 mbar

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds

	Cross interference
	<0.5% vol. O2 response to 5% CO2 balance N2

	Response time
	< 5 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 50 °C, 5% relative error

	Nominal sensor life
	( 500,000 % oxygen hours

	Long term output drift
	< 1% vol oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	( 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes

	Conformal coating on PCB
	Yes 

	Pressure equalization - PCB
	Yes

	Pressure equalization - output cap
	None

	Weight (approximate)
	28 g

	Flow diverter
	No

	Packaging
	Gas barrier bag - clear

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.



Page 2 of 3

	Labelling


Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics - black

[image: image4.jpg]


 
Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

[image: image2.jpg]IMAGINATION
,/,—’.\J‘ INNOVATION
Y REALISATION

wwwi.rnorkedsr90.com
SENSOR TYPE: NaNSO01

] 2011-05 [1i] C¢€
SIN S01123456 A [e@ — =
|||||||||||||||||||||||||||| A 8

This oxygen ustic liquid,

and skin D tp t st aple this ba g D t remove th oxygen
th bag u tI eady to Dependlng upon the amb nt condi t s the yg
may require 30 minutes t ach sig ability after insta II tio

of leakage avoid conta tW|th ys



 

 NOTE: Serial number prefix to be  S01 e.g. S01123456
Page 3 of 3
Vandagraph Sensor Technologies Ltd.

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

www.vandagraphst.com

