[image: image1.png]A\Z4)=1 Vobster Marine Systems ~ YMS1700

. Oxygen Sensor
Vobster Quay Diving Centre

Upper Vobster, Radstock &I SIN VMS123456
s o /N CE€  ooe  INMUMIITI

www.vmsrebreathers.com






OEM Specification Datasheet - 8010017

Vobster Marine Systems

VMS1700 Oxygen Sensor
© Vandagraph Sensor Technologies ltd.

v1.0

14th April 2014

	Specifications


	VST part number
	8010017

	OEM model number
	VMS1700 Oxygen Sensor

	Output
	9 – 13 mV

	Output connector
	*Swicthcraft 0.141" (3.57mm) range 

miniature phone jack, type No. 41

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Range
	0 – 2 bar PPO2

	Pressure range
	600 to 2500 mbar

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds

	Cross interference
	<0.5% vol. O2 response to 5% CO2 balance N2

	Response time
	< 6 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 40 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 40 °C, 5% relative error

	Average sensor life
	36 months in air at 25°C and 50% R.H.

	Long term output drift
	< 1% vol oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	≥ 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes - Zitex®

	Conformal coating on PCB
	Yes – Kontakt Chemie Plastik 70

	Pressure equalization - PCB
	Yes – 2 x 3mm holes

	Pressure equalization - output cap
	Yes – 1 x 2mm hole


	Weight (approximate)
	28 g

	Flow diverter
	No

	Packaging
	Gas barrier bag - transparent

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.



* Standard steel nickel-plated washer S-1790-1 (8030490) replaced with copper alloy nickel-plated washer P-2061 (8030491)
Specifications subject to change

Page 2 of 4

	Drawings/images


[image: image2.png]Vandagraph
>ensor Technologies





[image: image3.jpg]


[image: image4.jpg]



Page 3 of 4

	Labelling


Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics - black

Font: Century Schoolbook Greek BT (True Type)
Font file name: GrCnskn

[image: image6.jpg]345

41

M16 x 1 (5/8-24 UNEF)

PHONE JACK




Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

Font: Century Schoolbook Greek BT (True Type)
Font file name: GrCnskn

[image: image5.png]h0=] Vobster Marine Systems

Vobster Quay Diving Centre
Upper Vobster, Radstock
Somerset BA3 5SD, UK
www.vmsrebreathers.com

VMS1700 Oxygen Sensor

SIN VMS123456 I
This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from

this bag until ready to use. Depending upon the ambient conditions the oxygen sensor
may require 30 minutes to reach signal stability after installation.

Ce





Page 4 of 4
Vandagraph Sensor Technologies

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

