[image: image1.png]Vandagraph
>ensor Technologies






OEM Specification Datasheet - 8010011

VR Technology Ltd. 

VR2200 Oxygen Sensor
[image: image2.png]



Image for illustration purposes only.

Shown with VST labelling
© Vandagraph Sensor Technologies :td.

v1.1

08/12/12

	Specifications


	VST part number
	8010011

	OEM model number
	VR2200 Oxygen Sensor

	Output
	9 – 13 mV

	Output connector
	3 pin Molex

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Input casing cap
	8030101

	Sensor body casing
	8030200

	Output casing cap
	8030300

	Range
	0 – 2 bar PPO2

	Pressure range
	600 to 2500 mbar

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds

	Cross interference
	<0.5% vol. O2 response to 5% CO2 balance N2

	Response time
	< 6 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 40 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 40 °C, 5% relative error

	Average sensor life
	36 months in air at 25°C and 50% R.H.

	Long term output drift
	< 1% vol oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	> 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes - Zitex®

	Conformal coating on PCB
	Yes – Kontakt Chemie Plastik 70

	Pressure equalization - PCB
	Yes – 2 x 3mm holes

	Pressure equalization - output cap
	None

	Weight (approximate)
	28 g

	Flow diverter
	No

	Packaging
	Gas barrier bag - transparent

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.



Specifications subject to change

Page 2 of 4

	Drawings/images


[image: image3.png]



[image: image4.jpg]=== VR Technology Ltd

Py
23/25 Bridge Street,
Bailie Gate Industrial Estate
Sturminster Marshall, Dorset
BH214DB, UK

Class VR2200 Oxygen Sensor

(M 201242 [ Cé€
SIN VR123456 —
(R TR -

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor
may require 30 minutes to reach signal stability after installation.





[image: image5.jpg]VR Technology Ltd

[NESip=y=ss
Bridge Street
Bailie Gate Incustrial Estate
Sturminster Marshall. Dorset

A

3

e

Class VR2200

Oxygen Sensor

(M SN VR123456
201212 IINNINANIWNET)

BH21 4DB, UK




	PCB layout


[image: image6.jpg]@145
|
T
~ i
|
|
| w0
| 3 -
T
|
\
X
! M16 x 1 (5/8-24 UNEF)
@28
1 _—Locking Laich
(-)f%/ TN
h==a
$29

%\N MOLEX CONNECTOR



[image: image7.jpg]



Page 3 of 4

	Labelling


Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics - black

Font: Century Schoolbook Greek BT (True Type)
Font file name: GrCnskn

[image: image8.jpg]



Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

Font: Century Schoolbook Greek BT (True Type)
Font file name: GrCnskn

[image: image9.jpg]



Serial number prefix = VR       

Page 4 of 4
Vandagraph Sensor Technologies Ltd.

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

www.vandagraphst.com

