[image: image1.png]Vandagraph
>ensor Technologies






OEM Specification Datasheet - 8010003

VR Technology Ltd. VR1700 Oxygen Sensor
[image: image2.png]



© Vandagraph Sensor Technologies

v1.3

26/10/10

[image: image3.png]



	Specifications


	VST part number
	8010003

	OEM model number
	VR1700 Oxygen Sensor

	Output
	9 – 13 mV

	Output connector
	Swicthcraft 0.141" (3.57mm) range 

miniature phone jack, type No. 41

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Input casing cap
	8030101

	Sensor body casing
	8030200

	Output casing cap
	8030301

	Range
	0 – 2 bar PPO2

	Pressure range
	600 to 2500 mbar

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds

	Cross interference
	<0.5% vol. O2 response to 5% CO2 balance N2

	Response time
	< 6 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 40 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 40 °C, 5% relative error

	Average sensor life
	36 months in air at 25°C and 50% R.H.

	Long term output drift
	< 1% vol oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	( 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes - Zitex®

	Conformal coating on PCB
	Yes – Kontakt Chemie Plastik 70

	Pressure equalization - PCB
	Yes – 2 x 3mm holes

	Pressure equalization - output cap
	Yes – 1 x 2mm hole

	Weight (approximate)
	28 g

	Flow diverter
	No

	Packaging
	Gas barrier bag - transparent

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.



Specifications subject to change

Page 2 of 4

	Drawings/images


[image: image4.jpg](o)
o
—
il =
o
=)
w

41

34.5

T

M16 x 1 (5/8-24 UNEF)

PHONE JACK




[image: image5.jpg]



[image: image6.jpg]


[image: image7.jpg]i VR1700
OXygen Sensor

¢ M SN VR

» |||||||||4||mu|mn





	PCB layout


[image: image8.jpg]


[image: image9.emf]
Page 3 of 4

	Labelling


Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics - black

Font: Century Schoolbook Greek BT (True Type)
Font file name: GrCnskn

[image: image10.emf]
Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

Font: Century Schoolbook Greek BT (True Type)
Font file name: GrCnskn


Page 4 of 4
Vandagraph Sensor Technologies

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

