QC33e


	[image: image1.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid





	
	[image: image2.png]


 Ltd.
15 Station Road, Cross Hills, Keighley, West Yorkshire, BD20 7DT, UK.

Website : www.viamed.co.uk.  Email : info@viamed.co.uk.
Tel : +44 (0)1535 634542.  Fax : +44 (0)1535 635542.


TOM THUMB CALIBRATION / TEST & Q.A. SHEET.

	Description.
	Tom Thumb.
	Time & Date of Calibration / Test.
	

	Model.
	TT 495.
	Time & Date of QA.
	

	Serial No.
	
	Serial no. : O2 flowmeter.
	

	
	
	Serial no. : Air flowmeter.
	


Do not start QA check within 1 hour of the calibration / test.  Labelling is to be attached after calibration / test.

Record the manometer reading in millibars below, ensuring the TT495 meets the limits specified.

	Test Equipment
	Test
	Specification
	Reading
	P/F

	
	
	
	Cal
	Q.A.
	

	
	
	
	
	
	

	Snoop liquid.
	Check all ports / connections for leaks.
	No bubbling.
	
	
	

	
	
	
	
	
	

	CE 078.
	Adjustable Valve :

@ 15 Lpm
	Minimum :

( 7 cmH2O.
	
	
	

	
	
	
	
	
	

	CE 078.
	Adjustable Valve :

@ 15 Lpm.
	Maximum :

( 43 & ( 47 cmH2O.
	
	
	

	
	
	
	
	
	

	CE 078.
	Precision Valve :

@ 15 Lpm.
	Maximum :

+3.0 cmH2O over the adjustable valve setting.
	
	
	

	
	
	
	
	
	

	CE 078.
	Pressure Gauge Test.
	Pressure :

@ 50 cmH2O.
	
	
	

	
	
	
	
	
	

	CE079
	O2 flowmeter accuracy
	10 Lpm on O2 set on TT

( 9 & ( 11 Lpm
	
	
	

	
	
	
	
	
	

	CE079
	Air flow meter accuracy
	10 Lpm on air set on TT

( 9 & ( 11 Lpm
	
	
	

	
	
	
	
	
	

	Visual check.
	Gauge cannot be removed without tools.
	
	
	
	

	
	
	
	
	
	

	Visual check.


	Gauge appears straight.
	
	
	
	

	
	
	
	
	
	

	Visual check.
	All adjustable settings are set to a minimum.
	
	
	
	

	
	
	
	
	
	

	Visual check.
	Labels are attached.
	CE label.
	
	
	

	
	
	Viamed Flowmeter label.
	
	
	

	
	
	Serial No. label.
	
	
	

	
	
	Tom Thumb label
	
	
	


Calibration :




Q.A check :
Signed :  …………………………………                   Signed :  ………………………………           






0

