REPORT ON DATA 

TELEDYNE ANALYTICAL INSTRUMENTS MEDICAL OXYGEN SENSORS

FOR 510K COMPLIANCE

TABLE OF CONTENTS

Section
Page
  I.   
R17MED, R22MED, R13, R15, T1, T2, and T4 Sensor Tests
3 - 6

 II.
T7 Sensor Tests
7 - 8

III.
R24MED Sensor Tests
9 - 12

IV.
UFO130 Sensor Tests
13 - 36
 V.
Effects of Anesthetics Agents
37 - 39

VI.
Sensor Life
40 - 41

I

PERFORMANCE TESTS (R17MED, R22MED, R13, R15, T1, T2, T4, and T7)

INTRODUCTION:  


Groups of R17MED, R22MED, R13, R15, T1, T2, T4, and T7 sensors were subjected to the following tests: output in air, 90% response time, and accuracy. The sensors (less the T7 sensor) were tested for accuracy during operational and for accuracy after storage temperatures.

EQUIVALENCY:  


The C1R and C2R sensors are essentially the same as the T1 and T2 sensors. The only difference is that the C1R and C2R have two wire pin-connectors attached to the back-plate.  The T1 and T2 do not have the pin-connectors.  The C1R sensors utilize the same sensing membrane, cathode, electrolyte, and anode as the T1 sensors.  The C2R sensors utilize the same sensing membrane, cathode, electrolyte, and anode as the T2 sensors.  Therefore, the data for the T1 and T2 sensors can be considered equivalent for the C1R and C2R sensors.

METHODS:  


The output in each test gas was measured using a digital multimeter capable of resolving 0.01 millivolts (mV).  Response was measured by producing a step change in the flow of the sample gas between nitrogen or air (20.9% oxygen) and 100% oxygen.  The sensors were connected to flow-adaptors (in series) inside an environmental chamber and then flowing test gas at 3 SCFH.  During the response, the chamber was set at 25 deg. C and the measuring intervals of the chart recorders was set at a chart speed of 30 cm/min (equivalent to 2 seconds per cm).  The remainder of time a chart speed of 6 cm/hr (equivalent to 10 minutes per cm) was used.  Output readings were measured by flowing each test gas over the sensing surface of the test sensors and noting the mV output after stable readings were obtained (usually 10-15 minutes).  The accuracy errors were determined by noting the 100% oxygen reading (the reference reading) and then switching to 60%, 40%, 20.9% oxygen and nitrogen gas.  For air (20.9% oxygen), a gas sampling pump was used to pass room ambient air over the test sensors at a flow rate of 3 SCFH.  The rest of the test gases came from certified gas cylinders. The test sensors were allowed to stabilize with each test gas before the mV outputs were recorded.  To eliminate back pressure from an incoming test gas, a four-way switching valve was used in order to keep the gas flow constant: one direction to the cells and the other direction to by-pass vent. 


To demonstrate that the sensor meets the operation temperature requirements, the sensors were tested for accuracy by calibrating with 100% oxygen at 25 deg. C.  Then, the sensor was checked with other test gases (i.e. various oxygen concentrations).  Compressed air was allowed to flow through at a flow of 3.0 scfh for 10-15 minutes.  Then, the chamber was set to 0, 10, 15, and 40 deg. C.  After a stabilization time of 1.0-1.5 hours at each temperature, the sensor accuracy was checked again by the same procedure that was done before introduced into chamber.


To demonstrate that the sensor meets the storage temperature requirements, the sensor was stored inside the environmental chamber at 0 deg. C for 12 hours tested and then at 50 deg. C for another 12 hours.  The chamber was returned to 25 deg. C and then the sensor was tested for accuracy by calibrating with 100% oxygen at 25 deg. C.  Next, the sensor was checked with other calibration gases.  


Sensor measurement accuracy (percentage of full scale) at each temperature was determined by calculating the deviation between the 100% oxygen readings (the reference readings) and the other test gas readings one by one.  The test sensor was allowed to stabilize in each test gas before the mV output was recorded.  

EQUIPMENT AND MATERIAL:

3
R17MED SENSORS (SN 415210, 415211, and  415212)

3
R22MED SENSORS (SN 397854, 397855, and  397856)

3
R15 SENSORS (SN 393313, 393314, and  393315)

3
R13 SENSORS (SN 423430, 423432, and  423433)

3
T1 SENSORS (SN 416183, 416184, and  416185)

2
T2 SENSORS (SN 324320 and  324321)

4
T4 SENSORS (SN 412522, 412523, 412524 and  412525)

2
T4 TEST UNITS (SEN 111 - TED 80  and SEN 112 - TED 100)

1
R13 TEST UNITS (SEN 110 - TED80)

1 CAL GAS #1, 100% OXYGEN (EXTRA-DRY GRADE)

1 TEST GAS #2, 59.9% OXYGEN IN NITROGEN (certified)

1
TEST GAS #3, 40.1% OXYGEN IN NITROGEN (certified)

1
TEST GAS #4, COMPRESSED AIR, 20.9% OXYGEN 

1
TEST GAS #5, 100% NITROGEN (ULTRA-HIGH PURITY)

1
TENNEY JR. ENVIRONMENTAL CHAMBER (MODEL  TJR-17 SN 25529-16)

1
DIGITAL MULTIMETERs, FLUKE, MDL. 8020A SN CC23741and 8060A SN CC05832

RESULTS:  

The following tables summarize the results of all tests.  Attached pages contain total test results.

Table 1:  Summary of the average Air Output and 90% Response Time Test Data at 25 (C:

	Sensor
	Output in Air, 25 (C

Specification
	Average Output in Air, 

Test Results
	90% Response

Specification
	Average 90% Response Test Results

	R17MED
	10 +/-3 mV
	9.9 mV
	< 6 seconds
	3

	R22MED
	10 +/-3 mV
	9.9 mV
	< 6 seconds
	3

	R13
	0.4 +/-0.1 uA
	Test unit 0-100% O2
	< 6 seconds
	3

	R15
	59 +/-11 uA
	51 uA
	< 15 seconds
	9

	T1
	200 +/-40 uA
	213 uA
	< 30 seconds
	22

	T2
	200 +/-40 uA
	235 uA
	< 30 seconds
	20

	T4
	0.4 +/-0.2 uA
	Test unit 0-100% O2
	< 6 seconds
	4

	T7
	0.4 +/-0.1 uA
	Test unit 0-100% O2
	< 6 seconds
	4



NOTE: See next page for individual test performance results.

Table 2:  Summary of the average Operational Temperature and Accuracy Test Data at 0-40 deg. C:

	Sensor
	Accuracy Specification
	Accuracy Test Results

	R17MED
	+/- 1%
	Passed within +/- 1%

	R22MED
	+/- 1%
	Passed within +/- 1%

	R13
	+/- 2%
	Passed within +/- 2%

	R15
	+/- 1%
	Passed within +/- 1%

	T1
	+/- 1%
	Passed within +/- 1%

	T2
	+/- 1%
	Passed within +/- 1%

	T4
	+/- 2%
	Passed within +/- 2%


[image: image1.wmf]Reading Errors Observed in TAI Medical Electrochemical 

Oxygen Sensors in Various Anesthetic Agents

Sensor

Agent

Conc  

(%)

Error In 

100% O2        

(-% O2)

In 30% O2, 

70% N2O         

(-%O2)

In 30% O2, 70% 

N2O 

(calculated)*

Sens Memb 

Thickness 

(mils)

T-4

Halothane

1.5

3.5

1.75

0.6

Isoflurane

3

3

1.5

0.6

Enflurane

3

2.7

1.35

0.6

R-24MED

Halothane

6

3.4 - 3.6

0.9 - 1

0.5

Halothane

4

1 - 1.2

0.5

Halothane

3

0.6 - 0.7

0.5

Isoflurane

6

0.9 - 1.3

0.5

Isoflurane

5

2.6 - 3.6

1 - 1.1

0.5

Enflurane

6

1 - 1.8

0.5 - 0.7

0.5

Sevoflurane

7

1.1 - 1.3

0.5

Sevoflurane

6

1.1 - 1.3

0.5

Sevoflurane

5

1.1 - 1.3

0.5

Desflurane

20

1.9 - 2.2

0.2 - 0.5

0.5

Desflurane

15

1.1

0.5

R-15

Halothane

3

2.3

1.2

0.75

Isoflurane

1.4

0.7

0.4

0.75

Enflurane

2.8

1.5

0.8

0.75

C-1(R)

Halothane

5

2

1.0

1.5

Isoflurane

5

2

1.0

1.5

Enflurane

3

1

0.5

1.5

UFO 130

Halothane

5.42

2.0-2.7

1.4

0.2

Isoflurane

4.96

1.6 - 2.2

1.1

0.2

Enflurane

5

1.2 - 1.9

1

0.2

*Error in 100% O2/2



Table 3:  Summary of the average Accuracy Test Data after Storage Temperature of 0-50 deg. C:

	Sensor
	Accuracy Specification
	Accuracy Test Results

	R17MED
	+/- 1%
	Passed within +/- 1%

	R22MED
	+/- 1%
	Passed within +/- 1%

	R13
	+/- 2%
	Passed within +/- 2%

	R15
	+/- 1%
	Passed within +/- 1%

	T1
	+/- 1%
	Passed within +/- 1%

	T2
	+/- 1%
	Passed within +/- 1%

	T4
	+/- 2%
	Passed within +/- 2%


CONCLUSIONS: 

Output in air- All oxygen sensors/cells passed and had outputs within the specification range. 

90% response- All sensors/cells passed and had 90% response within the specification range.

Accuracy- Accuracy is specified as accuracy within (1% or (2% of full scale at constant temperature and pressure.  All sensors/cells passed and were within the (1% or (2% of full scale specification range.

II

TEMPERATURE TESTS

T7 SENSORS

INTRODUCTION:  


The objective of this test is to demonstrate that the T7 sensor will meet the operation temperature requirements of 0 to +40 degrees Centigrade and the storage requirement for the T7 sensor is 0 to +50 Centigrade degrees.  The storage temperature of the TED200-T7 unit requirement is -40 to +70 degrees Centigrade. Since both the unit and the sensor were tested together to the TED200-T7 unit requirement, this extreme temperature range should also covers the sensor temperature requirement. 

METHOD:


To demonstrate that the T7 sensor meets the operation temperature requirements, a TED200-T7 unit with a T7 sensor was taken out of stock and tested for accuracy.  The sensor was mounted onto a flow adaptor inside an environmental chamber at 25 deg. C.  Calibration gas (100% oxygen) was allowed to flow through at a rate of 1.0 scfh.  After the reading was stable in 100%, the TED200-T7 unit and chart recorder were calibrated to 100% full scale.  Next, the sensor was checked with other calibration gases (60%, 20.9%, and 100% nitrogen) for accuracy. 


Then, the chamber was set to +40 deg. C.  After 4 hours at this temperature, the sensor accuracy was checked again by the same procedure that was done before by calibrating with 100% oxygen and then checking the other calibration gases.  Next, the chamber was set to 0 deg. C.  After 4 hours at this temperature, the sensor accuracy was checked again by the same procedure that was done before.


To demonstrate that the sensor meets the storage temperature requirements, the TED200-T7 unit with T7 sensor was tested for accuracy by calibrating with compressed air (20.9% oxygen) at 25 deg. C.  Then, the sensor was checked with other calibration gases for accuracy.  The unit and sensor were package in its storage box in the same manner as if it was going to be stored or shipped.


Then, the storage box was placed inside the environmental chamber and the chamber was set to -40 deg. C.  After 12 hours at this temperature, the chamber was allowed to return to room temperature to 23.0 deg. C.  After removing the storage box from the chamber, the sensor was re-connected to the test unit.  The sensor accuracy was checked again by the same procedure that was done before introduced into chamber.  The sensor was re-bagged and package with the unit in its storage box in the same manner as if it was going to be stored or shipped.


Next, the storage box was placed inside the chamber and the chamber was set to +70 deg. C.  After 12 hours at this temperature, the chamber was allowed to return to room temperature to 23.8 deg. C.  After removing the storage box from the chamber, the sensor was re-connected to the test unit.  The sensor accuracy was checked again by the same procedure that was done before introduced into chamber.

EQUIPMENT AND MATERIAL:

1
TED200-T7 OXYGEN MONITOR (SN 171841)

1 T7 OXYGEN SENSOR (SN 209959)

1
COMPRESSED AIR, 20.9% OXYGEN

2 BOTTLE OF CAL GAS 100% OXYGEN

2 BOTTLE OF CAL GAS 60% OXYGEN IN NITROGEN

1 ZERO GAS, 100% NITROGEN

1
TENNEY JR. ENVIRONMENTAL CHAMBER (MODEL  TJR-17 SN 25529-16)

RESULTS:


After initial 100% Oxygen calibration at room temperature, the sensor was exposed to calibration gases of air, 60% Oxygen in Nitrogen, and Nitrogen.  After exposure to +40 deg. C (Test #1) and then to +0 deg. C (Test #2), the sensor was exposed to the calibration gases again.  The accuracy test results were ( 1% Oxygen for Test #1 and (2% Oxygen for Test #2.  The results met the (2% Oxygen accuracy requirements of the TED200-T7 Oxygen Monitor. 

A. Operational Temperature: The test results are summarized in the table below.


After initial 100% Oxygen calibration at room temperature, the sensor was exposed to calibration gases of air, 60% Oxygen in Nitrogen, and Nitrogen.  After exposure to +40 deg. C (Test #1) and then to +0 deg. C (Test #2), the sensor was exposed to the calibration gases again.  The accuracy test results were ( 1% Oxygen for Test #1 and (2% Oxygen for Test #2.  The results met the (2% Oxygen accuracy requirements of the TED200-T7 Oxygen Monitor.

Table 1:

	Test
	 
	 Accuracy Test

    temperature
	 before

cycling
	
	Accuracy

exposure to 

 for
	Test, after

+40 deg. C

4 hrs.
	Accuracy

exposure to 

 for
	Test, after

 0 deg. C 

 4 hours

	No.
	Cal with

 100% O2
	60% O2 

Reading
	20.9% O2 

Reading
	100% N2

Reading 
	60% O2 

Reading
	 100% O2

Reading
	60% O2 Reading
	100% O2 

Reading

	1
	100
	60
	21
	0.0
	61
	101
	-
	-

	2
	100
	60
	21
	0.0
	-
	-
	59
	102


B. Storage Temperature: The test results are summarized in the table below.


After each air calibration, the sensor was exposed to calibration gases of 100% Oxygen, 60% Oxygen in Nitrogen and 100% Nitrogen.  After exposure to –40 deg. C (Test #1) and then to +70 deg. C (Test #2), the accuracy test results were (1% Oxygen for both tests.  The results met the (2% Oxygen accuracy requirements of the TED200-T7 Oxygen Monitor.  

Table 2:

	Test
	 
	 Accuracy Test

    temperature
	 before

cycling
	
	Accuracy

exposure to 

 for
	Test, after

–40 deg. C

12 hrs.
	Accuracy

exposure to 

 for
	Test, after

 +70 deg. C  12 hours

	No.
	Cal with

 21% O2
	60% O2 

Reading
	100% O2 

Reading
	100% N2

Reading 
	60% O2 

Reading
	100% O2 

Reading
	60% O2 Reading
	100% O2 

Reading

	1
	21
	60
	101
	0.0
	61
	101
	-
	-

	2
	21
	60
	101
	0.0
	-
	-
	60
	101


CONCLUSION:

The sensor met the ( 2% Oxygen accuracy requirements after being subjected to the extreme operation temperature conditions of 0 to +40 degrees Centigrade.  In addition, the sensor met the (2% Oxygen accuracy requirements after being subjected to the extreme storage temperature conditions of –40 and +70 deg. C.

III

PERFORMANCE TESTS

R24MED SENSORS

INTRODUCTION:  


A group of 92 R-24s was made using final configuration molded parts and A58446 clamp rings and woven mesh membrane clamps (A55791).  A randomly selected group of 22 cells was subjected to the following tests:  output in air, 90% response time, 90% recovery time, offset in nitrogen, linearity, and vacuum linearity error.  The remaining 71 cells were labeled for shipment to selected customers and tested for output in air and vacuum linearity only. Five cells from the group of 22 were tested for temperature compensation error.

METHODS:  


The output in air was measured using a digital multimeter capable of resolving 0.01 millivolts;  the readings were rounded to the nearest 0.1 millivolt.  Response and recovery were measured by producing a step change in the flow of the sample gas between oxygen and nitrogen and visa-versa.  Three cells were tested at one time by connecting the cells in series and flowing the gases at 2 SCFH.  During the response and recovery measuring intervals a chart speed of 12 cm/min was used;  the remainder of time a chart speed of 12 cm/hr was used to conserve chart paper.  Offsets were measured by flowing nitrogen over the sensing surface of the test cells and noting the reading after 1-2 minutes.  The linearity error was determined by noting the reading above 21% in two minutes after switching from 100% oxygen to air.  A sampling gas pump was used to pass room air over the test cells at a flow rate of 2 SCFH.  Oxygen from a gas cylinder was flown at the same rate so that the back pressure at each cell (again, three sensors in series) remained constant.  The vacuum linearity error was determined by noting the output of each cell in air at ambient pressure followed by subjecting the sensing surface to a 10% vacuum.  The error was determined by computing the percent additional drop beyond the expected 10% drop:  

[image: image2.wmf]at 26ºC

0

20

40

60

80

100

120

0

50

100

Oxygen Concentration, %

Sensor Output, mV

[image: image3.wmf]at 26ºC

0

20

40

60

80

100

120

0

50

100

Oxygen Concentration, %

Sensor Output, mV

(reading at 10% vac) - (expected reading)

[image: image4.wmf]at 50ºC

0

20

40

60

80

100

120

0

50

100

Oxygen Concentration, %

Sensor Output, mV


(reading at ambient pressure)

The temperature compensation error was determined by subjecting the test cells to five temperatures within the range 0-50 deg C (0, 18, 25, 38, and 50).  25 deg C was use used as the reference temperature; the test cells remained at each temperature for 1.5-2.0 hours. Sensor measurement accuracy (percent error of full scale) were determined on the strip chart recording by measuring the deviation of each temperature reading, at the end of the 1.5-2.0 hour equilibration period, from the 25 deg C reading.

RESULTS:  The following tables list the results of all tests:

Table 1:

	Cell #
	Air Output (mv)
	90% Resp

(sec)
	90%

Rec

(sec)
	Offset

in N2
(mv)
	Linearity

Error

(%O2)
	Vac Lin

Error

(%Read)

	1
	9.7
	5.0
	5.0
	0.14
	0.0
	0.0

	2
	9.5
	5.0
	5.5
	0.13
	0.3
	0.0

	3
	9.8
	5.0
	6.0
	0.15
	0.5
	0.0

	4
	9.2
	4.5
	4.5
	0.15
	0.2
	-1.2

	5
	9.4
	6.0
	6.0
	0.14
	0.5
	0.0

	6
	9.1
	5.5
	6.0
	0.13
	0.4
	0.0

	7
	9.4
	5.0
	4.5
	0.13
	0.2
	0.0

	8
	9.5
	5.5
	5.0
	0.13
	0.7
	-1.7

	9
	9.7
	5.0
	5.5
	0.14
	0.7
	0.0

	10
	9.7
	4.5
	4.0
	0.15
	0.4
	0.0

	11
	9.3
	5.0
	6.0
	0.13
	0.4
	0.0

	12
	8.8
	6.5
	7.0
	0.14
	0.7
	0.0

	13
	8.9
	4.5
	4.5
	0.13
	0.7
	0.0

	14
	9.9
	5.0
	5.0
	0.13
	0.3
	0.0

	15
	9.7
	5.0
	5.0
	0.11
	0.4
	0.0

	16
	9.4
	4.0
	4.0
	0.12
	0.5
	0.0

	17
	9.5
	5.5
	5.5
	0.13
	0.3
	0.0

	18
	9.1
	5.0
	5.5
	0.13
	0.7
	-6.1

	19
	9.6
	5.0
	4.0
	0.11
	0.2
	0.0

	20
	9.6
	5.5
	5.5
	0.15
	0.4
	0.0

	21
	9.2
	5.0
	5.5
	0.12
	0.5
	0.0

	22
	9.9
	4.0
	4.5
	0.14
	0.4
	0.0


Table 2:
	Temperature Compensation Error (% Full Scale)

	Cell #
	25 deg C

(reference)
	0 deg C
	18 deg C
	38 deg C
	50 deg C

	1
	0.0
	0.0
	-3.0
	-2.5
	1.0

	2
	0.0
	0.0
	-2.5
	-1.5
	1.5

	3
	0.0
	-0.5
	-3.0
	-1.5
	1.0

	19
	0.0
	1.0
	-2.5
	-2.5
	1.0

	20
	0.0
	0.0
	-2.5
	-2.0
	1.5


Table 3:
	Cell # 
	Air

Output

(mv)
	Vac Lin Error

(% read)
	Cell # 
	Air

Output

(mv)
	Vac Lin Error

(% read)
	Cell # 
	Air

Output

(mv)
	Vac Lin Error

(% read)

	469009
	9.8
	0.0
	469032
	10.2
	0.0
	468960
	10.2
	0.0

	10
	9.6
	0.0
	33
	9.5
	-0.5
	61
	10.0
	-0.6

	11
	9.5
	-0.5
	34
	9.4
	0.0
	62
	9.3
	-0.7

	12
	9.9
	0.0
	35
	9.7
	0.0
	63
	9.7
	-0.1

	13
	9.6
	0.0
	36
	9.2
	-0.5
	64
	9.8
	-0.3

	14
	9.5
	0.0
	37
	9.5
	0.0
	65
	9.6
	-0.7

	15
	10.1
	0.0
	38
	9.6
	-0.5
	66
	9.3
	-0.7

	16
	9.6
	0.0
	39
	9.4
	0.0
	67
	9.6
	-1.2

	17
	9.4
	0.0
	40
	8.8
	-0.6
	68
	9.9
	0.0

	18
	9.7
	0.0
	41
	9.5
	-0.4
	69
	9.1
	0.0

	19
	9.2
	0.0
	42
	10.2
	-0.4
	70
	9.7
	0.0

	20
	9.6
	0.0
	43
	9.5
	0.0
	71
	9.4
	-0.4

	21
	9.3
	0.0
	44
	9.5
	0.0
	72
	9.2
	-0.5

	22
	9.3
	0.0
	45
	8.9
	0.0
	73
	9.8
	0.0

	23
	10.0
	0.0
	46
	9.6
	0.0
	74
	9.6
	0.0

	24
	9.4
	0.0
	47
	9.5
	-0.6
	75
	9.1
	-0.4

	25
	9.7
	0.0
	48
	8.7
	-0.4
	76
	9.4
	-0.5

	26
	10.1
	0.0
	49
	8.6
	-1.2
	77
	9.4
	-0.6

	27
	9.6
	0.0
	50
	9.3
	-0.4
	78
	9.1
	-0.1

	28
	9.9
	0.0
	51
	8.9
	0.0
	79
	9.6
	0.0

	29
	8.9
	0.0
	52
	9.3
	-0.1
	80
	9.0
	-0.4

	30
	9.1
	-0.4
	53
	10.0
	-0.2
	81
	9.1
	-0.2

	31
	9.1
	-0.5
	468959
	9.6
	0.0
	82
	9.6
	-0.5


CONCLUSIONS:  

Output in air-  The specification for the output in air of the R24 at sea level is 7-13 mv.  All 92 cells had outputs within this specified range. (8.6- 10.2 mv)

90% response and recovery-  The specification for 90% response time is less than 10 seconds.  The group of 22 cells that were tested for response and recovery time showed a range of 4.0-6.0 seconds for 90% response and a range of 4.0-7.0 seconds for 90% recovery.

Offset voltage-  The specified offset voltage is less than 200 microvolts (0.200 millivolts).  The offset in nitrogen after 2 minutes for the 22 cells tested ranged from 0.11-0.15 millivolts, all less then the specified 0.200 millivolts.

Linearity- Linearity is specified as accuracy which should be within (1% of full scale at constant temperature and pressure.  The 22 test cells ranged in error from 0.0-0.7 %O2(same as % full scale  since full scale is 100%O2.
Vacuum linearity error-  Vacuum linearity is not specified for the R24 (or any other cell except the R22cc).  In this case it was used as a research tool to determine when the clamping grid was dimensioned properly. All 92 cells were tested for this parameter.  Errors ranged from 0.0 to        -1.7% with the exception of Cell #18; which had a vacuum linearity error of -6.1%.  Even this cell would have passed the R22cc test which allows a -10% vacuum linearity error.  

Temperature compensation error- The temperature compensation error is specified at (5% full scale.  For the five cells tested, the errors ranged from-3.0 to 1.5% full scale, well within the specified range.

It is important to note that of those cells made for this production test run (92 pieces), no 

leaks were found at any of the three sealing areas:  the sensing membrane (square o-ring seal), the expansion membrane (conventional o-ring seal) and the epoxy seals around the anode and cathode wires.  During the development phase where the clamping ring was being optimized, the clamp ring was in direct contact with the sensing membrane and sometimes the membrane was cut during installation.  In the final configuration, the clamp ring does not directly touch the sensing membrane and no potential for cutting exists.

IV

PERFORMANCE TESTS 

UFO130 OXYGEN SENSOR

 INTRODUCTION:  


The objective of this test is to demonstrate that the UFO130 oxygen sensor will meet the operation temperature requirements of 10 to +50 degrees Centigrade and storage temperature requirements of 0 to +50 degrees Centigrade.  The storage requirement for the UFO sensor is 0 to +50 degrees.  

METHOD:


Different oxygen concentrations were generated by mixing compressed nitrogen and oxygen with two 0 - 300 ml/min mass flow controllers (Teledyne Hasting-Raydist, Model HFC-202E).  Both mass flow controllers were one-point calibrated with a bubble flow meter at 500 setting before use.  The total flow rate of the mixed gases was constant through the test at about 290 ml/min.  

Eight prototype sensors without sampling cap were assemble in R24-MED contact and divert caps (P/N: C60395, C60396).  By using a flow adapter, the sensors were connected in series in the gas flow with a 1/32 inch ID PE/PVC sampling tubing.  To control the operation temperature, a modified Aqua ZT-1 constant temperature chamber was used.  Sensor output was taken 5 minutes after concentration change.
For the linearity test, if the sensor is a pure diffusion controlled device, the sensor output should have a linear relationship with the oxygen partial pressure.   However, due to the electrochemical process included in the sensor, the actual output could deviate from this linear relationship.

RESULTS:  

1.  For the Linearity Test:

Figures 1 and 2 show the calibration curves of the sensors at 26ºC and 50ºC, respectively.  By using these data, percentage error of full scale due to nonlinearity when calibrated with air and 100% oxygen can be calculated by:  

[image: image5.wmf]¶

¶

¶

¶

C

t

D

C

x

=

2

2

2

4

(

.

)

[image: image6.wmf]J

d

t

DC

d

e

n

n

Dt

d

(

,

)

(

)

(

.

)

=

+

-

å

-

0

1

2

1

2

5

2

2

2

n

q

p

[image: image7.wmf]0

0.1

0.2

0.3

0.4

0.5

0.6

0.7

0.8

0.9

1

1.1

0

0.1

0.2

0.3

0.4

0.5

0.6

0.7

0.8

Time, d^2/D

Oxygen Flux, Dco/d

 


C - CR



100

   CR  =  .       79     . x (R - R100) + 100




R100 - R21


where R is the output of a sensor, R21 the output at 21% oxygen, and R100 the 


output at 100% oxygen.  Results of the calculation are listed in Table 2.1.

 [image: image8.wmf]T

C

d

D

=

2

2

6

(

.

)

[image: image9.wmf]-0.1

0

0.1

0.2

0.3

0.4

0.5

0.6

0.7

0.8

0.9

1

1.1

0

10

20

30

40

50

60

Time/sec.

Percentage Response

at 0.1Hz

at 1Hz

at 2Hz

at 3Hz


Fig. 1.  Sensor calibration curves at 26ºC.

[image: image10.wmf]Frequency, Hz

Response Magnitude


Fig. 2.  Sensor calibration curves at 50ºC.

  Table 1.  
Full scale percentage error of sensors when calibrated with air and 100% oxygen 


at the operation temperatures of 26ºC and 50ºC.

At 26ºC

	Sensor ID
	P14
	P15
	P16
	P4
	P5
	P7
	P8
	P9

	Oxygen %
	
	
	
	Error%
	
	
	
	

	21
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00

	40
	0.56
	0.45
	0.62
	0.57
	0.65
	0.58
	0.42
	0.68

	50
	0.54
	0.50
	0.68
	0.65
	0.79
	0.69
	0.41
	0.82

	60
	0.54
	0.45
	0.78
	0.78
	0.89
	0.80
	0.41
	0.85

	80
	0.44
	0.53
	0.70
	0.76
	0.82
	0.79
	0.41
	0.73

	100
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00


At 50ºC

	Sensor ID
	P14
	P15
	P16
	P4
	P5
	P7
	P8
	P9

	Oxygen %
	
	
	
	Error%
	
	
	
	

	21
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00

	40
	0.40
	0.49
	0.47
	0.62
	0.76
	0.68
	0.45
	-2.66

	60
	0.34
	0.57
	0.67
	0.85
	0.97
	0.97
	0.45
	-1.48

	80
	0.62
	0.70
	0.70
	0.92
	0.85
	0.99
	0.54
	-0.36

	100
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00


As shown in table 2.1, when calibrated with air and oxygen, the biggest error shows at 60% oxygen and is less than 1%.  However, when one-point calibration is used at 100% oxygen, the full scale percentage error can then be calculated as:

[image: image11.wmf]0

5

10

15

20

25

490

510

530

550

570

590

610

630

T10-90/ms

Percentage

[image: image12.wmf]0

5

10

15

20

25

320

350

380

410

440

470

500

530

560

T10-90/ms

Percentage

[image: image13.wmf]450

500

550

600

650

58

60

62

64

66

68

70

72

Sensor Output in Oxygen, mV

T

10-90

, ms

 


C - CR



100

   CR  =        100      x R 




   R100 
As shown in Table 2, with one point calibration at 100% oxygen, the biggest error happens at 50% oxygen and is as high as 1.6%.

  Table 2. Full scale percentage error of sensors when one-point calibrated with 100% oxygen at the operation temperatures of 26ºC and 50ºC.

At 26ºC

	Sensor ID
	P14
	P15
	P16
	P4
	P5
	P7
	P8
	P9

	Oxygen %
	
	
	
	Error%
	
	
	
	

	0
	0.24
	0.21
	0.24
	0.21
	0.23
	0.21
	0.28
	0.22

	21
	0.96
	1.04
	1.27
	1.25
	1.32
	1.36
	1.09
	1.23

	40
	1.28
	1.24
	1.58
	1.50
	1.64
	1.61
	1.24
	1.61

	50
	1.14
	1.15
	1.47
	1.43
	1.61
	1.54
	1.10
	1.59

	60
	1.02
	0.97
	1.41
	1.40
	1.54
	1.48
	0.96
	1.46

	80
	0.67
	0.79
	1.01
	1.06
	1.14
	1.12
	0.68
	1.03

	100
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00


At 50ºC

	Sensor ID
	P14
	P15
	P16
	P4
	P5
	P7
	P8
	P9

	Oxygen %
	
	
	
	Error%
	
	
	
	

	0
	0.22
	0.20
	0.23
	0.00
	0.00
	0.20
	0.23
	0.20

	21
	0.85
	0.84
	0.95
	1.06
	1.09
	1.11
	0.81
	8.99

	40
	1.04
	1.12
	1.19
	1.42
	1.58
	1.51
	1.06
	4.47

	60
	0.76
	0.99
	1.15
	1.37
	1.51
	1.52
	0.85
	3.24

	80
	0.83
	0.90
	0.94
	1.18
	1.12
	1.26
	0.74
	1.96

	100
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00
	0.00


2.  For the 10 - 90% Response Time Test:

Fick's first law can only provide us information of diffusion at static state.  To study the transient behavior of the diffusion, we have to depend on Fick's second law

[image: image14.wmf]0.000

500.000

1000.000

1500.000

2000.000

2500.000

3000.000

3500.000

4000.000

0.00

0.50

1.00

1.50

2.00

2.50

3.00

Time/sec.

Sensor Output, mV

[image: image15.wmf]15

20

25

30

35

40

20

25

30

35

40

45

50

Temperature, ºC

Output in Air, mV   


By setting the boundary condition of C(x, t=0)=0 and C(0, t=0+)=C0, equation (2.4) can be solved as:

[image: image16.wmf](

'

)

(

'

)

(

.

)

S

S

T

T

S

-

-

23

3

4


[image: image17.wmf]15

20

25

30

35

40

20

25

30

35

40

45

50

Operation Temperature, ºC

Output in Air, mV

Without DC Temperature

Compensation

With DC Temperature

Compensation


Equation (2) describes the relationship between the sensor output and time when oxygen concentration changes from zero to C0.  Figure 3.1 shows the plot of J(d, t) in normalized units of DC0/d versus t in normalized units of d2/D.  

[image: image18.wmf]S-11 

400

450

500

550

600

650

50

100

150

200

250

300

350

Flow Rate, cc/min

T10-90%, ms

Cap-1

Cap-2

Cap-3


     Fig. 3. 
Response curve at the the boundary condition of C(x, t=0)=0 and

 C(0, t=0+)=C0 with Fick's second law.

As shown in equation (2) and figure 3, the response time T of a sensor in only the function of the diffusion coefficient and the membrane thickness which can be described as:

[image: image19.wmf]-90

-80

-70

-60

-50

-40

-30

-20

-10

0

0

50

100

150

200

250

300

350

Time, hour

Weight Loss, mg

40ºC

23ºC

40ºC

[image: image20.wmf]Output and Response Time (air to 100% O2) Testing:

Gases:

Compressor Air and 100% O2 flows @ 3 scfh

A.  R17MED Sensor:

E

. 

T1 Sensor:

Air, mV

Response

Air, uA

Response

S/N

 @25 deg. C

90%, sec.

S/N

 @25 deg. C

90%, sec.

415210

10.0

3

416183.0

217

21

415211

9.6

3

416184

222

20

415212

10.1

3

416185

200

24

Avg.=

9.9

3

Avg.=

213

22

Spec. =

10 +/-3 mV

< 6 sec.

Spec. =

200 +/-40 uA

< 30 sec.

B.  R22MED Sensor:

F

. 

T2 Sensor:

Air, mV

Response

Air, mV

Response

S/N

 @25 deg. C

90%, sec.

S/N

 @25 deg. C

90%, sec.

415210

10.0

3

324320

239

19

415211

9.6

3

324321

230

20

415212

10.1

3

Avg.=

235

20

Avg.=

9.9

3

Spec. =

200 +/-40 uA

< 30 sec.

Spec. =

10 +/-3 mV

< 6 sec.

C. R13 Sensor:

G. T4 Sensor:

Reading Air

Response

Reading 100%O2

Response

S/N

 @25 deg. C

90%, sec.

S/N

 @25 deg. C

90%, sec.

423430

100

4

412522

100

4

423432

100

3

412523

100

3

423433

100

3

Avg.=

100

4

Avg.=

100

3

Spec. =

100

< 6 sec.

Spec. =

100

< 6 sec.

D.  R15 Sensor:

H. T7 Sensor:

Air, uA

Response

Reading 100%O2

Response

S/N

 @25 deg. C

90%, sec.

S/N

 @25 deg. C

90%, sec.

393313

51.1

10

171841

100

4

393314

50.8

9

Spec. =

100

< 6 sec.

393315

51.8

9

Avg.=

51.2

9

Spec. =

59 +/-11 uA

< 15 sec.


Where C is a constant.  From equation (3),it is clear that with the same sensing membrane material, the thinner the sensing membrane, the shorter the response time.  From the supplier of our oxygen sensing membrane, the thinnest membrane available is 0.17 mil in thickness.  

Our current product R24-MED uses 0.5mil sensing membrane, and it's 10% to 90% response time T10-90 were measure to be 1.35 second.  The thickness of the 0.17 mil membrane was measured to be 0.20 mil.  With these testing data and the relation shown in equation (3), a sensor with a T10-90 at several hundreds millisecond range can be achieved.

The response speed of the micro-fuel cell oxygen sensor can also be represented by the frequency response of the sensor.  Figure 4. shows a typical response curve of a UFO130 sensor at different frequency of oxygen concentration change from 21% to 10%.  

[image: image21.wmf]at 26ºC

0

20

40

60

80

100

120

0

50

100

Oxygen Concentration, %

Sensor Output, mV


Fig. 2.2.  
Response of a prototype UFO130 sensor at different oxygen switching 

frequencies.

As we can see, at low frequency, the sensor exhibits a full response to the oxygen concentration changes.  However, when the frequency get closer to and higher than  1/T10-90%, the sensor response starts to attenuate.  If we plot the amplitude of the sensor response versus the frequency, we'll get a curve somewhat like the one shown in Figure 2.3.  Because the sensor response is controlled by a well-defined process, the attenuation of the response can be compensated electronically.  In other word, the response speed of the sensor can be electronically enhanced.  The frequency compensation is a well 

[image: image22.wmf]at 26ºC

0

20

40

60

80

100

120

0

50

100

Oxygen Concentration, %

Sensor Output, mV

[image: image23.wmf]at 50ºC

0

20

40

60

80

100

120

0

50

100

Oxygen Concentration, %

Sensor Output, mV


Fig. 2.3.  Frequency response curve of a micro-fuel oxygen sensor.

developed technology.  The combination of this technology with the micro-fuel cell oxygen sensor is a novel approach in this product development.

A good way to visualize the concept is to view the actual micro fuel cell sensor as series combination of an ideal sensor (one with unlimited high frequency response) and a sensor low pass filter which matches the high frequency attenuation of the actual micro-fuel cell sensor.  

Viewing the frequency domain, the ideal sensor would have a flat frequency response.  That is its output amplitude for a sinusoidal input oxygen concentration would be the same for any frequency of input oxygen concentration.   Viewing the time domain, an ideal step of oxygen concentration into the ideal sensor would yield an ideal step output (zero rise time with no overshoot or ringing).

On the other hand, in the frequency domain, the sensor low pass filter would attenuate the high frequency components of the ideal sensor output.  In the time domain, the sensor low pass output would be a step response with a finite rise time when its input had zero rise time. 

This visualization is shown in Figure 2.4 below:

[image: image24.wmf]¶

¶

¶

¶

C

t

D

C

x

=

2

2

2

4

(

.

)


[image: image25.wmf]J

d

t

DC

d

e

n

n

Dt

d

(

,

)

(

)

(

.

)

=

+

-

å

-

0

1

2

1

2

5

2

2

2

n

q

p

[image: image26.wmf]T

C

d

D

=

2

2

6

(

.

)


[image: image27.wmf]0

0.1

0.2

0.3

0.4

0.5

0.6

0.7

0.8

0.9

1

1.1

0

0.1

0.2

0.3

0.4

0.5

0.6

0.7

0.8

Time, d^2/D

Oxygen Flux, Dco/d

[image: image28.wmf]Frequency, Hz

Response Magnitude


[image: image29.wmf]-0.1

0

0.1

0.2

0.3

0.4

0.5

0.6

0.7

0.8

0.9

1

1.1

0

10

20

30

40

50

60

Time/sec.

Percentage Response

at 0.1Hz

at 1Hz

at 2Hz

at 3Hz


[image: image30.wmf]0

5

10

15

20

25

320

350

380

410

440

470

500

530

560

T10-90/ms

Percentage

[image: image31.wmf]0

5

10

15

20

25

490

510

530

550

570

590

610

630

T10-90/ms

Percentage

[image: image32.wmf]450

500

550

600

650

58

60

62

64

66

68

70

72

Sensor Output in Oxygen, mV

T

10-90

, ms

[image: image33.wmf]Output and Response Time (air to 100% O2) Testing:

Gases:

Compressor Air and 100% O2 flows @ 3 scfh

A.  R17MED Sensor:

E

. 

T1 Sensor:

Air, mV

Response

Air, uA

Response

S/N

 @25 deg. C

90%, sec.

S/N

 @25 deg. C

90%, sec.

415210

10.0

3

416183.0

217

21

415211

9.6

3

416184

222

20

415212

10.1

3

416185

200

24

Avg.=

9.9

3

Avg.=

213

22

Spec. =

10 +/-3 mV

< 6 sec.

Spec. =

200 +/-40 uA

< 30 sec.

B.  R22MED Sensor:

F

. 

T2 Sensor:

Air, mV

Response

Air, mV

Response

S/N

 @25 deg. C

90%, sec.

S/N

 @25 deg. C

90%, sec.

415210

10.0

3

324320

239

19

415211

9.6

3

324321

230

20

415212

10.1

3

Avg.=

235

20

Avg.=

9.9

3

Spec. =

200 +/-40 uA

< 30 sec.

Spec. =

10 +/-3 mV

< 6 sec.

C. R13 Sensor:

G. T4 Sensor:

Reading Air

Response

Reading 100%O2

Response

S/N

 @25 deg. C

90%, sec.

S/N

 @25 deg. C

90%, sec.

423430

100

4

412522

100

4

423432

100

3

412523

100

3

423433

100

3

Avg.=

100

4

Avg.=

100

3

Spec. =

100

< 6 sec.

Spec. =

100

< 6 sec.

D.  R15 Sensor:

H. T7 Sensor:

Air, uA

Response

Reading 100%O2

Response

S/N

 @25 deg. C

90%, sec.

S/N

 @25 deg. C

90%, sec.

393313

51.1

10

171841

100

4

393314

50.8

9

Spec. =

100

< 6 sec.

393315

51.8

9

Avg.=

51.2

9

Spec. =

59 +/-11 uA

< 15 sec.


[image: image34.wmf]0.000

500.000

1000.000

1500.000

2000.000

2500.000

3000.000

3500.000

4000.000

0.00

0.50

1.00

1.50

2.00

2.50

3.00

Time/sec.

Sensor Output, mV

[image: image35.wmf]15

20

25

30

35

40

20

25

30

35

40

45

50

Temperature, ºC

Output in Air, mV   

[image: image36.wmf](

'

)

(

'

)

(

.

)

S

S

T

T

S

-

-

23

3

4


[image: image37.wmf]15

20

25

30

35

40

20

25

30

35

40

45

50

Operation Temperature, ºC

Output in Air, mV

Without DC Temperature

Compensation

With DC Temperature

Compensation


Fig. 2.4.  
Schematic view of the operation of a micro-fuel sensor at a step 

change of oxygen concentration.

Ideally, the electronic compensation circuit would be the exact inverse of the sensor low pass filter and thus cancel the effect of the sensor low pass filter and yield  a product with perfectly flat frequency response to infinitely high frequency.  

The system transfer function (in the "s" plane) is the product of the transfer function of the cell low pass filter and the electronic compensation circuit.  If the transfer function of the cell low pass filter were a simple single pole factor, for this idealized case the transfer function of the electronic compensation circuit would be a simple single zero factor, which would cancel the pole and yield an ideal transfer function of "1".  This is shown below:

system transfer function = 

           (cell low pass filter transfer function) x (electronic compensation circuit transfer function)

Gsys(s) = Gcell(s) x Gcomp(s)

Gsys(s) = (   1 / (1 + sT1)   ) x  (   (1 + sT1)   ) = 1

Now an electronic compensation circuit transfer function of "(1 + sT1)" can not be implemented because of finite amplifier gain and finite operational amplifier bandwidth.  Also it would not be desirable to implement because it would greatly amplify noise voltages.

However:        Gcomp(s) = (   (1 + sT1) / (   (1 + sT2) x (1+sT3)   )   )   can be implemented.

Let T2 = T1 / N   (where N > 1) and  T3 < T2 / 5

Now:

Gsys(s) = (   1 / (1 + sT1)   ) x (   (1 + sT1) / (  (1 + sT2) x (1+sT3)  )   ) 

Gsys(s) = (   1 /  (  (1 + sT2) x (1+sT3)  )   )

Since T3 < T2 / 5,        Gsys(s) 

 (1 /  (1 + sT2)) = (1 / (1 + s(T1 / N) ) ) [for the frequencies of interest]

The net result is that the rise time of the system is "N" times smaller than the cell alone!

The actual sensor low pass filter transfer function is more complicated than a simple single pole and it varies with temperature and from one cell to another, however, the basic compensation procedure is the same: for the frequencies of interest, make the electronic compensation circuit  transfer function close to the inverse of the sensor low pass filter.  

1.2. Validation.

3.2.1. Experiment setup.

The sensor response time was measured by using a gas switching device made in house.  Figure 3.1 shows a schematic drawing and working mode of the gas switching device.

[image: image38.wmf]S-11 

400

450

500

550

600

650

50

100

150

200

250

300

350

Flow Rate, cc/min

T10-90%, ms

Cap-1

Cap-2

Cap-3

[image: image39.wmf]-90

-80

-70

-60

-50

-40

-30

-20

-10

0

0

50

100

150

200

250

300

350

Time, hour

Weight Loss, mg

40ºC

23ºC

40ºC


[image: image40.wmf]Reading Errors Observed in TAI Medical Electrochemical 

Oxygen Sensors in Various Anesthetic Agents

Sensor

Agent

Conc  

(%)

Error In 

100% O2        

(-% O2)

In 30% O2, 

70% N2O         

(-%O2)

In 30% O2, 70% 

N2O 

(calculated)*

Sens Memb 

Thickness 

(mils)

T-4

Halothane

1.5

3.5

1.75

0.6

Isoflurane

3

3

1.5

0.6

Enflurane

3

2.7

1.35

0.6

R-24MED

Halothane

6

3.4 - 3.6

0.9 - 1

0.5

Halothane

4

1 - 1.2

0.5

Halothane

3

0.6 - 0.7

0.5

Isoflurane

6

0.9 - 1.3

0.5

Isoflurane

5

2.6 - 3.6

1 - 1.1

0.5

Enflurane

6

1 - 1.8

0.5 - 0.7

0.5

Sevoflurane

7

1.1 - 1.3

0.5

Sevoflurane

6

1.1 - 1.3

0.5

Sevoflurane

5

1.1 - 1.3

0.5

Desflurane

20

1.9 - 2.2

0.2 - 0.5

0.5

Desflurane

15

1.1

0.5

R-15

Halothane

3

2.3

1.2

0.75

Isoflurane

1.4

0.7

0.4

0.75

Enflurane

2.8

1.5

0.8

0.75

C-1(R)

Halothane

5

2

1.0

1.5

Isoflurane

5

2

1.0

1.5

Enflurane

3

1

0.5

1.5

UFO 130

Halothane

5.42

2.0-2.7

1.4

0.2

Isoflurane

4.96

1.6 - 2.2

1.1

0.2

Enflurane

5

1.2 - 1.9

1

0.2

*Error in 100% O2/2












Position-1





Position-2

Fig. 3.1.  Schematic drawing and working mode of the gas switching device.  At position-1, the device is delivering gas-1, while at position-2, the device is delivering gas-2. 

The full travel length of the piston from position-1 to position -2 was about 1/4 inch.  The opening for gas delivering was about 1/32 inch.  Two pull type tubular solenoid (series S-15-75, Magnetic Sensor System, Van Nuys, CA) were used to drive the piston.  Transient current change of the solenoid during operation showed that the travel time of the piston from position-1 to position-2 was about 7 milliseconds.  

To measure the sensor response time, sensor sampling cap was attached to the sensor.  A PE/PVA, 1/16 inch in OD and 2 inch in length, was used to connect the sampling cap and the gas switch outlet.  By switching the gas from gas-1 to gas-2, sensor response curve was recorded through a data acquisition system.  Unless it is specified, gas-1 is compressed air and gas-2 is compressed oxygen.  Response time T10-90 is defined as the time taken from 10% to 90% of the full output change.

3.2.2.  Results.

(a) (b)

     Fig. 3.2. 
Response time distribution of (a) sensor S-XX + P-XX, and (b) sensor 

PT-XX at room temperature.

Figure 3.2 shows the response time distribution of different batches of sensors.  The sensor response time varies from 320 to 630 milliseconds.  By comparing Figure 4.4. (a) and (b), it is clear that the distribution range also changes from batch to batch.  The reason for this variation is not clear yet.

DISCUSSION: According to equations (2.3) and (2.6), both sensor output and sensor response time is related to the properties of the sensing film.  So, any variation of the output should be reflected by the variation of the response time.  Assuming the diffusion coefficient more constant from sensor to sensor, then we should expect that the higher the sensor output, the shorter the response time.  In reality, this relationship was not strictly followed.  Figure 3.3 shows a typical scattered relationship between sensor output and response time at room temperature.  This scattered relationship also changes from sensors batch to batch.


    Fig. 3.3.  
Relationship between output and response time of P-XX sensors at 

room temperature.

Test results of the prototype sensors after frequency compensation at room temeprature and 100 ml/min gas flow rate are shown in the following table.

Table. 3.1. Test results of prototype UFO130 unit.

	Sensor S/N
	T10-90, ms
	Output in Oxygen, V
	Sensor T10-90 Range

	PT2
	100
	3.73
	3

	PT3
	100
	3.54
	3

	PT4
	80
	3.66
	3

	PT5
	90
	3.67
	3

	PT6
	100
	3.62
	3

	PT7
	100
	3.81
	2

	PT8
	105
	3.67
	3

	PT9
	100
	3.43
	3

	PT10
	90
	3.63
	3

	PT11
	140
	3.49
	3

	PT12
	90
	3.52
	3

	PT13
	100
	3.52
	3

	PT14
	100
	3.53
	3

	PT15
	100
	3.73
	2

	PT16
	100
	3.88
	2

	PT17
	100
	3.78
	3

	PT18
	100
	3.46
	3

	PT19
	120
	3.48
	3

	PT20
	100
	3.66
	3

	PT21
	110
	3.52
	3

	PT22
	130
	3.54
	3

	PT23
	100
	3.69
	3

	PT24
	150
	3.48
	3

	PT25
	120
	3.55
	3

	PT26
	110
	3.67
	3

	PT27
	130
	3.60
	3

	PT28
	120
	3.60
	3

	PT29
	110
	3.83
	3

	PT30
	120
	3.34
	3

	PT31
	110
	3.51
	3

	PT32
	100
	3.62
	3

	PT33
	100
	3.62
	3

	PT34
	110
	3.75
	3

	PT35
	90
	3.67
	3

	PT36
	140
	3.42
	3

	PT37
	100
	3.66
	3

	PT38
	80
	3.71
	3

	PT39
	105
	3.67
	3

	PT40
	115
	3.12
	3

	PT41
	100
	3.74
	3

	PT42
	100
	3.67
	3

	PT43
	95
	3.77
	3

	PT44
	100
	3.66
	3

	PT45
	90
	3.71
	3

	PT46
	90
	3.47
	3

	PT47
	160
	3.10
	3

	PT48
	100
	3.75
	3


4. Cross Interference.

4.1. Verification.  None

1.3. Validation.

4.2.1. Experiment Setup.

4.2.1.a. Gas Mixture.  Table 4.1 summarizes information of the gas mixtures used for this test.

Table 4.1. Concentration of the Test Gas Mixtures

	Gas Mixture
	Certified concentration
	Manufacturer

	
	
	

	Halothane/oxygen
	5.42
	Scott Medical Product

	Enflurane/oxygen
	5.00
	Scott Medical Product

	Isoflurane/oxygen
	4.96
	Scott Medical Product

	N2O/oxygen
	80.12%
	Mathson

	Helium/oxygen
	79.4%
	Mathson

	CO2/oxygen
	10.09%
	Mathson


The concentration of halothane from ISO7767 is required to be 6%.  However, Technical staff from Scott informed us that 6% of halothane in oxygen is over the low explosion limit and unsafe.  The highest concentration of halothane in oxygen is 5.5%.

4.2.1.b. Sensor and Gas Flow System.  Eight prototype sensors, P1, P2, P3, 

P4, P5, P7, P8, and P9, were used for the test. The gas flow connection was the same as described in section 4.2.  The gas flow rate is controlled by using tube flow meters.  Tubing used were 1/4 inch OD poly flo (Imperial EastmanTM) and 1/16 inch OD PE/PVA tubing.  No other material tubing should be used because of the corrosive property of the anaesthetic agent.

4.2.1.c.  Temperature Control.  Except for that the CO2 test was done at room temperature, all the other tests were done at controlled temperature.  When tested at controlled temperature, the sensors were placed in a modified Aqua ZT-1 chamber with temperature controlled by an RTD and Omron temperature controller.

4.2.1.d. Test Procedure.  Set the temperature of the chamber and flow the blank gas 

Fig. 4.1.  Test data during a isoflurane interference test at 40ºC.

(100% oxygen or compressed air) through the sensors.  Let the sensor reach equilibrium with the chamber.  Start data recording when the sensor gives a steady output, and switch the gas to the anesthesia gas mixture and hold it for two hours.  Switch the gas back to the blank gas.  Stop data record when the sensors give a stable output.  The error is calculated by calibrate the sensor at the beginning of the test with the blank gas, then substrate the oxygen concentration the sensor reads with the real oxygen concentration during the test.  Figure 4.1 shows the typical data of this test.

4.2.2.  Results

Table 4.2 summarizes the test results of errors caused by the interfering gases at different temperatures.  As shown in table 4.2, at 34ºC for halothane, all the sensors still could not meet the 2% error requirement.  At 37ºC for halothane, some sensor met and some did not meet the requirement.  At 40ºC, all the sensors met the 2% requirement for halothane, and enflurane gases.  For Isoflurane at 40ºC, two sensors showed 2.2% error at the worst case.  

Table 4.2.  Errors caused by anesthesia gases at different temperatures.

	Interfering gases
	Error at different temperatures, %

	
	23ºC
	34ºC
	37ºC
	40ºC

	Halothane
	
	2.0 - 2.7
	1.9 - 2.4
	1.5 - 1.9

	Enflurane
	
	
	
	1.2 - 1.9

	Isoflurane
	
	
	
	1.6 - 2.2

	Helium
	
	
	
	0.5 - 0.8

	Nitrous oxide
	
	
	
	0.3 - 0.5

	Carbon dioxide
	1.3 - 1.5
	
	
	


5. Storage Tempaerature.

5.1. Verification. Refer to the storage temperature of  standard Teledyne oxygen sensors.

5.1. Validation. Refer to the next section (operation temperature).

6. Operating Temeprature.

6.1. Verification. Refer to the operating temperature of standard Teledyne oxygen sensors.

6.2. Validation.  

6.2.1. Experiment setup.  The test was done by placing the UFO130 unit in a environmental chamber and use the test setup of section 3.

6.2.2. Results.  Table 6.1 shows the response speed and output of 4 prototype sensors tested at 10, 23, and 50ºC.

Table 6.1.  Sensor performance at different operating temperatures.

	Sensor
	T10-90 Response Time, ms
	Output in air
	Output in O2

	
	10ºC
	23ºC
	50ºC
	10ºC
	23ºC
	50ºC
	10ºC
	23ºC
	50ºC

	PT35
	110
	100
	100
	0.76
	0.77
	0.79
	3.54
	3.60
	3.69

	PT36
	120
	120
	80
	0.69
	0.69
	7.51
	3.24
	3.26
	3.5

	PT41
	110
	90
	70
	0.75
	0.75
	0.79
	3.50
	3.50
	3.69

	PT42
	110
	100
	80
	0.72
	0.74
	0.77
	3.34
	3.45
	3.58




Fig. 6.1.  Response curves of PT35 unit at different temperatures.

Figure 6.1 shows the response curves of the PT35 unit at different operation temperatures.  The results demonstrate that the UFO130 is operatable at the temperature range of 10 -50ºC.

7. Operating Humidity.

7.1. Verification.  Refer to R24-MED ISO test results.

7.2. Validation. Refer to R24-MED ISO test results.

8. Temperature Compensation Error

8.1. Verification. 


  Fig. 8.1.  Temperature effect on the output of nine prototype sensors.

Ten sensors were built with R24 bodies.  By connecting the sensor output leads with a 100(1% ohm resistor, sensor output at different temperature in air were measured.  Figure 3.4 shows the temperature effect on the output of the nine prototype sensors at the temperature range from 23ºC to 45ºC.  In this temperature range, test showed that the change of the resistor value was less than 0.2%.  The following formula was used to calculate the percentage change per degree centigrade:


where S23 represents the sensor output AT 23ºC, S' represents the output at T', and  S represents the output at  temperature T.  This percentage change per degree centigrade actually represents the slop of the plots shown in Figure 3.4.  The calculated results are listed in table 8.1.

  Table 8.1.  
Percentage output change per degree centigrade at the temperature range 

From 23 to 45ºC.

	Sensor ID
	S11
	S12
	S13
	S14
	S15
	S16
	S17
	S18
	S19

	
	
	
	
	
	
	
	
	
	

	23 - 30 ºC
	2.1
	2.2
	2.3
	2.3
	2.0
	2.1
	2.0
	2.0
	2.1

	30 - 35 ºC
	2.6
	2.7
	2.7
	2.8
	2.9
	3.1
	3.1
	3.1
	3.3

	35 - 40 ºC
	3.3
	3.2
	3.7
	3.7
	3.2
	3.1
	3.0
	3.0
	3.2

	40 - 45 ºC
	3.3
	3.7
	3.1
	3.5
	3.6
	3.6
	3.6
	3.6
	3.9


As being discussed in concept section, temperature dependence of the sensor output results from temperature effects on the diffusion coefficient.  All sensor with the cast-Teflon sensing membrane should have the same temperature dependence.  

The data showed in Table 8.1 agrees with the behavior of R24 MED sensors.  So, the temperature compensation network (P/N B65388) design for R24-MED sensor were used in this project.  

8.2. Validation. Typical temperature compensation effect is shown in Figure 8.2.

  Fig. 3.5.  Output of S18 sensor in air at different temperature with and without the temperature compensation.

Based on the test data shown in table 6.1, the temperature compensation error in the operation temperature range is calculated and listed in table 8.2.


Table 8.2.  DC temperature compensation error (percentage of reading) in the 10 - 

50ºC range.

	Sensor Unit
	PT35
	PT36
	PT41
	PT42

	Total error, %
	4.0
	8.0
	5.4
	6.9


As shown in table 8.2, the total error from 10ºC to 50ºC is less than 10%.  So when calibrated at certain temperature within the temperature range, the temperature compensation will meet (5%.

9. Flow Rate of Gas Sampling.

9.1. Verification.

In this project, gas sampling system was called sampling cap.  Three gas sampling systems were developed and tested.  Figure 9.1 shows the schematic view of the caps.

























   Cross-section



        Bottom

Fig. 9.1.  Schematic view of the sampling caps.

Cap-1 was the original design.  It forms a big dead volume above the sensing surface and introduces an assymmetric gas flow between the cap and sensing surface.  Cap-2 reduces the dead volume but still introduces an assymmetric gas flow.  Cap-3 reduces the dead volume and also introduces an symmetric gas flow.  Figure 9.2 shows the flow rate dependence of a prototype sensor T10-90 response time with the three sampling caps.


  Fig. 9.2.  Flow rate dependence of a prototype sensor (S-11) response time with three different sampling caps.

As shown in Figure 3.6, both the reduction of the dead volume and the distribution of a symmetrical gas flow improved the flow rate dependence of the sensor response time. The sensor showed the least flow rate dependence with cap-3.

For the final unit design, the flow rate dependence of the response time is critical because the frequency compensation circuit can only be optimized at one flow rate, but in real application the unit can be used at any flow rate from 100 to 300 ml/min.

9.2. Validation.

The test results shown in Table 3.1 in section 3 demonstrated that the UFO130 unit can be operated at the flow rate as low as 100ml/min and still meet the 130ms response time requirement.

10. Range of Output.

The output range required by SpaceLabs is 3.0 - 3.8 V.  Test results shown in Table 3.1 demonstrated that the UFO130 unit meet this requirement.

11. Power Supply of Unit.

The unit requires a +12V, -12V power supply.  The tolerance is (2V.

12. Weight of Unit.

Typical weight of the unit is 50 ( 10 g.

13. Volume of unit.

Dimension of the unit is 4"(L) x 1.8"(W) x 1"(H).

14. Mounting Kit.  

A mounting bracket was designed for SpaceLab. (DWG# B68831)

15. Sensor Life Time.

15.1. Verification.

Sensor life are limited by both the effective lead amount and effective volume inside the sensor.  The effective volume is defined as the volume change allowed without building significant pressure inside the sensor or the volume change allowed by the back expansion membrane.

15.1.1. Experiment setup.

Four model sensor with not lead inside were assembled together with PT-XX sensors.  After assembly, the sensors passed standard production leaking tests.  The sensors were then place in ambient air at room temperature or in a oven at 40ºC.  For every 24 hours or more, the weight of the sensor was measured with a high sensitive balance (0.1mg).  

Drill a hole through the wall of a standard R24-MED sensor, and attach a syringe through the hole sealed with epoxy.  By plunging or sucking electrolyte with the syringe, effective volume was measured.

15.1.2. Results

    Fig. 15.1.  
The weight change of the four model sensors during the water lost test at 

different temperatures.  

The effective volume was measured to be 0.57 ml.  Figure 15.1 shows the weight change of the four model sensors.  The ambient temperature during the test was about 23ºC, and relative humidity varied from 30% to 50%.

From the data, the water loss rate can be calculated and summarized in Table 15.1.  Dividing the effective volume, sensor life due to water loss is also calculated.

  Table 15.1.  
Water loss rate of the sensor and sensor life due to water loss at different 

temperatures.

	
	Water Loss Rate, mg/day
	Sensor life due to water loss, day

	Sensor No.
	1
	2
	3
	4
	1
	2
	3
	4

	23ºC
	-2.26
	-2.26
	-2.16
	-2.28
	252
	252
	263
	250

	40ºC
	-8.97
	-8.33
	-8.4
	-8.57
	63
	68
	68
	66


Based on Jay Lauer's study, approximately 80% lead in the sensor is effective. The lead amount in the sensor is approximately 10g.   A sensor output about 250 A in air at 23ºC and about 330 A in air at 40ºC.  According to equation (3.1), the life time due to lead consumption at different condition can be calculated and is listed in Table 15.2.


Table 15.2.  Calculated life time due to lead consumption at different conditions.

	Conditions
	Calculated sensor life due to lead consumption, day

	At 23ºC in 21% oxygen
	375

	At 23ºC in 50% oxygen
	160

	At 40ºC in 21% oxygen
	288

	At 40ºC in 50% oxygen
	122


Tables 15.1 and 15.2 show the estimated life of the sensor.  In real application, both water loss and lead consumption can limit the sensor life.  Whichever come first will end the sensor's normal operation. 

V

EFFECTS OF ANESTHETICS AGENTS

on TELEDYNE ANALYTICAL SENSORS

Compliance with ISO 7767 (1997),

Anesthetic Agents (Section 11.1)

HISTORY-


The effect of anesthetic agents on electrochemical sensors has been known for some time.  Standards have been written defining the permissible errors since l979, ANSI A79. 10-1979.  This standard was withdrawn and replaced with ISO 7767 (dated Dec 1988).  ISO 7767 was revised in May of 1997.  It was not until this last standard change, that the testing was modified to change the carrier gas for testing, from 100% oxygen to a mixture of 30% oxygen/70% nitrous oxide.

It was found early on that the errors, due to the presence of anesthetic agents, were significantly less in 30% oxygen than in 100% oxygen (this is the normal gas mixture into which the various anesthetic agents are vaporized).  The T4 data was obtained in 1983, the R15 and C1(R) data in 1985 and the UFO130 data in September of 1997.  All of these data were collected using 100% oxygen as the background gas.  Only the R-24 data (which was taken in mid 1995) was run using the 30% oxygen/70% nitrous oxide mixture (anticipating the release of the revised ISO standard).


The concentration of the various agents used in the tests varied considerably.  This was due to the fact that earlier drafts and standards used concentrations that were less, due to the opinions of individuals on the writing committees at the time.  There was also debate as to whether to use induction or maintenance levels in the testing.

ANALYSIS OF THE DATA (Refer to attached Table)-


For purposes of comparison, since the most current standard uses 30% oxygen/70% nitrous oxide as the background mixture, those data taken in 100% oxygen were factored down by 50% (worst case).  In those few instances where data was taken in both backgrounds (see R-24 data), the factors ranged from 10.5% (in 20% Desflurane) to 50% (in 6% Enflurane).  

The following relationships were observed:

1) The errors are not directly proportional to the oxygen concentration, but the errors are less at 30% oxygen than 100% by the above observed factors.

2) The error vs agent concentration is not linear (e.g. in the case of R-24s in Halothane, the errors at 4 % are greater than those at 6%).

3) All agents produce negative errors, implying the same or similar mechanism.  

4) All exposure vs time charts for all agents show a slow increase in error over the 2 hour test period, followed by a similar or longer recovery period.  (Further evidence of a similar mechanism).

5) All errors in the 30% oxygen/70% nitrous oxide background (actual and worst case calculated) are less than the (2% oxygen allowed by the ISO 7767 (1997) Standard.

EQUIVALENCY-


Data from the above agent tests can be extrapolated to medical sensors that are essentially identical to those actually tested.  The differing feature(s) is specified in the “Same as Except” column of following table:

	Oxygen Sensor
	Equivalent to
	Same as Except

	C1R
	C1
	Has wires attached to contact plate.

	C2R
	C1
	Same as above, plus internal diffusion barrier for improved life in nitrous oxide.

	R13
	T4
	Galvanically driven rather than polarographically, KCl electrolyte rather than KOH. Has identical cathode structure and temperature compensation method.

	R15
	N/A
	Actually tested

	R17MED
	R24
	Proportionally larger (same cathode size and sensing membrane), different connector.

	R22MED
	R24
	Proportionally larger (same cathode size and sensing membrane),

	R23*
	R15
	Dual Cathode (two semi circles) in place of single cathode

	R24MED
	R24
	No difference. Actually tested.

	R29MED*
	N/A
	N/A

	T1
	C1
	Label

	T2
	C1
	Same as C1 except has internal diffusion barrier.

	T4
	N/A
	Actually tested

	T7
	R24
	Proportionally larger (same cathode size and sensing membrane), different connector.

	UFO130
	N/A
	Actually tested



*Purchased Part/Private Labeling


VI

SENSOR LIFE FOR 

THE MEDICAL OXYGEN SENSORS

The medical oxygen sensor is based on an electro-chemical system with a noble metal such as rhodium as the cathode, lead as the anode, and an aqueous solution of potassium hydroxide as electrolyte.  Figure 1 is a schematic showing a cross-sectional view of the sensor.

In the presence of oxygen, the sensor creates a current that flows from cathode to the anode as a result of the following electro-chemical reactions:


Cathode:




Anode:




Water and lead are two consumables inside a sensor. The medical oxygen sensor is designed to have enough water to last. The amount of lead is the limit of sensor life. Therefore, the sensor life can be calculated according to the following equation:

L = 3.59x10-4 WPb/i  *

Here, L is the life time of a sensor in the unit of month; 3.59x10-4 is a constant deduced from Faraday law ( see appendix); WPb is the amount of lead inside the sensor in the unit of grams; i is the current through the sensor in the unit of ampere. The higher the current i, the shorter the sensor life L. When substitute i with the sensor current in air, the equation give the sensor life time in month/air; when substitute in I with the sensor current in 100% oxygen, the equation give the sensor life time in month/100%oxygen.


Appendix

* Faraday law: Q = nFM

Q is the charge in coulomb; n is the number of electron transfer per molecule, for the lead electrode n=2; F is the Faraday constant, 96500; M is the quantity of material in unit of mole; For the lead electrode M = WPb/207; WPb is the weight of lead in grams.

On the other hand, Q = i x L.   i and L is the current and time, respectively.

Therefore, 
L = 2x96500/207 WPb/i = 932 WPb/i (seconds)


 Or  L = 3.59x10-4 WPb/i (months)

NOTE: The above equation assumes one hundred percent usage of the lead electrode. For conservative consideration, 70% anode efficiency should be used to estimate the sensor life.

Table:  Expected Life Calculation assuming 70% lead anode efficiency and 25 deg. C

	Sensor
	Amount of Lead, in grams
	Air Current,

in microamps
	Expected life in Air,  

in months

	C1R
	15
	240
	18

	C2R
	8
	240
	10

	R13 *
	NA
	0.4
	12

	R15
	14
	70
	36

	R17MED
	14
	127
	40

	R22MED
	14
	127
	40

	R23
	NA
	30
	33

	R24MED
	10
	118
	24

	R29MED
	NA
	NA
	24

	T1 
	15
	240
	18

	T2
	8
	240
	10

	T4 **
	NA
	0.4
	NA

	T7
	14
	157
	29

	UFO130
	10
	340
	6


 * Galvanically driven rather than polarographically, KCl electrolyte rather than KOH.

** Polarographically driven rather than Galvanically, KCl electrolyte rather than KOH.

    Life is determined by the amount of water loss from the KCl electrolyte.


 EMBED Word.Picture.8  


Figure 1

x 100





� EMBED Excel.Sheet.8  ���





(2)





% Error  =





x 100





� EMBED Excel.Sheet.8  ���





� EMBED Excel.Sheet.8  ���





� EMBED Excel.Sheet.8  ���





% Error  =





(1)





� EMBED Equation  ���





� EMBED Equation  ���





� EMBED Excel.Sheet.8  ���





� EMBED Equation  ���





� EMBED Excel.Sheet.8  ���





� EMBED Excel.Sheet.8  ���





Input of a step change of oxygen concentration





Signal from ideal sensor





Signal Output





Sensor Low 


Pass Filter





Ideal Sensor





Gas outlet 


delivering to sensor





Gas outlet


delivering  to sensor





Gas-1 inlet 





Gas-1 





Gas-2 inlet 





Gas-2 





Piston





Cylinder





� EMBED Excel.Sheet.8  ���





� EMBED Excel.Sheet.8  ���





� EMBED Excel.Sheet.8  ���





% Error  =





10ºC, 23ºC, 50ºC





� EMBED Excel.Sheet.8  ���





� EMBED Excel.Sheet.8  ���





� EMBED Equation  ���





� EMBED Excel.Sheet.8  ���





Gas out





Gas in





Sampling 


Cap-1





Gas inlet





Sensing 


surface





Sensor body





Gas outlet





Gas out





Gas in





Sampling 


Cap-2





Gas inlet





Sensing 


surface





Gas outlet





Sensor body





Gas in





Gas out





Gas out





Sampling 


Cap-3





Gas inlet





Sensing 


surface





Sensor body





Gas outlets





� EMBED Excel.Sheet.8  ���





� EMBED Excel.Sheet.8  ���





x 100





� EMBED Excel.Sheet.8  ���





(3)








PAGE  
39

_940749222.xls
Chart2

		0		0		0		0		0		0		0		0

		21		21		21		21		21		21		21		21

		40		40		40		40		40		40		40		40

		50		50		50		50		50		50		50		50

		60		60		60		60		60		60		60		60

		80		80		80		80		80		80		80		80

		100		100		100		100		100		100		100		100



Oxygen Concentration, %

Sensor Output, mV

at 26ºC

0.2515

0.2417

0.2417

0.2466

0.2637

0.249

0.2759

0.249

23.2715

24.7974

22.0825

25.52

26.0645

26.3745

21.9995

25.4517

43.7451

46.3989

41.2207

47.6099

48.6304

49.0771

41.0718

47.6294

54.1895

57.5488

51.0303

59.0015

60.2783

60.7959

50.8911

59.0576

64.6631

68.6035

60.8813

70.437

71.875

72.5122

60.7153

70.354

85.4834

90.9009

80.3174

92.9932

94.7754

95.6812

80.3564

92.7588

105.9619

112.5171

99.1406

114.7144

116.7993

117.9517

99.5996

114.4678



Sheet1

		0		0.2515		0.2417		0.2417		0.2466		0.2637		0.249		0.2759		0.249

		21		23.2715		24.7974		22.0825		25.52		26.0645		26.3745		21.9995		25.4517

		40		43.7451		46.3989		41.2207		47.6099		48.6304		49.0771		41.0718		47.6294

		50		54.1895		57.5488		51.0303		59.0015		60.2783		60.7959		50.8911		59.0576

		60		64.6631		68.6035		60.8813		70.437		71.875		72.5122		60.7153		70.354

		80		85.4834		90.9009		80.3174		92.9932		94.7754		95.6812		80.3564		92.7588

		100		105.9619		112.5171		99.1406		114.7144		116.7993		117.9517		99.5996		114.4678

		a=		0.9553708774		0.9005958753		1.0252004656		0.8857058291		0.8706692471		0.8626601381		1.0180399252		0.8874799053

		b=		-1.2329133733		-1.3324361574		-1.6389892821		-1.6032127578		-1.6935585905		-1.7522298127		-1.3963693346		-1.5878723062

						Non-Linearity Error When Calibrated with air and 100% O2

		0		-0.9926375976		-1.1147621344		-1.3911983296		-1.3847977003		-1.4639631101		-1.5374274383		-1.1154921192		-1.3668898098

		21		0		0		0		0		0		0		0		0

		40		0.5598811954		0.4542217997		0.6204915512		0.5651531935		0.6474351627		0.5846280515		0.4163628655		0.6822630962

		50		0.5381567872		0.4957757493		0.6772980388		0.6547597159		0.7889034858		0.693969678		0.4128023031		0.8245609502

		60		0.5443292087		0.4515929717		0.7765478256		0.7832487241		0.8857935434		0.8011546542		0.4142301363		0.8498889527

		80		0.4354374873		0.532539441		0.7024465955		0.7614065457		0.8244675692		0.7881273942		0.4096541113		0.7336987354

		100		0		0		0		0		0		0		0		0





Sheet1

		



Oxygen Concentration, %

Sensor Output, mV

at 26ºC



Sheet2

		





Sheet3

		






_940831947.xls
Chart5

		490

		510

		530

		550

		570

		590

		610

		630



T10-90/ms

Percentage

2.0833333333

18.75

20.8333333333

22.9166666667

20.8333333333

8.3333333333

4.1666666667

2.0833333333



Sheet1

		320		1						320		1		1.4492753623				490		1		2.0833333333		380

		330		2						350		4		5.7971014493				510		9		18.75		370

		330		3						380		10		14.4927536232				530		10		20.8333333333		420

		340		4						410		17		24.6376811594				550		11		22.9166666667		390

		350		5						440		12		17.3913043478				570		10		20.8333333333		400

		360		6						470		17		24.6376811594				590		4		8.3333333333		370

		360		7						500		6		8.6956521739				610		2		4.1666666667		390

		370		8						530		1		1.4492753623				630		1		2.0833333333		410

		370		9						560		1		1.4492753623										330

		370		10								69												360

		370		11																				350

		370		12																				380

		380		13																				330

		380		14																				360

		380		15																				320

		390		16																				340

		390		17																				400

		390		18																				670

		390		19																				390

		390		20																				390

		390		21																				460

		390		22																				380

		400		23																				410

		400		24																				370

		400		25																				470

		400		26																				490

		400		27																				490

		410		28																				490

		410		29																				400

		410		30																				460

		410		31

		410		32

		420		33

		420		34

		420		35

		420		36

		420		37

		420		38

		420		39

		430		40

		430		41

		430		42

		440		43

		440		44

		450		45

		450		46

		450		47

		450		48

		450		49

		450		50

		450		51

		460		52

		460		53

		460		54

		460		55

		460		56

		460		57

		460		58

		460		59

		470		60

		470		61

		480		62

		490		63

		490		64

		490		65

		490		66

		490		67

		510		68

		570		69

		670		70





Sheet1

		



T10-90/ms

Percentage



Sheet2

		



T10-90/ms

Percentage



Sheet3

		





		






_948379409

_948379413.unknown

_1005565422.xls
Sheet1

								Reading Errors Observed in TAI Medical Electrochemical

								Oxygen Sensors in Various Anesthetic Agents

		Sensor		Agent		Conc  (%)		Error In 100% O2        (-% O2)		In 30% O2, 70% N2O         (-%O2)		In 30% O2, 70% N2O (calculated)*		Sens Memb Thickness (mils)

		T-4		Halothane		1.5		3.5				1.75		0.6

				Isoflurane		3		3				1.5		0.6

				Enflurane		3		2.7				1.35		0.6

		R-24MED		Halothane		6		3.4 - 3.6		0.9 - 1				0.5

				Halothane		4				1 - 1.2				0.5

				Halothane		3				0.6 - 0.7				0.5

				Isoflurane		6				0.9 - 1.3				0.5

				Isoflurane		5		2.6 - 3.6		1 - 1.1				0.5

				Enflurane		6		1 - 1.8		0.5 - 0.7				0.5

				Sevoflurane		7				1.1 - 1.3				0.5

				Sevoflurane		6				1.1 - 1.3				0.5

				Sevoflurane		5				1.1 - 1.3				0.5

				Desflurane		20		1.9 - 2.2		0.2 - 0.5				0.5

				Desflurane		15				1.1				0.5

		R-15		Halothane		3		2.3				1.2		0.75

				Isoflurane		1.4		0.7				0.4		0.75

				Enflurane		2.8		1.5				0.8		0.75

		C-1(R)		Halothane		5		2				1.0		1.5

				Isoflurane		5		2				1.0		1.5

				Enflurane		3		1				0.5		1.5

		UFO 130		Halothane		5.42		2.0-2.7				1.4		0.2

				Isoflurane		4.96		1.6 - 2.2				1.1		0.2

				Enflurane		5		1.2 - 1.9				1		0.2

				*Error in 100% O2/2





Sheet2

		





Sheet3

		






_1005539333.xls
Response 

		

		Output and Response Time (air to 100% O2) Testing:

				Gases:		Compressor Air and 100% O2 flows @ 3 scfh

		A.  R17MED Sensor:

						Air, mV		Response						Air, uA		Response

				S/N		@25 deg. C		90%, sec.				S/N		@25 deg. C		90%, sec.

				415210		10.0		3				416183.0		217		21

				415211		9.6		3				416184		222		20

				415212		10.1		3				416185		200		24

				Avg.=		9.9		3				Avg.=		213		22

				Spec. =		10 +/-3 mV		< 6 sec.				Spec. =		200 +/-40 uA		< 30 sec.

		B.  R22MED Sensor:

						Air, mV		Response						Air, mV		Response

				S/N		@25 deg. C		90%, sec.				S/N		@25 deg. C		90%, sec.

				415210		10.0		3				324320		239		19

				415211		9.6		3				324321		230		20

				415212		10.1		3				Avg.=		235		20

				Avg.=		9.9		3				Spec. =		200 +/-40 uA		< 30 sec.

				Spec. =		10 +/-3 mV		< 6 sec.

		C. R13 Sensor:										G. T4 Sensor:

						Reading Air		Response						Reading 100%O2		Response

				S/N		@25 deg. C		90%, sec.				S/N		@25 deg. C		90%, sec.

				423430		100		4				412522		100		4

				423432		100		3				412523		100		3

				423433		100		3				Avg.=		100		4

				Avg.=		100		3				Spec. =		100		< 6 sec.

				Spec. =		100		< 6 sec.

		D.  R15 Sensor:										H. T7 Sensor:

						Air, uA		Response						Reading 100%O2		Response

				S/N		@25 deg. C		90%, sec.				S/N		@25 deg. C		90%, sec.

				393313		51.1		10				171841		100		4

				393314		50.8		9				Spec. =		100		< 6 sec.

				393315		51.8		9

				Avg.=		51.2		9

				Spec. =		59 +/-11 uA		< 15 sec.






_948379410

_940848366.xls
Chart1

		0		0		0		0

		24		24		24		24

		48		48		48		48

		72		72		72		72

		144		144		144		144

		168		168		168		168

		192		192		192		192

		216		216		216		216

		240		240		240		240

		336		336		336		336



40ºC

23ºC

40ºC

Time, hour

Weight Loss, mg

0

0

0

0

-8.4

-8.4

-8

-8.6

-7.9

-7.7

-7.4

-8

-10.1

-9.8

-9.4

-10.1

-16.8

-16.8

-16

-17

-19.2

-19

-18.2

-19.4

-29.2

-29

-27.9

-29.6

-37.5

-36.8

-35.8

-37.3

-46.1

-44

-43.4

-45.1

-82.5

-74.6

-74.9

-78.1



Sheet1

		0		7.6771		7.8642		7.6156		7.9011		0		0		0		0

		24		7.6687		7.8558		7.6076		7.8925		-8.4		-8.4		-8		-8.6

		48		7.6692		7.8565		7.6082		7.8931		-7.9		-7.7		-7.4		-8

		72		7.667		7.8544		7.6062		7.891		-10.1		-9.8		-9.4		-10.1

		144		7.6603		7.8474		7.5996		7.8841		-16.8		-16.8		-16		-17

		168		7.6579		7.8452		7.5974		7.8817		-19.2		-19		-18.2		-19.4

		192		7.6479		7.8352		7.5877		7.8715		-29.2		-29		-27.9		-29.6

		216		7.6396		7.8274		7.5798		7.8638		-37.5		-36.8		-35.8		-37.3

		240		7.631		7.8202		7.5722		7.856		-46.1		-44		-43.4		-45.1

		336		7.5946		7.7896		7.5407		7.823		-82.5		-74.6		-74.9		-78.1





Sheet1

		



40ºC

23ºC

40ºC

Time, hour

Weight Loss, mg



Sheet2

		





Sheet3

		






_946283731.xls
Chart1

		0		0		0

		0.01		0.01		0.01

		0.02		0.02		0.02

		0.03		0.03		0.03

		0.04		0.04		0.04

		0.05		0.05		0.05

		0.0600000001		0.0600000001		0.0600000001

		0.0699999999		0.0699999999		0.0699999999

		0.08		0.08		0.08

		0.09		0.09		0.09

		0.1		0.1		0.1

		0.11		0.11		0.11

		0.12		0.12		0.12

		0.13		0.13		0.13

		0.14		0.14		0.14

		0.15		0.15		0.15

		0.16		0.16		0.16

		0.17		0.17		0.17

		0.1800000001		0.1800000001		0.1800000001

		0.1899999999		0.1899999999		0.1899999999

		0.2		0.2		0.2

		0.21		0.21		0.21

		0.22		0.22		0.22

		0.23		0.23		0.23

		0.24		0.24		0.24

		0.25		0.25		0.25

		0.26		0.26		0.26

		0.27		0.27		0.27

		0.28		0.28		0.28

		0.29		0.29		0.29

		0.3		0.3		0.3

		0.3100000001		0.3100000001		0.3100000001

		0.3199999999		0.3199999999		0.3199999999

		0.33		0.33		0.33

		0.34		0.34		0.34

		0.35		0.35		0.35

		0.36		0.36		0.36

		0.37		0.37		0.37

		0.38		0.38		0.38

		0.39		0.39		0.39

		0.4		0.4		0.4

		0.41		0.41		0.41

		0.42		0.42		0.42

		0.4300000001		0.4300000001		0.4300000001

		0.4399999999		0.4399999999		0.4399999999

		0.45		0.45		0.45

		0.46		0.46		0.46

		0.47		0.47		0.47

		0.48		0.48		0.48

		0.49		0.49		0.49

		0.5		0.5		0.5

		0.51		0.51		0.51

		0.52		0.52		0.52

		0.53		0.53		0.53

		0.54		0.54		0.54

		0.55		0.55		0.55

		0.5600000001		0.5600000001		0.5600000001

		0.5699999999		0.5699999999		0.5699999999

		0.58		0.58		0.58

		0.59		0.59		0.59

		0.6		0.6		0.6

		0.61		0.61		0.61

		0.62		0.62		0.62

		0.63		0.63		0.63

		0.64		0.64		0.64

		0.65		0.65		0.65

		0.66		0.66		0.66

		0.67		0.67		0.67

		0.6800000001		0.6800000001		0.6800000001

		0.6899999999		0.6899999999		0.6899999999

		0.7		0.7		0.7

		0.71		0.71		0.71

		0.72		0.72		0.72

		0.73		0.73		0.73

		0.74		0.74		0.74

		0.75		0.75		0.75

		0.76		0.76		0.76

		0.77		0.77		0.77

		0.78		0.78		0.78

		0.79		0.79		0.79

		0.8		0.8		0.8

		0.8100000001		0.8100000001		0.8100000001

		0.8199999999		0.8199999999		0.8199999999

		0.83		0.83		0.83

		0.84		0.84		0.84

		0.85		0.85		0.85

		0.86		0.86		0.86

		0.87		0.87		0.87

		0.88		0.88		0.88

		0.89		0.89		0.89

		0.9		0.9		0.9

		0.91		0.91		0.91

		0.92		0.92		0.92

		0.9300000001		0.9300000001		0.9300000001

		0.9399999999		0.9399999999		0.9399999999

		0.95		0.95		0.95

		0.96		0.96		0.96

		0.97		0.97		0.97

		0.98		0.98		0.98

		0.99		0.99		0.99

		1		1		1

		1.01		1.01		1.01

		1.02		1.02		1.02

		1.03		1.03		1.03

		1.04		1.04		1.04

		1.05		1.05		1.05

		1.0600000001		1.0600000001		1.0600000001

		1.0699999999		1.0699999999		1.0699999999

		1.08		1.08		1.08

		1.09		1.09		1.09

		1.1		1.1		1.1

		1.11		1.11		1.11

		1.12		1.12		1.12

		1.13		1.13		1.13

		1.14		1.14		1.14

		1.15		1.15		1.15

		1.16		1.16		1.16

		1.17		1.17		1.17

		1.1800000001		1.1800000001		1.1800000001

		1.1899999999		1.1899999999		1.1899999999

		1.2		1.2		1.2

		1.21		1.21		1.21

		1.22		1.22		1.22

		1.23		1.23		1.23

		1.24		1.24		1.24

		1.25		1.25		1.25

		1.26		1.26		1.26

		1.27		1.27		1.27

		1.28		1.28		1.28

		1.29		1.29		1.29

		1.3		1.3		1.3

		1.3100000001		1.3100000001		1.3100000001

		1.3199999999		1.3199999999		1.3199999999

		1.33		1.33		1.33

		1.34		1.34		1.34

		1.35		1.35		1.35

		1.36		1.36		1.36

		1.37		1.37		1.37

		1.38		1.38		1.38

		1.39		1.39		1.39

		1.4		1.4		1.4

		1.41		1.41		1.41

		1.42		1.42		1.42

		1.4300000001		1.4300000001		1.4300000001

		1.4399999999		1.4399999999		1.4399999999

		1.45		1.45		1.45

		1.46		1.46		1.46

		1.47		1.47		1.47

		1.48		1.48		1.48

		1.49		1.49		1.49

		1.5		1.5		1.5

		1.51		1.51		1.51

		1.52		1.52		1.52

		1.53		1.53		1.53

		1.54		1.54		1.54

		1.55		1.55		1.55

		1.5600000001		1.5600000001		1.5600000001

		1.5699999999		1.5699999999		1.5699999999

		1.58		1.58		1.58

		1.59		1.59		1.59

		1.6		1.6		1.6

		1.61		1.61		1.61

		1.62		1.62		1.62

		1.63		1.63		1.63

		1.64		1.64		1.64

		1.65		1.65		1.65

		1.66		1.66		1.66

		1.67		1.67		1.67

		1.6800000001		1.6800000001		1.6800000001

		1.6899999999		1.6899999999		1.6899999999

		1.7		1.7		1.7

		1.71		1.71		1.71

		1.72		1.72		1.72

		1.73		1.73		1.73

		1.74		1.74		1.74

		1.75		1.75		1.75

		1.76		1.76		1.76

		1.77		1.77		1.77

		1.78		1.78		1.78

		1.79		1.79		1.79

		1.8		1.8		1.8

		1.8100000001		1.8100000001		1.8100000001

		1.8199999999		1.8199999999		1.8199999999

		1.83		1.83		1.83

		1.84		1.84		1.84

		1.85		1.85		1.85

		1.86		1.86		1.86

		1.87		1.87		1.87

		1.88		1.88		1.88

		1.89		1.89		1.89

		1.9		1.9		1.9

		1.91		1.91		1.91

		1.92		1.92		1.92

		1.9300000001		1.9300000001		1.9300000001

		1.9399999999		1.9399999999		1.9399999999

		1.95		1.95		1.95

		1.96		1.96		1.96

		1.97		1.97		1.97

		1.98		1.98		1.98

		1.99		1.99		1.99

		2		2		2

		2.01		2.01		2.01

		2.02		2.02		2.02

		2.03		2.03		2.03

		2.04		2.04		2.04

		2.05		2.05		2.05

		2.0600000001		2.0600000001		2.0600000001

		2.0699999999		2.0699999999		2.0699999999

		2.08		2.08		2.08

		2.09		2.09		2.09

		2.1		2.1		2.1

		2.11		2.11		2.11

		2.12		2.12		2.12

		2.13		2.13		2.13

		2.14		2.14		2.14

		2.15		2.15		2.15

		2.16		2.16		2.16

		2.17		2.17		2.17

		2.1800000001		2.1800000001		2.1800000001

		2.1899999999		2.1899999999		2.1899999999

		2.2		2.2		2.2

		2.21		2.21		2.21

		2.22		2.22		2.22

		2.23		2.23		2.23

		2.24		2.24		2.24

		2.25		2.25		2.25

		2.26		2.26		2.26

		2.27		2.27		2.27

		2.28		2.28		2.28

		2.29		2.29		2.29

		2.3		2.3		2.3

		2.3100000001		2.3100000001		2.3100000001

		2.3199999999		2.3199999999		2.3199999999

		2.33		2.33		2.33

		2.34		2.34		2.34

		2.35		2.35		2.35

		2.36		2.36		2.36

		2.37		2.37		2.37

		2.38		2.38		2.38

		2.39		2.39		2.39

		2.4		2.4		2.4

		2.41		2.41		2.41

		2.42		2.42		2.42

		2.4300000001		2.4300000001		2.4300000001

		2.4399999999		2.4399999999		2.4399999999

		2.45		2.45		2.45

		2.46		2.46		2.46

		2.47		2.47		2.47

		2.48		2.48		2.48

		2.49		2.49		2.49

		2.5		2.5		2.5

		2.51		2.51		2.51

		2.52		2.52		2.52

		2.53		2.53		2.53

		2.54		2.54		2.54

		2.55		2.55		2.55

		2.5600000001		2.5600000001		2.5600000001

		2.5699999999		2.5699999999		2.5699999999

		2.58		2.58		2.58

		2.59		2.59		2.59

		2.6		2.6		2.6

		2.61		2.61		2.61

		2.62		2.62		2.62

		2.63		2.63		2.63

		2.64		2.64		2.64

		2.65		2.65		2.65

		2.66		2.66		2.66

		2.67		2.67		2.67

		2.6800000001		2.6800000001		2.6800000001

		2.6899999999		2.6899999999		2.6899999999

		2.7		2.7		2.7

		2.71		2.71		2.71

		2.72		2.72		2.72

		2.73		2.73		2.73

		2.74		2.74		2.74

		2.75		2.75		2.75

		2.76		2.76		2.76

		2.77		2.77		2.77

		2.78		2.78		2.78

		2.79		2.79		2.79

		2.8		2.8		2.8

		2.8100000001		2.8100000001		2.8100000001

		2.8199999999		2.8199999999		2.8199999999

		2.83		2.83		2.83

		2.84		2.84		2.84

		2.85		2.85		2.85

		2.86		2.86		2.86

		2.87		2.87		2.87

		2.88		2.88		2.88

		2.89		2.89		2.89

		2.9		2.9		2.9

		2.91		2.91		2.91

		2.92		2.92		2.92

		2.9300000001		2.9300000001		2.9300000001

		2.9399999999		2.9399999999		2.9399999999

		2.95		2.95		2.95

		2.96		2.96		2.96

		2.97		2.97		2.97

		2.98		2.98		2.98

		2.99		2.99		2.99

		3		3		3

		3.01		3.01		3.01

		3.02		3.02		3.02

		3.03		3.03		3.03

		3.04		3.04		3.04

		3.05		3.05		3.05

		3.0600000001		3.0600000001		3.0600000001

		3.0699999999		3.0699999999		3.0699999999

		3.08		3.08		3.08

		3.09		3.09		3.09

		3.1		3.1		3.1

		3.11		3.11		3.11

		3.12		3.12		3.12

		3.13		3.13		3.13

		3.14		3.14		3.14

		3.15		3.15		3.15

		3.16		3.16		3.16

		3.17		3.17		3.17

		3.1800000001		3.1800000001		3.1800000001

		3.1899999999		3.1899999999		3.1899999999

		3.2		3.2		3.2

		3.21		3.21		3.21

		3.22		3.22		3.22

		3.23		3.23		3.23

		3.24		3.24		3.24

		3.25		3.25		3.25

		3.26		3.26		3.26

		3.27		3.27		3.27

		3.28		3.28		3.28

		3.29		3.29		3.29

		3.3		3.3		3.3

		3.3100000001		3.3100000001		3.3100000001

		3.3199999999		3.3199999999		3.3199999999

		3.33		3.33		3.33

		3.34		3.34		3.34

		3.35		3.35		3.35

		3.36		3.36		3.36

		3.37		3.37		3.37

		3.38		3.38		3.38

		3.39		3.39		3.39

		3.4		3.4		3.4

		3.41		3.41		3.41

		3.42		3.42		3.42

		3.4300000001		3.4300000001		3.4300000001

		3.4399999999		3.4399999999		3.4399999999

		3.45		3.45		3.45

		3.46		3.46		3.46

		3.47		3.47		3.47

		3.48		3.48		3.48

		3.49		3.49		3.49

		3.5		3.5		3.5

		3.51		3.51		3.51

		3.52		3.52		3.52

		3.53		3.53		3.53

		3.54		3.54		3.54

		3.55		3.55		3.55

		3.5600000001		3.5600000001		3.5600000001

		3.5699999999		3.5699999999		3.5699999999

		3.58		3.58		3.58

		3.59		3.59		3.59

		3.6		3.6		3.6

		3.61		3.61		3.61

		3.62		3.62		3.62

		3.63		3.63		3.63

		3.64		3.64		3.64

		3.65		3.65		3.65

		3.66		3.66		3.66

		3.67		3.67		3.67

		3.6800000001		3.6800000001		3.6800000001

		3.6899999999		3.6899999999		3.6899999999

		3.7		3.7		3.7

		3.71		3.71		3.71

		3.72		3.72		3.72

		3.73		3.73		3.73

		3.74		3.74		3.74

		3.75		3.75		3.75

		3.76		3.76		3.76

		3.77		3.77		3.77

		3.78		3.78		3.78

		3.79		3.79		3.79

		3.8		3.8		3.8

		3.8100000001		3.8100000001		3.8100000001

		3.8199999999		3.8199999999		3.8199999999

		3.83		3.83		3.83

		3.84		3.84		3.84

		3.85		3.85		3.85

		3.86		3.86		3.86

		3.87		3.87		3.87

		3.88		3.88		3.88

		3.89		3.89		3.89

		3.9		3.9		3.9

		3.91		3.91		3.91

		3.92		3.92		3.92

		3.9300000001		3.9300000001		3.9300000001

		3.9399999999		3.9399999999		3.9399999999

		3.95		3.95		3.95

		3.96		3.96		3.96

		3.97		3.97		3.97

		3.98		3.98		3.98

		3.99		3.99		3.99

		4		4		4

		4.01		4.01		4.01

		4.02		4.02		4.02

		4.03		4.03		4.03

		4.04		4.04		4.04

		4.05		4.05		4.05

		4.0600000001		4.0600000001		4.0600000001

		4.0699999999		4.0699999999		4.0699999999

		4.08		4.08		4.08

		4.09		4.09		4.09

		4.1		4.1		4.1

		4.11		4.11		4.11

		4.12		4.12		4.12

		4.13		4.13		4.13

		4.14		4.14		4.14

		4.15		4.15		4.15

		4.16		4.16		4.16

		4.17		4.17		4.17

		4.1800000001		4.1800000001		4.1800000001

		4.1899999999		4.1899999999		4.1899999999

		4.2		4.2		4.2

		4.21		4.21		4.21

		4.22		4.22		4.22

		4.23		4.23		4.23

		4.24		4.24		4.24

		4.25		4.25		4.25

		4.26		4.26		4.26

		4.27		4.27		4.27

		4.28		4.28		4.28

		4.29		4.29		4.29

		4.3		4.3		4.3

		4.3100000001		4.3100000001		4.3100000001

		4.3199999999		4.3199999999		4.3199999999

		4.33		4.33		4.33

		4.34		4.34		4.34

		4.35		4.35		4.35

		4.36		4.36		4.36

		4.37		4.37		4.37

		4.38		4.38		4.38

		4.39		4.39		4.39

		4.4		4.4		4.4

		4.41		4.41		4.41

		4.42		4.42		4.42

		4.4300000001		4.4300000001		4.4300000001

		4.4399999999		4.4399999999		4.4399999999

		4.45		4.45		4.45

		4.46		4.46		4.46

		4.47		4.47		4.47

		4.48		4.48		4.48

		4.49		4.49		4.49

		4.5		4.5		4.5

		4.51		4.51		4.51

		4.52		4.52		4.52

		4.53		4.53		4.53

		4.54		4.54		4.54

		4.55		4.55		4.55

		4.5600000001		4.5600000001		4.5600000001

		4.5699999999		4.5699999999		4.5699999999

		4.58		4.58		4.58

		4.59		4.59		4.59

		4.6		4.6		4.6

		4.61		4.61		4.61

		4.62		4.62		4.62

		4.63		4.63		4.63

		4.64		4.64		4.64

		4.65		4.65		4.65

		4.66		4.66		4.66

		4.67		4.67		4.67

		4.6800000001		4.6800000001		4.6800000001

		4.6899999999		4.6899999999		4.6899999999

		4.7		4.7		4.7

		4.71		4.71		4.71

		4.72		4.72		4.72

		4.73		4.73		4.73

		4.74		4.74		4.74

		4.75		4.75		4.75

		4.76		4.76		4.76

		4.77		4.77		4.77

		4.78		4.78		4.78

		4.79		4.79		4.79

		4.8		4.8		4.8

		4.8100000001		4.8100000001		4.8100000001

		4.8199999999		4.8199999999		4.8199999999

		4.83		4.83		4.83

		4.84		4.84		4.84

		4.85		4.85		4.85

		4.86		4.86		4.86

		4.87		4.87		4.87

		4.88		4.88		4.88

		4.89		4.89		4.89

		4.9		4.9		4.9

		4.91		4.91		4.91

		4.92		4.92		4.92

		4.9300000001		4.9300000001		4.9300000001

		4.9399999999		4.9399999999		4.9399999999

		4.95		4.95		4.95

		4.96		4.96		4.96

		4.97		4.97		4.97

		4.98		4.98		4.98

		4.99		4.99		4.99

		5		5		5

		5.01		5.01		5.01

		5.02		5.02		5.02

		5.03		5.03		5.03

		5.04		5.04		5.04

		5.05		5.05		5.05

		5.0600000001		5.0600000001		5.0600000001

		5.0699999999		5.0699999999		5.0699999999

		5.08		5.08		5.08

		5.09		5.09		5.09

		5.1		5.1		5.1

		5.11		5.11		5.11

		5.12		5.12		5.12

		5.13		5.13		5.13

		5.14		5.14		5.14

		5.15		5.15		5.15

		5.16		5.16		5.16

		5.17		5.17		5.17

		5.1800000001		5.1800000001		5.1800000001

		5.1899999999		5.1899999999		5.1899999999

		5.2		5.2		5.2

		5.21		5.21		5.21

		5.22		5.22		5.22

		5.23		5.23		5.23

		5.24		5.24		5.24

		5.25		5.25		5.25

		5.26		5.26		5.26

		5.27		5.27		5.27

		5.28		5.28		5.28

		5.29		5.29		5.29

		5.3		5.3		5.3

		5.3100000001		5.3100000001		5.3100000001

		5.3199999999		5.3199999999		5.3199999999

		5.33		5.33		5.33

		5.34		5.34		5.34

		5.35		5.35		5.35

		5.36		5.36		5.36

		5.37		5.37		5.37

		5.38		5.38		5.38

		5.39		5.39		5.39

		5.4		5.4		5.4

		5.41		5.41		5.41

		5.42		5.42		5.42

		5.4300000001		5.4300000001		5.4300000001

		5.4399999999		5.4399999999		5.4399999999

		5.45		5.45		5.45

		5.46		5.46		5.46

		5.47		5.47		5.47

		5.48		5.48		5.48

		5.49		5.49		5.49

		5.5		5.5		5.5

		5.51		5.51		5.51

		5.52		5.52		5.52

		5.53		5.53		5.53

		5.54		5.54		5.54

		5.55		5.55		5.55

		5.5600000001		5.5600000001		5.5600000001

		5.5699999999		5.5699999999		5.5699999999

		5.58		5.58		5.58

		5.59		5.59		5.59

		5.6		5.6		5.6

		5.61		5.61		5.61

		5.62		5.62		5.62

		5.63		5.63		5.63

		5.64		5.64		5.64

		5.65		5.65		5.65

		5.66		5.66		5.66

		5.67		5.67		5.67

		5.6800000001		5.6800000001		5.6800000001

		5.6899999999		5.6899999999		5.6899999999

		5.7		5.7		5.7

		5.71		5.71		5.71

		5.72		5.72		5.72

		5.73		5.73		5.73

		5.74		5.74		5.74

		5.75		5.75		5.75

		5.76		5.76		5.76

		5.77		5.77		5.77

		5.78		5.78		5.78

		5.79		5.79		5.79

		5.8		5.8		5.8

		5.8100000001		5.8100000001		5.8100000001

		5.8199999999		5.8199999999		5.8199999999

		5.83		5.83		5.83

		5.84		5.84		5.84

		5.85		5.85		5.85

		5.86		5.86		5.86

		5.87		5.87		5.87

		5.88		5.88		5.88

		5.89		5.89		5.89

		5.9		5.9		5.9

		5.91		5.91		5.91

		5.92		5.92		5.92

		5.9300000001		5.9300000001		5.9300000001

		5.9399999999		5.9399999999		5.9399999999

		5.95		5.95		5.95

		5.96		5.96		5.96

		5.97		5.97		5.97

		5.98		5.98		5.98

		5.99		5.99		5.99

		6		6		6

		6.01		6.01		6.01

		6.02		6.02		6.02

		6.03		6.03		6.03

		6.04		6.04		6.04

		6.05		6.05		6.05

		6.0600000001		6.0600000001		6.0600000001

		6.0699999999		6.0699999999		6.0699999999

		6.08		6.08		6.08

		6.09		6.09		6.09

		6.1		6.1		6.1

		6.11		6.11		6.11

		6.12		6.12		6.12

		6.13		6.13		6.13

		6.14		6.14		6.14

		6.15		6.15		6.15

		6.16		6.16		6.16

		6.17		6.17		6.17

		6.1800000001		6.1800000001		6.1800000001

		6.1899999999		6.1899999999		6.1899999999

		6.2		6.2		6.2

		6.21		6.21		6.21

		6.22		6.22		6.22

		6.23		6.23		6.23

		6.24		6.24		6.24

		6.25		6.25		6.25

		6.26		6.26		6.26

		6.27		6.27		6.27

		6.28		6.28		6.28

		6.29		6.29		6.29

		6.3		6.3		6.3

		6.3100000001		6.3100000001		6.3100000001

		6.3199999999		6.3199999999		6.3199999999

		6.33		6.33		6.33

		6.34		6.34		6.34

		6.35		6.35		6.35

		6.36		6.36		6.36

		6.37		6.37		6.37

		6.38		6.38		6.38

		6.39		6.39		6.39

		6.4		6.4		6.4

		6.41		6.41		6.41

		6.42		6.42		6.42

		6.4300000001		6.4300000001		6.4300000001

		6.4399999999		6.4399999999		6.4399999999

		6.45		6.45		6.45

		6.46		6.46		6.46

		6.47		6.47		6.47

		6.48		6.48		6.48

		6.49		6.49		6.49

		6.5		6.5		6.5

		6.51		6.51		6.51

		6.52		6.52		6.52

		6.53		6.53		6.53

		6.54		6.54		6.54

		6.55		6.55		6.55

		6.5600000001		6.5600000001		6.5600000001

		6.5699999999		6.5699999999		6.5699999999

		6.58		6.58		6.58

		6.59		6.59		6.59

		6.6		6.6		6.6

		6.61		6.61		6.61

		6.62		6.62		6.62

		6.63		6.63		6.63

		6.64		6.64		6.64

		6.65		6.65		6.65

		6.66		6.66		6.66

		6.67		6.67		6.67

		6.6800000001		6.6800000001		6.6800000001

		6.6899999999		6.6899999999		6.6899999999

		6.7		6.7		6.7

		6.71		6.71		6.71

		6.72		6.72		6.72

		6.73		6.73		6.73

		6.74		6.74		6.74

		6.75		6.75		6.75

		6.76		6.76		6.76

		6.77		6.77		6.77

		6.78		6.78		6.78

		6.79		6.79		6.79

		6.8		6.8		6.8

		6.8100000001		6.8100000001		6.8100000001

		6.8199999999		6.8199999999		6.8199999999

		6.83		6.83		6.83

		6.84		6.84		6.84

		6.85		6.85		6.85

		6.86		6.86		6.86

		6.87		6.87		6.87

		6.88		6.88		6.88

		6.89		6.89		6.89

		6.9		6.9		6.9

		6.91		6.91		6.91

		6.92		6.92		6.92

		6.9300000001		6.9300000001		6.9300000001

		6.9399999999		6.9399999999		6.9399999999

		6.95		6.95		6.95

		6.96		6.96		6.96

		6.97		6.97		6.97

		6.98		6.98		6.98

		6.99		6.99		6.99

		7		7		7

		7.01		7.01		7.01

		7.02		7.02		7.02

		7.03		7.03		7.03

		7.04		7.04		7.04

		7.05		7.05		7.05

		7.0600000001		7.0600000001		7.0600000001

		7.0699999999		7.0699999999		7.0699999999

		7.08		7.08		7.08

		7.09		7.09		7.09

		7.1		7.1		7.1

		7.11		7.11		7.11

		7.12		7.12		7.12

		7.13		7.13		7.13

		7.14		7.14		7.14

		7.15		7.15		7.15

		7.16		7.16		7.16

		7.17		7.17		7.17

		7.1800000001		7.1800000001		7.1800000001

		7.1899999999		7.1899999999		7.1899999999

		7.2		7.2		7.2

		7.21		7.21		7.21

		7.22		7.22		7.22

		7.23		7.23		7.23

		7.24		7.24		7.24

		7.25		7.25		7.25

		7.26		7.26		7.26

		7.27		7.27		7.27

		7.28		7.28		7.28

		7.29		7.29		7.29

		7.3		7.3		7.3

		7.3100000001		7.3100000001		7.3100000001

		7.3199999999		7.3199999999		7.3199999999

		7.33		7.33		7.33

		7.34		7.34		7.34

		7.35		7.35		7.35

		7.36		7.36		7.36

		7.37		7.37		7.37

		7.38		7.38		7.38

		7.39		7.39		7.39

		7.4		7.4		7.4

		7.41		7.41		7.41

		7.42		7.42		7.42

		7.4300000001		7.4300000001		7.4300000001

		7.4399999999		7.4399999999		7.4399999999

		7.45		7.45		7.45

		7.46		7.46		7.46

		7.47		7.47		7.47

		7.48		7.48		7.48

		7.49		7.49		7.49

		7.5		7.5		7.5

		7.51		7.51		7.51

		7.52		7.52		7.52

		7.53		7.53		7.53

		7.54		7.54		7.54

		7.55		7.55		7.55

		7.5600000001		7.5600000001		7.5600000001

		7.5699999999		7.5699999999		7.5699999999

		7.58		7.58		7.58

		7.59		7.59		7.59

		7.6		7.6		7.6

		7.61		7.61		7.61

		7.62		7.62		7.62

		7.63		7.63		7.63

		7.64		7.64		7.64

		7.65		7.65		7.65

		7.66		7.66		7.66

		7.67		7.67		7.67

		7.6800000001		7.6800000001		7.6800000001

		7.6899999999		7.6899999999		7.6899999999

		7.7		7.7		7.7

		7.71		7.71		7.71

		7.72		7.72		7.72

		7.73		7.73		7.73

		7.74		7.74		7.74

		7.75		7.75		7.75

		7.76		7.76		7.76

		7.77		7.77		7.77

		7.78		7.78		7.78

		7.79		7.79		7.79

		7.8		7.8		7.8

		7.8100000001		7.8100000001		7.8100000001

		7.8199999999		7.8199999999		7.8199999999

		7.83		7.83		7.83

		7.84		7.84		7.84

		7.85		7.85		7.85

		7.86		7.86		7.86

		7.87		7.87		7.87

		7.88		7.88		7.88

		7.89		7.89		7.89

		7.9		7.9		7.9

		7.91		7.91		7.91

		7.92		7.92		7.92

		7.9300000001		7.9300000001		7.9300000001

		7.9399999999		7.9399999999		7.9399999999

		7.95		7.95		7.95

		7.96		7.96		7.96

		7.97		7.97		7.97

		7.98		7.98		7.98

		7.99		7.99		7.99

		8		8		8

		8.01		8.01		8.01

		8.02		8.02		8.02

		8.03		8.03		8.03

		8.04		8.04		8.04

		8.05		8.05		8.05

		8.0600000001		8.0600000001		8.0600000001

		8.0699999999		8.0699999999		8.0699999999

		8.08		8.08		8.08

		8.09		8.09		8.09

		8.1		8.1		8.1

		8.11		8.11		8.11

		8.12		8.12		8.12

		8.13		8.13		8.13

		8.14		8.14		8.14

		8.15		8.15		8.15

		8.16		8.16		8.16

		8.17		8.17		8.17

		8.1800000001		8.1800000001		8.1800000001

		8.1899999999		8.1899999999		8.1899999999

		8.2		8.2		8.2

		8.21		8.21		8.21

		8.22		8.22		8.22

		8.23		8.23		8.23

		8.24		8.24		8.24

		8.25		8.25		8.25

		8.26		8.26		8.26

		8.27		8.27		8.27

		8.28		8.28		8.28

		8.29		8.29		8.29

		8.3		8.3		8.3

		8.3100000001		8.3100000001		8.3100000001

		8.3199999999		8.3199999999		8.3199999999

		8.33		8.33		8.33

		8.34		8.34		8.34

		8.35		8.35		8.35

		8.36		8.36		8.36

		8.37		8.37		8.37

		8.38		8.38		8.38

		8.39		8.39		8.39

		8.4		8.4		8.4

		8.41		8.41		8.41

		8.42		8.42		8.42

		8.4300000001		8.4300000001		8.4300000001

		8.4399999999		8.4399999999		8.4399999999

		8.45		8.45		8.45

		8.46		8.46		8.46

		8.47		8.47		8.47

		8.48		8.48		8.48

		8.49		8.49		8.49

		8.5		8.5		8.5

		8.51		8.51		8.51

		8.52		8.52		8.52

		8.53		8.53		8.53

		8.54		8.54		8.54

		8.55		8.55		8.55

		8.5600000001		8.5600000001		8.5600000001

		8.5699999999		8.5699999999		8.5699999999

		8.58		8.58		8.58

		8.59		8.59		8.59

		8.6		8.6		8.6

		8.61		8.61		8.61

		8.62		8.62		8.62

		8.63		8.63		8.63

		8.64		8.64		8.64

		8.65		8.65		8.65

		8.66		8.66		8.66

		8.67		8.67		8.67

		8.6800000001		8.6800000001		8.6800000001

		8.6899999999		8.6899999999		8.6899999999

		8.7		8.7		8.7

		8.71		8.71		8.71

		8.72		8.72		8.72

		8.73		8.73		8.73

		8.74		8.74		8.74

		8.75		8.75		8.75

		8.76		8.76		8.76

		8.77		8.77		8.77

		8.78		8.78		8.78

		8.79		8.79		8.79

		8.8		8.8		8.8

		8.8100000001		8.8100000001		8.8100000001

		8.8199999999		8.8199999999		8.8199999999

		8.83		8.83		8.83

		8.84		8.84		8.84

		8.85		8.85		8.85

		8.86		8.86		8.86

		8.87		8.87		8.87

		8.88		8.88		8.88

		8.89		8.89		8.89

		8.9		8.9		8.9

		8.91		8.91		8.91

		8.92		8.92		8.92

		8.9300000001		8.9300000001		8.9300000001

		8.9399999999		8.9399999999		8.9399999999

		8.95		8.95		8.95

		8.96		8.96		8.96

		8.97		8.97		8.97

		8.98		8.98		8.98

		8.99		8.99		8.99

		9		9		9

		9.01		9.01		9.01

		9.02		9.02		9.02

		9.03		9.03		9.03

		9.04		9.04		9.04

		9.05		9.05		9.05

		9.0600000001		9.0600000001		9.0600000001

		9.0699999999		9.0699999999		9.0699999999

		9.08		9.08		9.08

		9.09		9.09		9.09

		9.1		9.1		9.1

		9.11		9.11		9.11

		9.12		9.12		9.12

		9.13		9.13		9.13

		9.14		9.14		9.14

		9.15		9.15		9.15

		9.16		9.16		9.16

		9.17		9.17		9.17

		9.1800000001		9.1800000001		9.1800000001

		9.1899999999		9.1899999999		9.1899999999

		9.2		9.2		9.2

		9.21		9.21		9.21

		9.22		9.22		9.22

		9.23		9.23		9.23

		9.24		9.24		9.24

		9.25		9.25		9.25

		9.26		9.26		9.26

		9.27		9.27		9.27

		9.28		9.28		9.28

		9.29		9.29		9.29

		9.3		9.3		9.3

		9.3100000001		9.3100000001		9.3100000001

		9.3199999999		9.3199999999		9.3199999999

		9.33		9.33		9.33

		9.34		9.34		9.34

		9.35		9.35		9.35

		9.36		9.36		9.36

		9.37		9.37		9.37

		9.38		9.38		9.38

		9.39		9.39		9.39

		9.4		9.4		9.4

		9.41		9.41		9.41

		9.42		9.42		9.42

		9.4300000001		9.4300000001		9.4300000001

		9.4399999999		9.4399999999		9.4399999999

		9.45		9.45		9.45

		9.46		9.46		9.46

		9.47		9.47		9.47

		9.48		9.48		9.48

		9.49		9.49		9.49

		9.5		9.5		9.5

		9.51		9.51		9.51

		9.52		9.52		9.52

		9.53		9.53		9.53

		9.54		9.54		9.54

		9.55		9.55		9.55

		9.5600000001		9.5600000001		9.5600000001

		9.5699999999		9.5699999999		9.5699999999

		9.58		9.58		9.58

		9.59		9.59		9.59

		9.6		9.6		9.6

		9.61		9.61		9.61

		9.62		9.62		9.62

		9.63		9.63		9.63

		9.64		9.64		9.64

		9.65		9.65		9.65

		9.66		9.66		9.66

		9.67		9.67		9.67

		9.6800000001		9.6800000001		9.6800000001

		9.6899999999		9.6899999999		9.6899999999

		9.7		9.7		9.7

		9.71		9.71		9.71

		9.72		9.72		9.72

		9.73		9.73		9.73

		9.74		9.74		9.74

		9.75		9.75		9.75

		9.76		9.76		9.76

		9.77		9.77		9.77

		9.78		9.78		9.78

		9.79		9.79		9.79

		9.8		9.8		9.8

		9.8100000001		9.8100000001		9.8100000001

		9.8199999999		9.8199999999		9.8199999999

		9.83		9.83		9.83

		9.84		9.84		9.84

		9.85		9.85		9.85

		9.86		9.86		9.86

		9.87		9.87		9.87

		9.88		9.88		9.88

		9.89		9.89		9.89

		9.9		9.9		9.9

		9.91		9.91		9.91

		9.92		9.92		9.92

		9.9300000001		9.9300000001		9.9300000001

		9.9399999999		9.9399999999		9.9399999999

		9.95		9.95		9.95

		9.96		9.96		9.96

		9.97		9.97		9.97

		9.98		9.98		9.98

		9.99		9.99		9.99

		10		10		10



Time/sec.

Sensor Output, mV

769.04296875

769.04296875

756.8359375

766.6015625

766.6015625

771.484375

751.953125

771.484375

759.27734375

771.484375

769.04296875

773.92578125

756.8359375

769.04296875

773.92578125

769.04296875

751.953125

769.04296875

773.92578125

776.3671875

751.953125

773.92578125

764.16015625

771.484375

756.8359375

766.6015625

759.27734375

771.484375

766.6015625

771.484375

744.62890625

769.04296875

766.6015625

771.484375

749.51171875

771.484375

766.6015625

771.484375

754.39453125

773.92578125

761.71875

773.92578125

766.6015625

771.484375

759.27734375

764.16015625

771.484375

771.484375

754.39453125

771.484375

771.484375

769.04296875

756.8359375

771.484375

793.45703125

766.6015625

769.04296875

793.45703125

754.39453125

771.484375

793.45703125

764.16015625

773.92578125

788.57421875

761.71875

773.92578125

793.45703125

759.27734375

769.04296875

788.57421875

766.6015625

769.04296875

795.8984375

754.39453125

769.04296875

791.015625

769.04296875

771.484375

788.57421875

751.953125

771.484375

793.45703125

766.6015625

769.04296875

778.80859375

754.39453125

769.04296875

795.8984375

764.16015625

771.484375

783.69140625

754.39453125

773.92578125

795.8984375

761.71875

773.92578125

788.57421875

761.71875

771.484375

791.015625

751.953125

771.484375

791.015625

764.16015625

769.04296875

786.1328125

751.953125

771.484375

798.33984375

764.16015625

769.04296875

786.1328125

751.953125

771.484375

798.33984375

764.16015625

769.04296875

781.25

759.27734375

771.484375

793.45703125

764.16015625

776.3671875

786.1328125

764.16015625

773.92578125

795.8984375

756.8359375

771.484375

788.57421875

773.92578125

771.484375

791.015625

751.953125

771.484375

793.45703125

773.92578125

769.04296875

788.57421875

756.8359375

771.484375

793.45703125

771.484375

771.484375

783.69140625

773.92578125

771.484375

795.8984375

793.45703125

769.04296875

788.57421875

839.84375

773.92578125

793.45703125

908.203125

773.92578125

791.015625

1035.15625

773.92578125

791.015625

1186.5234375

769.04296875

793.45703125

1406.25

769.04296875

791.015625

1633.30078125

771.484375

795.8984375

1909.1796875

776.3671875

791.015625

2160.64453125

771.484375

793.45703125

2416.9921875

769.04296875

791.015625

2641.6015625

771.484375

791.015625

2832.03125

771.484375

773.92578125

3005.37109375

769.04296875

786.1328125

3125

771.484375

798.33984375

3247.0703125

769.04296875

786.1328125

3308.10546875

771.484375

800.78125

3386.23046875

771.484375

788.57421875

3420.41015625

771.484375

795.8984375

3464.35546875

776.3671875

786.1328125

3479.00390625

783.69140625

791.015625

3496.09375

805.6640625

791.015625

3513.18359375

864.2578125

788.57421875

3503.41796875

983.88671875

795.8984375

3530.2734375

1191.40625

786.1328125

3508.30078125

1472.16796875

798.33984375

3527.83203125

1806.640625

786.1328125

3510.7421875

2155.76171875

798.33984375

3525.390625

2480.46875

786.1328125

3518.06640625

2761.23046875

793.45703125

3518.06640625

2976.07421875

791.015625

3520.5078125

3134.765625

791.015625

3515.625

3242.1875

793.45703125

3525.390625

3317.87109375

786.1328125

3505.859375

3374.0234375

793.45703125

3525.390625

3405.76171875

788.57421875

3510.7421875

3427.734375

795.8984375

3520.5078125

3444.82421875

786.1328125

3515.625

3457.03125

791.015625

3515.625

3469.23828125

788.57421875

3518.06640625

3476.5625

788.57421875

3508.30078125

3483.88671875

795.8984375

3525.390625

3483.88671875

791.015625

3503.41796875

3496.09375

798.33984375

3518.06640625

3491.2109375

788.57421875

3503.41796875

3500.9765625

815.4296875

3515.625

3500.9765625

878.90625

3505.859375

3503.41796875

1035.15625

3510.7421875

3505.859375

1325.68359375

3510.7421875

3510.7421875

1728.515625

3505.859375

3515.625

2194.82421875

3515.625

3515.625

2595.21484375

3498.53515625

3518.06640625

2912.59765625

3515.625

3520.5078125

3103.02734375

3500.9765625

3522.94921875

3242.1875

3518.06640625

3532.71484375

3315.4296875

3500.9765625

3532.71484375

3371.58203125

3513.18359375

3530.2734375

3410.64453125

3508.30078125

3535.15625

3437.5

3505.859375

3532.71484375

3471.6796875

3515.625

3537.59765625

3481.4453125

3498.53515625

3540.0390625

3510.7421875

3522.94921875

3537.59765625

3510.7421875

3500.9765625

3540.0390625

3537.59765625

3518.06640625

3540.0390625

3540.0390625

3498.53515625

3544.921875

3554.6875

3518.06640625

3547.36328125

3557.12890625

3510.7421875

3544.921875

3566.89453125

3500.9765625

3544.921875

3579.1015625

3510.7421875

3544.921875

3574.21875

3496.09375

3544.921875

3591.30859375

3520.5078125

3549.8046875

3586.42578125

3496.09375

3549.8046875

3598.6328125

3510.7421875

3547.36328125

3598.6328125

3498.53515625

3547.36328125

3603.515625

3508.30078125

3547.36328125

3605.95703125

3508.30078125

3552.24609375

3603.515625

3498.53515625

3554.6875

3615.72265625

3510.7421875

3547.36328125

3603.515625

3491.2109375

3549.8046875

3620.60546875

3518.06640625

3552.24609375

3613.28125

3500.9765625

3554.6875

3615.72265625

3510.7421875

3549.8046875

3615.72265625

3498.53515625

3549.8046875

3613.28125

3505.859375

3549.8046875

3627.9296875

3508.30078125

3549.8046875

3613.28125

3503.41796875

3554.6875

3632.8125

3505.859375

3554.6875

3613.28125

3493.65234375

3547.36328125

3632.8125

3510.7421875

3549.8046875

3618.1640625

3498.53515625

3549.8046875

3625.48828125

3508.30078125

3554.6875

3623.046875

3496.09375

3554.6875

3625.48828125

3505.859375

3549.8046875

3632.8125

3493.65234375

3554.6875

3623.046875

3498.53515625

3552.24609375

3632.8125

3496.09375

3552.24609375

3623.046875

3496.09375

3554.6875

3630.37109375

3505.859375

3552.24609375

3620.60546875

3488.76953125

3552.24609375

3635.25390625

3505.859375

3554.6875

3627.9296875

3483.88671875

3552.24609375

3635.25390625

3515.625

3552.24609375

3632.8125

3488.76953125

3549.8046875

3625.48828125

3503.41796875

3554.6875

3637.6953125

3493.65234375

3549.8046875

3632.8125

3496.09375

3554.6875

3637.6953125

3508.30078125

3552.24609375

3627.9296875

3483.88671875

3554.6875

3635.25390625

3510.7421875

3552.24609375

3640.13671875

3483.88671875

3552.24609375

3635.25390625

3513.18359375

3554.6875

3635.25390625

3488.76953125

3559.5703125

3632.8125

3493.65234375

3557.12890625

3645.01953125

3488.76953125

3554.6875

3632.8125

3491.2109375

3554.6875

3645.01953125

3503.41796875

3549.8046875

3630.37109375

3479.00390625

3557.12890625

3640.13671875

3500.9765625

3557.12890625

3640.13671875

3479.00390625

3554.6875

3637.6953125

3503.41796875

3554.6875

3642.578125

3483.88671875

3554.6875

3635.25390625

3496.09375

3552.24609375

3649.90234375

3488.76953125

3557.12890625

3632.8125

3491.2109375

3559.5703125

3649.90234375

3493.65234375

3554.6875

3635.25390625

3481.4453125

3554.6875

3649.90234375

3505.859375

3554.6875

3642.578125

3486.328125

3557.12890625

3647.4609375

3500.9765625

3559.5703125

3645.01953125

3491.2109375

3554.6875

3640.13671875

3493.65234375

3557.12890625

3652.34375

3493.65234375

3552.24609375

3637.6953125

3493.65234375

3559.5703125

3654.78515625

3491.2109375

3559.5703125

3640.13671875

3491.2109375

3557.12890625

3657.2265625

3496.09375

3557.12890625

3640.13671875

3486.328125

3557.12890625

3649.90234375

3498.53515625

3557.12890625

3649.90234375

3481.4453125

3562.01171875

3649.90234375

3496.09375

3557.12890625

3649.90234375

3483.88671875

3559.5703125

3642.578125

3498.53515625

3552.24609375

3654.78515625

3483.88671875

3562.01171875

3647.4609375

3488.76953125

3562.01171875

3654.78515625

3493.65234375

3562.01171875

3647.4609375

3481.4453125

3559.5703125

3652.34375

3503.41796875

3559.5703125

3657.2265625

3479.00390625

3564.453125

3649.90234375

3503.41796875

3562.01171875

3657.2265625

3483.88671875

3559.5703125

3645.01953125

3503.41796875

3559.5703125

3662.109375

3491.2109375

3564.453125

3652.34375

3488.76953125

3562.01171875

3654.78515625

3500.9765625

3566.89453125

3654.78515625

3483.88671875

3562.01171875

3662.109375

3508.30078125

3562.01171875

3659.66796875

3479.00390625

3564.453125

3649.90234375

3500.9765625

3564.453125

3664.55078125

3481.4453125

3564.453125

3652.34375

3496.09375

3566.89453125

3669.43359375

3491.2109375

3564.453125

3649.90234375

3483.88671875

3566.89453125

3664.55078125

3500.9765625

3564.453125

3657.2265625

3479.00390625

3566.89453125

3664.55078125

3500.9765625

3569.3359375

3659.66796875

3479.00390625

3566.89453125

3654.78515625

3498.53515625

3564.453125

3664.55078125

3488.76953125

3566.89453125

3657.2265625

3491.2109375

3569.3359375

3671.875

3491.2109375

3566.89453125

3649.90234375

3488.76953125

3566.89453125

3669.43359375

3503.41796875

3566.89453125

3659.66796875

3486.328125

3564.453125

3671.875

3498.53515625

3569.3359375

3662.109375

3488.76953125

3569.3359375

3662.109375

3503.41796875

3569.3359375

3666.9921875

3496.09375

3569.3359375

3659.66796875

3498.53515625

3566.89453125

3666.9921875

3493.65234375

3571.77734375

3659.66796875

3500.9765625

3571.77734375

3669.43359375

3493.65234375

3574.21875

3662.109375

3493.65234375

3569.3359375

3664.55078125

3496.09375

3569.3359375

3666.9921875

3486.328125

3571.77734375

3662.109375

3503.41796875

3571.77734375

3674.31640625

3483.88671875

3571.77734375

3659.66796875

3500.9765625

3566.89453125

3671.875

3486.328125

3569.3359375

3659.66796875

3505.859375

3574.21875

3671.875

3488.76953125

3574.21875

3666.9921875

3496.09375

3574.21875

3664.55078125

3500.9765625

3571.77734375

3671.875

3486.328125

3569.3359375

3666.9921875

3513.18359375

3574.21875

3676.7578125

3481.4453125

3576.66015625

3659.66796875

3513.18359375

3574.21875

3676.7578125

3483.88671875

3574.21875

3666.9921875

3510.7421875

3574.21875

3679.19921875

3496.09375

3571.77734375

3662.109375

3493.65234375

3576.66015625

3671.875

3503.41796875

3576.66015625

3671.875

3491.2109375

3571.77734375

3671.875

3515.625

3574.21875

3674.31640625

3479.00390625

3571.77734375

3662.109375

3505.859375

3574.21875

3681.640625

3491.2109375

3581.54296875

3666.9921875

3505.859375

3576.66015625

3676.7578125

3500.9765625

3574.21875

3666.9921875

3488.76953125

3576.66015625

3679.19921875

3505.859375

3576.66015625

3676.7578125

3491.2109375

3581.54296875

3669.43359375

3510.7421875

3583.984375

3676.7578125

3486.328125

3571.77734375

3669.43359375

3503.41796875

3574.21875

3679.19921875

3498.53515625

3579.1015625

3669.43359375

3498.53515625

3579.1015625

3676.7578125

3505.859375

3581.54296875

3674.31640625

3498.53515625

3576.66015625

3676.7578125

3500.9765625

3576.66015625

3679.19921875

3500.9765625

3579.1015625

3666.9921875

3503.41796875

3583.984375

3679.19921875

3496.09375

3581.54296875

3669.43359375

3503.41796875

3579.1015625

3681.640625

3498.53515625

3579.1015625

3671.875

3508.30078125

3576.66015625

3679.19921875

3493.65234375

3583.984375

3681.640625

3503.41796875

3581.54296875

3674.31640625

3496.09375

3579.1015625

3681.640625

3500.9765625

3579.1015625

3669.43359375

3505.859375

3581.54296875

3684.08203125

3493.65234375

3583.984375

3676.7578125

3505.859375

3583.984375

3686.5234375

3491.2109375

3581.54296875

3666.9921875

3515.625

3581.54296875

3684.08203125

3488.76953125

3581.54296875

3676.7578125

3510.7421875

3576.66015625

3679.19921875

3498.53515625

3583.984375

3676.7578125

3500.9765625

3581.54296875

3669.43359375

3508.30078125

3581.54296875

3684.08203125

3493.65234375

3581.54296875

3671.875

3518.06640625

3581.54296875

3681.640625

3493.65234375

3586.42578125

3669.43359375

3515.625

3583.984375

3691.40625

3493.65234375

3576.66015625

3674.31640625

3513.18359375

3581.54296875

3676.7578125

3508.30078125

3588.8671875

3676.7578125

3503.41796875

3588.8671875

3679.19921875

3508.30078125

3586.42578125

3684.08203125

3491.2109375

3583.984375

3676.7578125

3518.06640625

3586.42578125

3684.08203125

3496.09375

3583.984375

3674.31640625

3513.18359375

3586.42578125

3688.96484375

3496.09375

3588.8671875

3674.31640625

3508.30078125

3586.42578125

3681.640625

3508.30078125

3583.984375

3684.08203125

3498.53515625

3586.42578125

3679.19921875

3508.30078125

3588.8671875

3679.19921875

3498.53515625

3586.42578125

3671.875

3520.5078125

3588.8671875

3686.5234375

3498.53515625

3581.54296875

3681.640625

3513.18359375

3583.984375

3684.08203125

3500.9765625

3583.984375

3679.19921875

3510.7421875

3588.8671875

3679.19921875

3510.7421875

3588.8671875

3686.5234375

3508.30078125

3586.42578125

3674.31640625

3510.7421875

3586.42578125

3684.08203125

3503.41796875

3586.42578125

3674.31640625

3520.5078125

3586.42578125

3691.40625

3503.41796875

3591.30859375

3674.31640625

3518.06640625

3586.42578125

3686.5234375

3500.9765625

3586.42578125

3674.31640625

3518.06640625

3586.42578125

3686.5234375

3503.41796875

3586.42578125

3684.08203125

3513.18359375

3593.75

3676.7578125

3505.859375

3586.42578125

3684.08203125

3508.30078125

3586.42578125

3674.31640625

3515.625

3588.8671875

3691.40625

3498.53515625

3593.75

3676.7578125

3520.5078125

3591.30859375

3688.96484375

3498.53515625

3588.8671875

3679.19921875

3522.94921875

3586.42578125

3691.40625

3505.859375

3586.42578125

3684.08203125

3513.18359375

3588.8671875

3684.08203125

3510.7421875

3591.30859375

3684.08203125

3513.18359375

3588.8671875

3679.19921875

3522.94921875

3588.8671875

3688.96484375

3500.9765625

3588.8671875

3674.31640625

3525.390625

3593.75

3691.40625

3503.41796875

3591.30859375

3681.640625

3522.94921875

3588.8671875

3684.08203125

3505.859375

3588.8671875

3686.5234375

3518.06640625

3591.30859375

3679.19921875

3513.18359375

3588.8671875

3691.40625

3510.7421875

3591.30859375

3679.19921875

3520.5078125

3591.30859375

3688.96484375

3500.9765625

3588.8671875

3676.7578125

3518.06640625

3586.42578125

3688.96484375

3505.859375

3593.75

3684.08203125

3520.5078125

3593.75

3684.08203125

3503.41796875

3593.75

3686.5234375

3515.625

3591.30859375

3679.19921875

3515.625

3591.30859375

3691.40625

3518.06640625

3586.42578125

3676.7578125

3515.625

3593.75

3693.84765625

3505.859375

3593.75

3676.7578125

3525.390625

3593.75

3696.2890625

3508.30078125

3588.8671875

3681.640625

3525.390625

3593.75

3688.96484375

3505.859375

3593.75

3686.5234375

3520.5078125

3591.30859375

3681.640625

3513.18359375

3591.30859375

3684.08203125

3520.5078125

3588.8671875

3676.7578125

3515.625

3593.75

3691.40625

3510.7421875

3593.75

3681.640625

3530.2734375

3593.75

3688.96484375

3505.859375

3593.75

3676.7578125

3527.83203125

3588.8671875

3688.96484375

3503.41796875

3591.30859375

3684.08203125

3522.94921875

3591.30859375

3684.08203125

3508.30078125

3593.75

3686.5234375

3518.06640625

3593.75

3679.19921875

3515.625

3591.30859375

3691.40625

3503.41796875

3593.75

3679.19921875

3525.390625

3588.8671875

3688.96484375

3503.41796875

3593.75

3681.640625

3525.390625

3596.19140625

3691.40625

3503.41796875

3593.75

3686.5234375

3522.94921875

3593.75

3681.640625

3513.18359375

3596.19140625

3686.5234375

3518.06640625

3596.19140625

3681.640625

3518.06640625

3596.19140625

3693.84765625

3513.18359375

3591.30859375

3679.19921875

3527.83203125

3588.8671875

3686.5234375

3508.30078125

3588.8671875

3684.08203125

3525.390625

3593.75

3688.96484375

3513.18359375

3593.75

3688.96484375

3522.94921875

3591.30859375

3681.640625

3518.06640625

3593.75

3691.40625

3515.625

3591.30859375

3681.640625

3518.06640625

3596.19140625

3693.84765625

3515.625

3596.19140625

3674.31640625

3522.94921875

3593.75

3691.40625

3513.18359375

3591.30859375

3681.640625

3518.06640625

3593.75

3691.40625

3510.7421875

3588.8671875

3684.08203125

3520.5078125

3596.19140625

3679.19921875

3515.625

3593.75

3691.40625

3520.5078125

3593.75

3681.640625

3515.625

3591.30859375

3691.40625

3522.94921875

3593.75

3679.19921875

3518.06640625

3601.07421875

3696.2890625

3518.06640625

3598.6328125

3684.08203125

3522.94921875

3593.75

3688.96484375

3515.625

3591.30859375

3684.08203125

3535.15625

3593.75

3688.96484375

3508.30078125

3591.30859375

3691.40625

3530.2734375

3591.30859375

3681.640625

3510.7421875

3593.75

3691.40625

3532.71484375

3593.75

3679.19921875

3515.625

3593.75

3691.40625

3515.625

3598.6328125

3684.08203125

3525.390625

3596.19140625

3686.5234375

3505.859375

3591.30859375

3688.96484375

3532.71484375

3591.30859375

3684.08203125

3500.9765625

3591.30859375

3691.40625

3527.83203125

3598.6328125

3676.7578125

3508.30078125

3596.19140625

3691.40625

3525.390625

3596.19140625

3681.640625

3520.5078125

3591.30859375

3693.84765625

3508.30078125

3593.75

3686.5234375

3530.2734375

3598.6328125

3684.08203125

3508.30078125

3596.19140625

3691.40625

3537.59765625

3596.19140625

3686.5234375

3505.859375

3593.75

3691.40625

3530.2734375

3596.19140625

3681.640625

3518.06640625

3591.30859375

3696.2890625

3525.390625

3598.6328125

3679.19921875

3525.390625

3593.75

3693.84765625

3515.625

3593.75

3679.19921875

3530.2734375

3591.30859375

3691.40625

3518.06640625

3596.19140625

3688.96484375

3535.15625

3596.19140625

3686.5234375

3513.18359375

3598.6328125

3688.96484375

3525.390625

3596.19140625

3681.640625

3525.390625

3591.30859375

3696.2890625

3520.5078125

3598.6328125

3679.19921875

3522.94921875

3596.19140625

3693.84765625

3515.625

3598.6328125

3679.19921875

3522.94921875

3598.6328125

3696.2890625

3520.5078125

3593.75

3686.5234375

3522.94921875

3598.6328125

3686.5234375

3515.625

3598.6328125

3688.96484375

3520.5078125

3598.6328125

3684.08203125

3520.5078125

3596.19140625

3693.84765625

3530.2734375

3593.75

3684.08203125

3518.06640625

3598.6328125

3686.5234375

3527.83203125

3598.6328125

3688.96484375

3520.5078125

3598.6328125

3688.96484375

3522.94921875

3596.19140625

3688.96484375

3530.2734375

3593.75

3681.640625

3515.625

3596.19140625

3688.96484375

3535.15625

3601.07421875

3684.08203125

3520.5078125

3596.19140625

3696.2890625

3537.59765625

3596.19140625

3679.19921875

3510.7421875

3598.6328125

3693.84765625

3527.83203125

3596.19140625

3688.96484375

3525.390625

3598.6328125

3688.96484375

3520.5078125

3598.6328125

3688.96484375

3527.83203125

3598.6328125

3684.08203125

3510.7421875

3596.19140625

3696.2890625

3535.15625

3596.19140625

3681.640625

3513.18359375

3598.6328125

3698.73046875

3530.2734375

3598.6328125

3679.19921875

3513.18359375

3598.6328125

3698.73046875

3527.83203125

3593.75

3679.19921875

3525.390625

3601.07421875

3691.40625

3520.5078125

3598.6328125

3686.5234375

3532.71484375

3598.6328125

3688.96484375

3515.625

3596.19140625

3693.84765625

3537.59765625

3596.19140625

3681.640625

3515.625

3593.75

3698.73046875

3532.71484375

3596.19140625

3684.08203125

3520.5078125

3601.07421875

3698.73046875

3532.71484375

3598.6328125

3684.08203125

3530.2734375

3596.19140625

3693.84765625

3525.390625

3598.6328125

3688.96484375

3527.83203125

3598.6328125

3688.96484375

3515.625

3601.07421875

3693.84765625

3537.59765625

3596.19140625

3679.19921875

3518.06640625

3596.19140625

3691.40625

3527.83203125

3591.30859375

3684.08203125

3518.06640625

3596.19140625

3693.84765625

3530.2734375

3596.19140625

3684.08203125

3525.390625

3601.07421875

3688.96484375

3520.5078125

3603.515625

3688.96484375

3525.390625

3598.6328125

3691.40625

3520.5078125

3591.30859375

3693.84765625

3542.48046875

3598.6328125

3684.08203125

3515.625

3601.07421875

3696.2890625

3535.15625

3601.07421875

3686.5234375

3520.5078125

3596.19140625

3693.84765625

3540.0390625

3596.19140625

3684.08203125

3522.94921875

3596.19140625

3691.40625

3532.71484375

3601.07421875

3693.84765625

3527.83203125

3598.6328125

3681.640625

3522.94921875

3596.19140625

3688.96484375

3542.48046875

3596.19140625

3681.640625

3515.625

3596.19140625

3698.73046875

3535.15625

3601.07421875

3681.640625

3515.625

3603.515625

3693.84765625

3537.59765625

3598.6328125

3684.08203125

3520.5078125

3596.19140625

3693.84765625

3525.390625

3598.6328125

3688.96484375

3525.390625

3601.07421875

3686.5234375

3525.390625

3603.515625

3688.96484375

3535.15625

3603.515625

3688.96484375

3515.625

3596.19140625

3696.2890625

3537.59765625

3596.19140625

3679.19921875

3520.5078125

3593.75

3696.2890625

3535.15625

3598.6328125

3688.96484375

3522.94921875

3598.6328125

3698.73046875

3527.83203125

3596.19140625

3684.08203125

3527.83203125

3596.19140625

3688.96484375

3527.83203125

3598.6328125

3693.84765625

3535.15625

3598.6328125

3691.40625

3518.06640625

3598.6328125

3691.40625

3535.15625

3596.19140625

3684.08203125

3522.94921875

3596.19140625

3696.2890625

3537.59765625

3596.19140625

3688.96484375

3522.94921875

3603.515625

3688.96484375

3532.71484375

3598.6328125

3688.96484375

3522.94921875

3598.6328125

3691.40625

3532.71484375

3596.19140625

3693.84765625

3522.94921875

3596.19140625

3686.5234375

3520.5078125

3603.515625

3696.2890625

3535.15625

3603.515625

3679.19921875

3520.5078125

3596.19140625

3696.2890625

3535.15625

3596.19140625

3688.96484375

3515.625

3598.6328125

3688.96484375

3537.59765625

3603.515625

3688.96484375

3518.06640625

3598.6328125

3688.96484375

3537.59765625

3598.6328125

3688.96484375

3525.390625

3596.19140625

3681.640625

3520.5078125

3598.6328125

3696.2890625

3540.0390625

3601.07421875

3684.08203125

3515.625

3601.07421875

3701.171875

3542.48046875

3598.6328125

3681.640625

3513.18359375

3593.75

3693.84765625

3542.48046875

3596.19140625

3688.96484375

3522.94921875

3596.19140625

3688.96484375

3535.15625

3598.6328125

3691.40625

3527.83203125

3601.07421875

3681.640625

3513.18359375

3601.07421875

3696.2890625

3540.0390625

3596.19140625

3684.08203125

3513.18359375

3598.6328125

3693.84765625

3542.48046875

3603.515625

3681.640625

3515.625

3601.07421875

3696.2890625

3535.15625

3596.19140625

3691.40625

3522.94921875

3598.6328125

3691.40625

3522.94921875

3598.6328125

3686.5234375

3527.83203125

3601.07421875

3686.5234375

3520.5078125

3598.6328125

3696.2890625

3540.0390625

3598.6328125

3684.08203125

3518.06640625

3598.6328125

3691.40625

3537.59765625

3598.6328125

3686.5234375

3522.94921875

3603.515625

3693.84765625

3532.71484375

3601.07421875

3686.5234375

3535.15625

3598.6328125

3686.5234375

3527.83203125

3598.6328125

3693.84765625

3532.71484375

3598.6328125

3686.5234375

3532.71484375

3601.07421875

3696.2890625

3532.71484375

3601.07421875

3679.19921875

3527.83203125

3598.6328125

3696.2890625

3530.2734375

3598.6328125

3688.96484375

3525.390625

3598.6328125

3693.84765625

3537.59765625

3598.6328125

3688.96484375

3525.390625

3601.07421875

3684.08203125

3532.71484375

3598.6328125

3696.2890625

3520.5078125

3596.19140625

3684.08203125

3540.0390625

3596.19140625

3696.2890625

3522.94921875

3598.6328125

3679.19921875

3527.83203125

3601.07421875

3698.73046875

3530.2734375

3601.07421875

3684.08203125

3525.390625

3596.19140625

3693.84765625

3542.48046875

3598.6328125

3686.5234375

3515.625

3601.07421875

3691.40625

3542.48046875

3601.07421875

3693.84765625

3518.06640625

3601.07421875

3686.5234375

3540.0390625

3598.6328125

3693.84765625

3525.390625

3601.07421875

3684.08203125

3522.94921875

3598.6328125

3701.171875

3540.0390625

3601.07421875

3684.08203125

3522.94921875

3601.07421875

3693.84765625

3549.8046875

3603.515625

3686.5234375

3513.18359375

3598.6328125

3696.2890625

3537.59765625

3598.6328125

3693.84765625

3520.5078125

3601.07421875

3686.5234375

3537.59765625

3603.515625

3691.40625

3530.2734375

3601.07421875

3686.5234375

3518.06640625

3598.6328125

3693.84765625

3537.59765625

3598.6328125

3686.5234375

3515.625

3598.6328125

3688.96484375

3542.48046875

3601.07421875

3691.40625

3515.625

3601.07421875

3691.40625

3532.71484375

3598.6328125

3696.2890625

3527.83203125

3603.515625

3686.5234375

3535.15625

3598.6328125

3696.2890625

3535.15625

3601.07421875

3684.08203125

3520.5078125

3598.6328125

3696.2890625

3540.0390625

3596.19140625

3684.08203125

3522.94921875

3596.19140625

3691.40625

3542.48046875

3598.6328125

3691.40625

3518.06640625

3601.07421875

3691.40625

3535.15625

3601.07421875

3691.40625

3532.71484375

3598.6328125

3684.08203125

3535.15625

3598.6328125

3701.171875

3532.71484375

3598.6328125

3681.640625

3527.83203125

3596.19140625

3696.2890625

3532.71484375

3601.07421875

3684.08203125

3527.83203125

3601.07421875

3696.2890625

3530.2734375

3601.07421875

3688.96484375

3525.390625

3598.6328125

3688.96484375

3537.59765625

3601.07421875

3693.84765625

3525.390625

3601.07421875

3684.08203125

3535.15625

3601.07421875

3696.2890625

3518.06640625

3601.07421875

3681.640625

3535.15625

3601.07421875

3693.84765625

3527.83203125

3603.515625

3684.08203125

3527.83203125

3598.6328125

3701.171875

3532.71484375

3601.07421875

3686.5234375

3520.5078125

3598.6328125

3693.84765625

3544.921875

3601.07421875

3684.08203125

3522.94921875

3601.07421875

3691.40625

3542.48046875

3603.515625

3688.96484375

3522.94921875

3598.6328125

3684.08203125

3537.59765625

3598.6328125

3693.84765625

3535.15625

3603.515625

3691.40625

3530.2734375

3601.07421875

3693.84765625

3540.0390625

3603.515625

3686.5234375

3518.06640625

3603.515625

3688.96484375

3542.48046875

3601.07421875

3696.2890625

3520.5078125

3601.07421875

3688.96484375

3537.59765625

3601.07421875

3693.84765625

3522.94921875

3601.07421875

3681.640625

3532.71484375

3601.07421875

3698.73046875

3532.71484375

3596.19140625

3681.640625

3520.5078125

3598.6328125

3691.40625

3532.71484375

3601.07421875

3691.40625

3525.390625

3603.515625

3688.96484375

3540.0390625

3601.07421875

3696.2890625

3522.94921875

3601.07421875

3684.08203125

3540.0390625

3598.6328125

3696.2890625

3525.390625

3598.6328125

3684.08203125

3537.59765625

3601.07421875

3701.171875

3530.2734375

3601.07421875

3681.640625

3530.2734375

3601.07421875

3693.84765625

3537.59765625

3603.515625

3686.5234375

3530.2734375

3598.6328125

3696.2890625

3547.36328125

3603.515625

3693.84765625

3525.390625

3598.6328125

3686.5234375

3540.0390625

3596.19140625

3693.84765625

3522.94921875

3601.07421875

3686.5234375

3537.59765625

3598.6328125

3698.73046875

3525.390625

3605.95703125

3681.640625

3532.71484375

3601.07421875

3696.2890625

3527.83203125

3598.6328125

3688.96484375

3525.390625

3598.6328125

3696.2890625

3542.48046875

3598.6328125

3691.40625

3518.06640625

3601.07421875

3688.96484375

3540.0390625

3601.07421875

3693.84765625

3518.06640625

3603.515625

3688.96484375

3544.921875

3601.07421875

3696.2890625

3525.390625

3598.6328125

3679.19921875

3535.15625

3596.19140625

3693.84765625

3532.71484375

3603.515625

3691.40625

3530.2734375

3603.515625

3696.2890625

3544.921875

3603.515625

3691.40625

3520.5078125

3598.6328125

3684.08203125

3544.921875

3603.515625

3696.2890625

3525.390625

3603.515625

3684.08203125

3542.48046875

3603.515625

3696.2890625

3530.2734375

3598.6328125

3684.08203125

3535.15625

3598.6328125

3696.2890625

3535.15625

3601.07421875

3691.40625

3530.2734375

3603.515625

3688.96484375

3537.59765625

3601.07421875

3691.40625

3520.5078125

3603.515625

3691.40625

3537.59765625

3601.07421875

3701.171875

3525.390625

3596.19140625

3679.19921875

3537.59765625

3601.07421875

3696.2890625

3525.390625

3603.515625

3681.640625

3535.15625

3601.07421875

3698.73046875

3530.2734375

3598.6328125

3686.5234375

3535.15625

3601.07421875

3693.84765625

3535.15625

3605.95703125

3691.40625

3527.83203125

3601.07421875

3691.40625

3537.59765625

3603.515625

3693.84765625

3527.83203125

3598.6328125

3681.640625

3537.59765625

3598.6328125

3698.73046875

3525.390625

3598.6328125

3688.96484375

3542.48046875

3601.07421875

3701.171875

3527.83203125

3601.07421875

3686.5234375

3542.48046875

3598.6328125

3693.84765625

3527.83203125

3598.6328125

3691.40625

3527.83203125

3593.75

3693.84765625

3537.59765625

3601.07421875

3691.40625

3520.5078125

3603.515625

3684.08203125

3542.48046875

3601.07421875

3698.73046875

3520.5078125

3598.6328125

3688.96484375

3544.921875

3598.6328125

3693.84765625

3522.94921875

3601.07421875

3684.08203125

3540.0390625

3601.07421875

3691.40625

3525.390625

3598.6328125

3696.2890625

3527.83203125

3598.6328125

3688.96484375

3537.59765625

3598.6328125

3696.2890625

3522.94921875

3603.515625

3684.08203125

3547.36328125

3603.515625

3698.73046875

3515.625

3598.6328125

3686.5234375

3544.921875

3601.07421875

3693.84765625

3527.83203125

3598.6328125

3688.96484375

3542.48046875

3601.07421875

3693.84765625

3535.15625

3601.07421875

3693.84765625

3530.2734375

3601.07421875

3686.5234375

3542.48046875

3598.6328125

3696.2890625

3525.390625

3601.07421875

3684.08203125

3542.48046875

3603.515625

3706.0546875

3520.5078125

3603.515625

3681.640625

3537.59765625

3601.07421875

3696.2890625

3527.83203125

3598.6328125

3686.5234375

3535.15625

3598.6328125

3696.2890625

3532.71484375

3601.07421875

3691.40625

3527.83203125

3601.07421875

3691.40625

3535.15625

3601.07421875

3693.84765625

3530.2734375

3601.07421875

3691.40625

3537.59765625

3601.07421875

3698.73046875

3525.390625

3598.6328125

3681.640625

3537.59765625

3603.515625

3696.2890625

3532.71484375

3603.515625

3686.5234375

3542.48046875

3601.07421875

3696.2890625

3530.2734375

3596.19140625

3686.5234375

3542.48046875

3598.6328125

3688.96484375

3532.71484375

3596.19140625

3696.2890625

3537.59765625

3601.07421875

3688.96484375

3535.15625

3601.07421875

3691.40625

3527.83203125

3596.19140625

3686.5234375

3544.921875

3603.515625

3693.84765625

3522.94921875

3603.515625

3688.96484375

3544.921875

3601.07421875

3693.84765625

3520.5078125

3596.19140625

3688.96484375

3547.36328125

3601.07421875

3691.40625

3530.2734375

3598.6328125

3698.73046875

3537.59765625

3605.95703125

3686.5234375

3532.71484375

3601.07421875

3696.2890625

3520.5078125

3598.6328125

3686.5234375

3549.8046875

3601.07421875

3698.73046875

3515.625

3596.19140625

3684.08203125

3547.36328125

3605.95703125

3688.96484375

3518.06640625

3601.07421875

3691.40625

3544.921875

3598.6328125

3688.96484375

3532.71484375

3603.515625

3696.2890625

3532.71484375

3598.6328125

3681.640625

3537.59765625

3603.515625

3698.73046875

3522.94921875

3601.07421875

3684.08203125

3547.36328125

3596.19140625

3701.171875

3515.625

3596.19140625

3684.08203125

3547.36328125

3598.6328125

3693.84765625

3527.83203125

3598.6328125

3688.96484375

3535.15625

3603.515625

3693.84765625

3532.71484375

3603.515625

3691.40625

3527.83203125

3596.19140625

3684.08203125

3535.15625

3596.19140625

3696.2890625

3522.94921875

3598.6328125

3686.5234375

3540.0390625

3603.515625

3696.2890625

3522.94921875

3601.07421875

3681.640625

3537.59765625

3603.515625

3696.2890625

3525.390625

3598.6328125

3691.40625

3537.59765625

3598.6328125

3691.40625

3537.59765625

3603.515625

3688.96484375

3535.15625

3601.07421875

3688.96484375

3532.71484375

3598.6328125

3698.73046875

3540.0390625

3598.6328125

3688.96484375

3537.59765625

3601.07421875

3696.2890625

3532.71484375

3601.07421875

3684.08203125

3540.0390625

3603.515625

3693.84765625

3530.2734375

3603.515625

3696.2890625

3547.36328125

3596.19140625

3686.5234375

3527.83203125

3596.19140625

3693.84765625

3547.36328125

3603.515625

3688.96484375

3522.94921875

3603.515625

3698.73046875

3540.0390625

3601.07421875

3681.640625

3527.83203125

3598.6328125

3698.73046875

3530.2734375

3601.07421875

3686.5234375

3537.59765625

3603.515625

3696.2890625

3518.06640625

3601.07421875

3688.96484375

3549.8046875

3601.07421875

3686.5234375

3515.625

3601.07421875

3698.73046875

3542.48046875

3598.6328125

3684.08203125

3520.5078125

3601.07421875

3703.61328125

3542.48046875

3601.07421875

3681.640625

3535.15625

3601.07421875

3696.2890625

3525.390625

3598.6328125

3686.5234375

3542.48046875

3598.6328125

3696.2890625

3520.5078125

3603.515625

3691.40625

3552.24609375

3603.515625

3691.40625

3518.06640625

3601.07421875

3693.84765625

3544.921875

3598.6328125

3686.5234375

3530.2734375

3598.6328125

3696.2890625

3540.0390625

3601.07421875

3684.08203125

3540.0390625

3601.07421875

3698.73046875

3527.83203125

3601.07421875

3691.40625

3542.48046875

3601.07421875

3693.84765625

3525.390625

3598.6328125

3688.96484375

3547.36328125

3601.07421875

3688.96484375

3520.5078125

3601.07421875

3698.73046875

3535.15625

3601.07421875

3688.96484375

3527.83203125

3598.6328125

3691.40625

3532.71484375

3601.07421875

3684.08203125

3537.59765625

3598.6328125

3693.84765625

3525.390625

3603.515625

3688.96484375

3535.15625

3603.515625

3688.96484375

3530.2734375

3601.07421875

3691.40625

3540.0390625

3598.6328125

3691.40625

3527.83203125

3598.6328125

3698.73046875

3537.59765625

3603.515625

3684.08203125

3532.71484375

3601.07421875

3696.2890625

3542.48046875

3601.07421875

3684.08203125

3530.2734375

3598.6328125

3698.73046875

3540.0390625

3598.6328125

3684.08203125

3530.2734375

3598.6328125

3691.40625

3537.59765625

3601.07421875

3691.40625

3542.48046875

3598.6328125

3688.96484375

3525.390625

3596.19140625

3693.84765625

3540.0390625

3601.07421875

3681.640625

3520.5078125

3601.07421875

3701.171875

3544.921875

3601.07421875

3684.08203125

3520.5078125

3603.515625

3698.73046875

3537.59765625

3598.6328125

3679.19921875

3530.2734375

3603.515625

3696.2890625

3535.15625

3603.515625

3693.84765625

3540.0390625

3601.07421875

3691.40625

3520.5078125

3603.515625

3691.40625

3544.921875

3601.07421875

3686.5234375

3522.94921875

3598.6328125

3698.73046875

3544.921875

3598.6328125

3688.96484375

3518.06640625

3601.07421875

3696.2890625

3540.0390625

3605.95703125

3684.08203125

3530.2734375

3598.6328125

3693.84765625

3535.15625

3601.07421875

3691.40625

3537.59765625

3603.515625

3688.96484375

3522.94921875

3603.515625

3691.40625

3544.921875

3601.07421875

3693.84765625

3527.83203125

3598.6328125

3698.73046875

3547.36328125

3596.19140625

3684.08203125

3520.5078125

3598.6328125

3693.84765625

3535.15625

3601.07421875

3688.96484375

3530.2734375

3601.07421875

3693.84765625

3532.71484375

3598.6328125

3691.40625

3535.15625

3601.07421875

3686.5234375

3525.390625

3598.6328125

3693.84765625

3537.59765625

3601.07421875

3688.96484375

3522.94921875

3601.07421875

3698.73046875

3544.921875

3598.6328125

3681.640625

3522.94921875

3601.07421875

3698.73046875

3535.15625

3601.07421875

3688.96484375

3530.2734375

3601.07421875

3693.84765625

3535.15625

3601.07421875

3686.5234375

3532.71484375

3603.515625

3686.5234375

3532.71484375

3598.6328125

3701.171875

3537.59765625

3596.19140625

3686.5234375

3532.71484375

3598.6328125

3698.73046875

3549.8046875

3601.07421875

3681.640625

3522.94921875

3603.515625

3701.171875

3542.48046875

3598.6328125

3691.40625

3527.83203125

3601.07421875

3696.2890625

3544.921875

3598.6328125

3691.40625

3527.83203125

3603.515625

3693.84765625

3532.71484375

3601.07421875

3693.84765625

3537.59765625

3598.6328125

3686.5234375

3525.390625

3603.515625

3696.2890625

3544.921875

3598.6328125

3686.5234375

3518.06640625

3601.07421875

3698.73046875

3542.48046875

3601.07421875

3691.40625

3527.83203125

3601.07421875

3693.84765625

3542.48046875

3601.07421875

3688.96484375

3527.83203125

3598.6328125

3691.40625

3535.15625

3603.515625

3696.2890625

3537.59765625

3603.515625

3686.5234375

3530.2734375

3601.07421875

3696.2890625

3544.921875

3603.515625

3688.96484375

3520.5078125

3598.6328125

3701.171875

3544.921875

3603.515625

3688.96484375

3530.2734375

3601.07421875

3688.96484375

3542.48046875

3598.6328125

3696.2890625

3535.15625

3601.07421875

3688.96484375

3532.71484375

3596.19140625

3696.2890625

3540.0390625

3598.6328125

3681.640625

3530.2734375

3601.07421875

3701.171875

3532.71484375

3601.07421875

3688.96484375

3525.390625

3603.515625

3698.73046875

3540.0390625

3601.07421875

3686.5234375

3527.83203125

3601.07421875

3696.2890625

3537.59765625

3603.515625

3691.40625

3525.390625

3603.515625

3688.96484375

3540.0390625

3603.515625

3693.84765625

3522.94921875

3601.07421875

3686.5234375

3537.59765625

3601.07421875

3703.61328125

3535.15625

3598.6328125

3686.5234375

3527.83203125

3603.515625

3698.73046875

3544.921875

3603.515625

3686.5234375

3525.390625

3598.6328125

3696.2890625

3547.36328125

3601.07421875

3688.96484375

3525.390625

3603.515625

3688.96484375

3544.921875

3603.515625

3691.40625

3522.94921875

3598.6328125

3691.40625

3540.0390625

3598.6328125

3696.2890625

3537.59765625

3601.07421875

3688.96484375

3525.390625

3601.07421875

3693.84765625

3544.921875

3598.6328125

3688.96484375

3522.94921875

3603.515625

3696.2890625

3547.36328125

3598.6328125

3688.96484375

3518.06640625

3603.515625

3688.96484375

3544.921875

3605.95703125

3698.73046875

3530.2734375

3603.515625

3688.96484375

3535.15625

3598.6328125

3696.2890625

3535.15625

3598.6328125

3684.08203125

3522.94921875

3598.6328125

3696.2890625

3547.36328125

3596.19140625

3693.84765625

3522.94921875

3603.515625

3693.84765625

3549.8046875

3603.515625

3688.96484375

3525.390625

3598.6328125

3688.96484375

3542.48046875

3601.07421875

3698.73046875

3537.59765625

3598.6328125

3688.96484375

3535.15625

3601.07421875

3696.2890625

3540.0390625

3601.07421875

3684.08203125

3527.83203125

3601.07421875

3701.171875

3544.921875

3598.6328125

3688.96484375

3532.71484375

3601.07421875

3693.84765625

3535.15625

3603.515625

3691.40625

3530.2734375

3603.515625

3691.40625

3537.59765625

3603.515625

3693.84765625

3535.15625

3598.6328125

3686.5234375

3532.71484375

3596.19140625

3693.84765625

3527.83203125

3603.515625

3684.08203125

3535.15625

3605.95703125

3701.171875

3532.71484375

3598.6328125

3684.08203125

3535.15625

3598.6328125

3696.2890625

3535.15625

3596.19140625

3691.40625

3527.83203125

3603.515625

3693.84765625

3540.0390625

3601.07421875

3693.84765625

3527.83203125

3601.07421875

3686.5234375

3544.921875

3601.07421875

3696.2890625

3530.2734375

3601.07421875

3688.96484375

3544.921875

3605.95703125

3698.73046875

3530.2734375

3601.07421875

3686.5234375

3535.15625

3598.6328125

3693.84765625

3544.921875

3598.6328125

3696.2890625

3527.83203125

3598.6328125

3691.40625

3554.6875

3601.07421875

3691.40625

3522.94921875

3603.515625

3684.08203125

3544.921875

3603.515625

3693.84765625

3522.94921875

3601.07421875

3693.84765625

3544.921875

3598.6328125

3696.2890625

3537.59765625

3598.6328125

3688.96484375

3527.83203125

3603.515625

3693.84765625

3540.0390625

3601.07421875

3693.84765625

3520.5078125

3601.07421875

3686.5234375

3549.8046875

3601.07421875

3693.84765625

3518.06640625

3603.515625

3684.08203125

3547.36328125

3601.07421875

3703.61328125

3525.390625

3603.515625

3684.08203125

3540.0390625

3601.07421875

3701.171875

3535.15625

3596.19140625

3679.19921875

3527.83203125

3603.515625

3698.73046875

3544.921875

3601.07421875

3691.40625

3525.390625

3601.07421875

3688.96484375

3549.8046875

3598.6328125

3693.84765625

3527.83203125

3598.6328125

3691.40625

3542.48046875

3598.6328125

3696.2890625

3535.15625

3601.07421875

3681.640625

3532.71484375

3603.515625

3701.171875

3535.15625

3601.07421875

3686.5234375

3527.83203125

3598.6328125

3703.61328125

3532.71484375

3601.07421875

3686.5234375

3527.83203125

3605.95703125

3691.40625

3540.0390625

3601.07421875

3693.84765625

3525.390625

3601.07421875

3693.84765625

3537.59765625

3598.6328125

3693.84765625

3522.94921875

3598.6328125

3684.08203125

3542.48046875

3601.07421875

3696.2890625

3527.83203125

3603.515625

3688.96484375

3537.59765625

3603.515625

3691.40625

3535.15625

3598.6328125

3691.40625

3535.15625

3596.19140625

3688.96484375

3542.48046875

3601.07421875

3696.2890625

3525.390625

3601.07421875

3684.08203125

3544.921875

3603.515625

3696.2890625

3527.83203125

3601.07421875

3679.19921875

3549.8046875

3598.6328125

3698.73046875

3530.2734375

3603.515625

3688.96484375

3537.59765625

3603.515625

3688.96484375

3530.2734375

3601.07421875

3693.84765625

3535.15625

3601.07421875

3691.40625

3540.0390625

3603.515625

3698.73046875



FO134

		0.00		769.043		769.043

		0.01		756.836		766.602

		0.02		766.602		771.484

		0.03		751.953		771.484

		0.04		759.277		771.484

		0.05		769.043		773.926

		0.06		756.836		769.043

		0.07		773.926		769.043

		0.08		751.953		769.043

		0.09		773.926		776.367

		0.10		751.953		773.926

		0.11		764.160		771.484

		0.12		756.836		766.602

		0.13		759.277		771.484

		0.14		766.602		771.484

		0.15		744.629		769.043

		0.16		766.602		771.484

		0.17		749.512		771.484

		0.18		766.602		771.484

		0.19		754.395		773.926

		0.20		761.719		773.926

		0.21		766.602		771.484

		0.22		759.277		764.160

		0.23		771.484		771.484

		0.24		754.395		771.484

		0.25		771.484		769.043

		0.26		756.836		771.484		793.457

		0.27		766.602		769.043		793.457

		0.28		754.395		771.484		793.457

		0.29		764.160		773.926		788.574

		0.30		761.719		773.926		793.457

		0.31		759.277		769.043		788.574

		0.32		766.602		769.043		795.898

		0.33		754.395		769.043		791.016

		0.34		769.043		771.484		788.574

		0.35		751.953		771.484		793.457

		0.36		766.602		769.043		778.809

		0.37		754.395		769.043		795.898

		0.38		764.160		771.484		783.691

		0.39		754.395		773.926		795.898

		0.40		761.719		773.926		788.574

		0.41		761.719		771.484		791.016

		0.42		751.953		771.484		791.016

		0.43		764.160		769.043		786.133

		0.44		751.953		771.484		798.340

		0.45		764.160		769.043		786.133

		0.46		751.953		771.484		798.340

		0.47		764.160		769.043		781.250

		0.48		759.277		771.484		793.457

		0.49		764.160		776.367		786.133

		0.50		764.160		773.926		795.898

		0.51		756.836		771.484		788.574

		0.52		773.926		771.484		791.016

		0.53		751.953		771.484		793.457

		0.54		773.926		769.043		788.574

		0.55		756.836		771.484		793.457

		0.56		771.484		771.484		783.691

		0.57		773.926		771.484		795.898

		0.58		793.457		769.043		788.574

		0.59		839.844		773.926		793.457

		0.60		908.203		773.926		791.016

		0.61		1035.156		773.926		791.016

		0.62		1186.523		769.043		793.457

		0.63		1406.250		769.043		791.016

		0.64		1633.301		771.484		795.898

		0.65		1909.180		776.367		791.016

		0.66		2160.645		771.484		793.457

		0.67		2416.992		769.043		791.016

		0.68		2641.602		771.484		791.016

		0.69		2832.031		771.484		773.926

		0.70		3005.371		769.043		786.133

		0.71		3125.000		771.484		798.340

		0.72		3247.070		769.043		786.133

		0.73		3308.105		771.484		800.781

		0.74		3386.230		771.484		788.574

		0.75		3420.410		771.484		795.898

		0.76		3464.355		776.367		786.133

		0.77		3479.004		783.691		791.016

		0.78		3496.094		805.664		791.016

		0.79		3513.184		864.258		788.574

		0.80		3503.418		983.887		795.898

		0.81		3530.273		1191.406		786.133

		0.82		3508.301		1472.168		798.340

		0.83		3527.832		1806.641		786.133

		0.84		3510.742		2155.762		798.340

		0.85		3525.391		2480.469		786.133

		0.86		3518.066		2761.230		793.457

		0.87		3518.066		2976.074		791.016

		0.88		3520.508		3134.766		791.016

		0.89		3515.625		3242.188		793.457

		0.90		3525.391		3317.871		786.133

		0.91		3505.859		3374.023		793.457

		0.92		3525.391		3405.762		788.574

		0.93		3510.742		3427.734		795.898

		0.94		3520.508		3444.824		786.133

		0.95		3515.625		3457.031		791.016

		0.96		3515.625		3469.238		788.574

		0.97		3518.066		3476.563		788.574

		0.98		3508.301		3483.887		795.898

		0.99		3525.391		3483.887		791.016

		1.00		3503.418		3496.094		798.340

		1.01		3518.066		3491.211		788.574

		1.02		3503.418		3500.977		815.430

		1.03		3515.625		3500.977		878.906

		1.04		3505.859		3503.418		1035.156

		1.05		3510.742		3505.859		1325.684

		1.06		3510.742		3510.742		1728.516

		1.07		3505.859		3515.625		2194.824

		1.08		3515.625		3515.625		2595.215

		1.09		3498.535		3518.066		2912.598

		1.10		3515.625		3520.508		3103.027

		1.11		3500.977		3522.949		3242.188

		1.12		3518.066		3532.715		3315.430

		1.13		3500.977		3532.715		3371.582

		1.14		3513.184		3530.273		3410.645

		1.15		3508.301		3535.156		3437.500

		1.16		3505.859		3532.715		3471.680

		1.17		3515.625		3537.598		3481.445

		1.18		3498.535		3540.039		3510.742

		1.19		3522.949		3537.598		3510.742

		1.20		3500.977		3540.039		3537.598

		1.21		3518.066		3540.039		3540.039

		1.22		3498.535		3544.922		3554.688

		1.23		3518.066		3547.363		3557.129

		1.24		3510.742		3544.922		3566.895

		1.25		3500.977		3544.922		3579.102

		1.26		3510.742		3544.922		3574.219

		1.27		3496.094		3544.922		3591.309

		1.28		3520.508		3549.805		3586.426

		1.29		3496.094		3549.805		3598.633

		1.30		3510.742		3547.363		3598.633

		1.31		3498.535		3547.363		3603.516

		1.32		3508.301		3547.363		3605.957

		1.33		3508.301		3552.246		3603.516

		1.34		3498.535		3554.688		3615.723

		1.35		3510.742		3547.363		3603.516

		1.36		3491.211		3549.805		3620.605

		1.37		3518.066		3552.246		3613.281

		1.38		3500.977		3554.688		3615.723

		1.39		3510.742		3549.805		3615.723

		1.40		3498.535		3549.805		3613.281

		1.41		3505.859		3549.805		3627.930

		1.42		3508.301		3549.805		3613.281

		1.43		3503.418		3554.688		3632.813

		1.44		3505.859		3554.688		3613.281

		1.45		3493.652		3547.363		3632.813

		1.46		3510.742		3549.805		3618.164

		1.47		3498.535		3549.805		3625.488

		1.48		3508.301		3554.688		3623.047

		1.49		3496.094		3554.688		3625.488

		1.50		3505.859		3549.805		3632.813

		1.51		3493.652		3554.688		3623.047

		1.52		3498.535		3552.246		3632.813

		1.53		3496.094		3552.246		3623.047

		1.54		3496.094		3554.688		3630.371

		1.55		3505.859		3552.246		3620.605

		1.56		3488.770		3552.246		3635.254

		1.57		3505.859		3554.688		3627.930

		1.58		3483.887		3552.246		3635.254

		1.59		3515.625		3552.246		3632.813

		1.60		3488.770		3549.805		3625.488

		1.61		3503.418		3554.688		3637.695

		1.62		3493.652		3549.805		3632.813

		1.63		3496.094		3554.688		3637.695

		1.64		3508.301		3552.246		3627.930

		1.65		3483.887		3554.688		3635.254

		1.66		3510.742		3552.246		3640.137

		1.67		3483.887		3552.246		3635.254

		1.68		3513.184		3554.688		3635.254

		1.69		3488.770		3559.570		3632.813

		1.70		3493.652		3557.129		3645.020

		1.71		3488.770		3554.688		3632.813

		1.72		3491.211		3554.688		3645.020

		1.73		3503.418		3549.805		3630.371

		1.74		3479.004		3557.129		3640.137

		1.75		3500.977		3557.129		3640.137

		1.76		3479.004		3554.688		3637.695

		1.77		3503.418		3554.688		3642.578

		1.78		3483.887		3554.688		3635.254

		1.79		3496.094		3552.246		3649.902

		1.80		3488.770		3557.129		3632.813

		1.81		3491.211		3559.570		3649.902

		1.82		3493.652		3554.688		3635.254

		1.83		3481.445		3554.688		3649.902

		1.84		3505.859		3554.688		3642.578

		1.85		3486.328		3557.129		3647.461

		1.86		3500.977		3559.570		3645.020

		1.87		3491.211		3554.688		3640.137

		1.88		3493.652		3557.129		3652.344

		1.89		3493.652		3552.246		3637.695

		1.90		3493.652		3559.570		3654.785

		1.91		3491.211		3559.570		3640.137

		1.92		3491.211		3557.129		3657.227

		1.93		3496.094		3557.129		3640.137

		1.94		3486.328		3557.129		3649.902

		1.95		3498.535		3557.129		3649.902

		1.96		3481.445		3562.012		3649.902

		1.97		3496.094		3557.129		3649.902

		1.98		3483.887		3559.570		3642.578

		1.99		3498.535		3552.246		3654.785

		2.00		3483.887		3562.012		3647.461

		2.01		3488.770		3562.012		3654.785

		2.02		3493.652		3562.012		3647.461

		2.03		3481.445		3559.570		3652.344

		2.04		3503.418		3559.570		3657.227

		2.05		3479.004		3564.453		3649.902

		2.06		3503.418		3562.012		3657.227

		2.07		3483.887		3559.570		3645.020

		2.08		3503.418		3559.570		3662.109

		2.09		3491.211		3564.453		3652.344

		2.10		3488.770		3562.012		3654.785

		2.11		3500.977		3566.895		3654.785

		2.12		3483.887		3562.012		3662.109

		2.13		3508.301		3562.012		3659.668

		2.14		3479.004		3564.453		3649.902

		2.15		3500.977		3564.453		3664.551

		2.16		3481.445		3564.453		3652.344

		2.17		3496.094		3566.895		3669.434

		2.18		3491.211		3564.453		3649.902

		2.19		3483.887		3566.895		3664.551

		2.20		3500.977		3564.453		3657.227

		2.21		3479.004		3566.895		3664.551

		2.22		3500.977		3569.336		3659.668

		2.23		3479.004		3566.895		3654.785

		2.24		3498.535		3564.453		3664.551

		2.25		3488.770		3566.895		3657.227

		2.26		3491.211		3569.336		3671.875

		2.27		3491.211		3566.895		3649.902

		2.28		3488.770		3566.895		3669.434

		2.29		3503.418		3566.895		3659.668

		2.30		3486.328		3564.453		3671.875

		2.31		3498.535		3569.336		3662.109

		2.32		3488.770		3569.336		3662.109

		2.33		3503.418		3569.336		3666.992

		2.34		3496.094		3569.336		3659.668

		2.35		3498.535		3566.895		3666.992

		2.36		3493.652		3571.777		3659.668

		2.37		3500.977		3571.777		3669.434

		2.38		3493.652		3574.219		3662.109

		2.39		3493.652		3569.336		3664.551

		2.40		3496.094		3569.336		3666.992

		2.41		3486.328		3571.777		3662.109

		2.42		3503.418		3571.777		3674.316

		2.43		3483.887		3571.777		3659.668

		2.44		3500.977		3566.895		3671.875

		2.45		3486.328		3569.336		3659.668

		2.46		3505.859		3574.219		3671.875

		2.47		3488.770		3574.219		3666.992

		2.48		3496.094		3574.219		3664.551

		2.49		3500.977		3571.777		3671.875

		2.50		3486.328		3569.336		3666.992

		2.51		3513.184		3574.219		3676.758

		2.52		3481.445		3576.660		3659.668

		2.53		3513.184		3574.219		3676.758

		2.54		3483.887		3574.219		3666.992

		2.55		3510.742		3574.219		3679.199

		2.56		3496.094		3571.777		3662.109

		2.57		3493.652		3576.660		3671.875

		2.58		3503.418		3576.660		3671.875

		2.59		3491.211		3571.777		3671.875

		2.60		3515.625		3574.219		3674.316

		2.61		3479.004		3571.777		3662.109

		2.62		3505.859		3574.219		3681.641

		2.63		3491.211		3581.543		3666.992

		2.64		3505.859		3576.660		3676.758

		2.65		3500.977		3574.219		3666.992

		2.66		3488.770		3576.660		3679.199

		2.67		3505.859		3576.660		3676.758

		2.68		3491.211		3581.543		3669.434

		2.69		3510.742		3583.984		3676.758

		2.70		3486.328		3571.777		3669.434

		2.71		3503.418		3574.219		3679.199

		2.72		3498.535		3579.102		3669.434

		2.73		3498.535		3579.102		3676.758

		2.74		3505.859		3581.543		3674.316

		2.75		3498.535		3576.660		3676.758

		2.76		3500.977		3576.660		3679.199

		2.77		3500.977		3579.102		3666.992

		2.78		3503.418		3583.984		3679.199

		2.79		3496.094		3581.543		3669.434

		2.80		3503.418		3579.102		3681.641

		2.81		3498.535		3579.102		3671.875

		2.82		3508.301		3576.660		3679.199

		2.83		3493.652		3583.984		3681.641

		2.84		3503.418		3581.543		3674.316

		2.85		3496.094		3579.102		3681.641

		2.86		3500.977		3579.102		3669.434

		2.87		3505.859		3581.543		3684.082

		2.88		3493.652		3583.984		3676.758

		2.89		3505.859		3583.984		3686.523

		2.90		3491.211		3581.543		3666.992

		2.91		3515.625		3581.543		3684.082

		2.92		3488.770		3581.543		3676.758

		2.93		3510.742		3576.660		3679.199

		2.94		3498.535		3583.984		3676.758

		2.95		3500.977		3581.543		3669.434

		2.96		3508.301		3581.543		3684.082

		2.97		3493.652		3581.543		3671.875

		2.98		3518.066		3581.543		3681.641

		2.99		3493.652		3586.426		3669.434

		3.00		3515.625		3583.984		3691.406

		3.01		3493.652		3576.660		3674.316

		3.02		3513.184		3581.543		3676.758

		3.03		3508.301		3588.867		3676.758

		3.04		3503.418		3588.867		3679.199

		3.05		3508.301		3586.426		3684.082

		3.06		3491.211		3583.984		3676.758

		3.07		3518.066		3586.426		3684.082

		3.08		3496.094		3583.984		3674.316

		3.09		3513.184		3586.426		3688.965

		3.10		3496.094		3588.867		3674.316

		3.11		3508.301		3586.426		3681.641

		3.12		3508.301		3583.984		3684.082

		3.13		3498.535		3586.426		3679.199

		3.14		3508.301		3588.867		3679.199

		3.15		3498.535		3586.426		3671.875

		3.16		3520.508		3588.867		3686.523

		3.17		3498.535		3581.543		3681.641

		3.18		3513.184		3583.984		3684.082

		3.19		3500.977		3583.984		3679.199

		3.20		3510.742		3588.867		3679.199

		3.21		3510.742		3588.867		3686.523

		3.22		3508.301		3586.426		3674.316

		3.23		3510.742		3586.426		3684.082

		3.24		3503.418		3586.426		3674.316

		3.25		3520.508		3586.426		3691.406

		3.26		3503.418		3591.309		3674.316

		3.27		3518.066		3586.426		3686.523

		3.28		3500.977		3586.426		3674.316

		3.29		3518.066		3586.426		3686.523

		3.30		3503.418		3586.426		3684.082

		3.31		3513.184		3593.750		3676.758

		3.32		3505.859		3586.426		3684.082

		3.33		3508.301		3586.426		3674.316

		3.34		3515.625		3588.867		3691.406

		3.35		3498.535		3593.750		3676.758

		3.36		3520.508		3591.309		3688.965

		3.37		3498.535		3588.867		3679.199

		3.38		3522.949		3586.426		3691.406

		3.39		3505.859		3586.426		3684.082

		3.40		3513.184		3588.867		3684.082

		3.41		3510.742		3591.309		3684.082

		3.42		3513.184		3588.867		3679.199

		3.43		3522.949		3588.867		3688.965

		3.44		3500.977		3588.867		3674.316

		3.45		3525.391		3593.750		3691.406

		3.46		3503.418		3591.309		3681.641

		3.47		3522.949		3588.867		3684.082

		3.48		3505.859		3588.867		3686.523

		3.49		3518.066		3591.309		3679.199

		3.50		3513.184		3588.867		3691.406

		3.51		3510.742		3591.309		3679.199

		3.52		3520.508		3591.309		3688.965

		3.53		3500.977		3588.867		3676.758

		3.54		3518.066		3586.426		3688.965

		3.55		3505.859		3593.750		3684.082

		3.56		3520.508		3593.750		3684.082

		3.57		3503.418		3593.750		3686.523

		3.58		3515.625		3591.309		3679.199

		3.59		3515.625		3591.309		3691.406

		3.60		3518.066		3586.426		3676.758

		3.61		3515.625		3593.750		3693.848

		3.62		3505.859		3593.750		3676.758

		3.63		3525.391		3593.750		3696.289

		3.64		3508.301		3588.867		3681.641

		3.65		3525.391		3593.750		3688.965

		3.66		3505.859		3593.750		3686.523

		3.67		3520.508		3591.309		3681.641

		3.68		3513.184		3591.309		3684.082

		3.69		3520.508		3588.867		3676.758

		3.70		3515.625		3593.750		3691.406

		3.71		3510.742		3593.750		3681.641

		3.72		3530.273		3593.750		3688.965

		3.73		3505.859		3593.750		3676.758

		3.74		3527.832		3588.867		3688.965

		3.75		3503.418		3591.309		3684.082

		3.76		3522.949		3591.309		3684.082

		3.77		3508.301		3593.750		3686.523

		3.78		3518.066		3593.750		3679.199

		3.79		3515.625		3591.309		3691.406

		3.80		3503.418		3593.750		3679.199

		3.81		3525.391		3588.867		3688.965

		3.82		3503.418		3593.750		3681.641

		3.83		3525.391		3596.191		3691.406

		3.84		3503.418		3593.750		3686.523

		3.85		3522.949		3593.750		3681.641

		3.86		3513.184		3596.191		3686.523

		3.87		3518.066		3596.191		3681.641

		3.88		3518.066		3596.191		3693.848

		3.89		3513.184		3591.309		3679.199

		3.90		3527.832		3588.867		3686.523

		3.91		3508.301		3588.867		3684.082

		3.92		3525.391		3593.750		3688.965

		3.93		3513.184		3593.750		3688.965

		3.94		3522.949		3591.309		3681.641

		3.95		3518.066		3593.750		3691.406

		3.96		3515.625		3591.309		3681.641

		3.97		3518.066		3596.191		3693.848

		3.98		3515.625		3596.191		3674.316

		3.99		3522.949		3593.750		3691.406

		4.00		3513.184		3591.309		3681.641

		4.01		3518.066		3593.750		3691.406

		4.02		3510.742		3588.867		3684.082

		4.03		3520.508		3596.191		3679.199

		4.04		3515.625		3593.750		3691.406

		4.05		3520.508		3593.750		3681.641

		4.06		3515.625		3591.309		3691.406

		4.07		3522.949		3593.750		3679.199

		4.08		3518.066		3601.074		3696.289

		4.09		3518.066		3598.633		3684.082

		4.10		3522.949		3593.750		3688.965

		4.11		3515.625		3591.309		3684.082

		4.12		3535.156		3593.750		3688.965

		4.13		3508.301		3591.309		3691.406

		4.14		3530.273		3591.309		3681.641

		4.15		3510.742		3593.750		3691.406

		4.16		3532.715		3593.750		3679.199

		4.17		3515.625		3593.750		3691.406

		4.18		3515.625		3598.633		3684.082

		4.19		3525.391		3596.191		3686.523

		4.20		3505.859		3591.309		3688.965

		4.21		3532.715		3591.309		3684.082

		4.22		3500.977		3591.309		3691.406

		4.23		3527.832		3598.633		3676.758

		4.24		3508.301		3596.191		3691.406

		4.25		3525.391		3596.191		3681.641

		4.26		3520.508		3591.309		3693.848

		4.27		3508.301		3593.750		3686.523

		4.28		3530.273		3598.633		3684.082

		4.29		3508.301		3596.191		3691.406

		4.30		3537.598		3596.191		3686.523

		4.31		3505.859		3593.750		3691.406

		4.32		3530.273		3596.191		3681.641

		4.33		3518.066		3591.309		3696.289

		4.34		3525.391		3598.633		3679.199

		4.35		3525.391		3593.750		3693.848

		4.36		3515.625		3593.750		3679.199

		4.37		3530.273		3591.309		3691.406

		4.38		3518.066		3596.191		3688.965

		4.39		3535.156		3596.191		3686.523

		4.40		3513.184		3598.633		3688.965

		4.41		3525.391		3596.191		3681.641

		4.42		3525.391		3591.309		3696.289

		4.43		3520.508		3598.633		3679.199

		4.44		3522.949		3596.191		3693.848

		4.45		3515.625		3598.633		3679.199

		4.46		3522.949		3598.633		3696.289

		4.47		3520.508		3593.750		3686.523

		4.48		3522.949		3598.633		3686.523

		4.49		3515.625		3598.633		3688.965

		4.50		3520.508		3598.633		3684.082

		4.51		3520.508		3596.191		3693.848

		4.52		3530.273		3593.750		3684.082

		4.53		3518.066		3598.633		3686.523

		4.54		3527.832		3598.633		3688.965

		4.55		3520.508		3598.633		3688.965

		4.56		3522.949		3596.191		3688.965

		4.57		3530.273		3593.750		3681.641

		4.58		3515.625		3596.191		3688.965

		4.59		3535.156		3601.074		3684.082

		4.60		3520.508		3596.191		3696.289

		4.61		3537.598		3596.191		3679.199

		4.62		3510.742		3598.633		3693.848

		4.63		3527.832		3596.191		3688.965

		4.64		3525.391		3598.633		3688.965

		4.65		3520.508		3598.633		3688.965

		4.66		3527.832		3598.633		3684.082

		4.67		3510.742		3596.191		3696.289

		4.68		3535.156		3596.191		3681.641

		4.69		3513.184		3598.633		3698.730

		4.70		3530.273		3598.633		3679.199

		4.71		3513.184		3598.633		3698.730

		4.72		3527.832		3593.750		3679.199

		4.73		3525.391		3601.074		3691.406

		4.74		3520.508		3598.633		3686.523

		4.75		3532.715		3598.633		3688.965

		4.76		3515.625		3596.191		3693.848

		4.77		3537.598		3596.191		3681.641

		4.78		3515.625		3593.750		3698.730

		4.79		3532.715		3596.191		3684.082

		4.80		3520.508		3601.074		3698.730

		4.81		3532.715		3598.633		3684.082

		4.82		3530.273		3596.191		3693.848

		4.83		3525.391		3598.633		3688.965

		4.84		3527.832		3598.633		3688.965

		4.85		3515.625		3601.074		3693.848

		4.86		3537.598		3596.191		3679.199

		4.87		3518.066		3596.191		3691.406

		4.88		3527.832		3591.309		3684.082

		4.89		3518.066		3596.191		3693.848

		4.90		3530.273		3596.191		3684.082

		4.91		3525.391		3601.074		3688.965

		4.92		3520.508		3603.516		3688.965

		4.93		3525.391		3598.633		3691.406

		4.94		3520.508		3591.309		3693.848

		4.95		3542.480		3598.633		3684.082

		4.96		3515.625		3601.074		3696.289

		4.97		3535.156		3601.074		3686.523

		4.98		3520.508		3596.191		3693.848

		4.99		3540.039		3596.191		3684.082

		5.00		3522.949		3596.191		3691.406

		5.01		3532.715		3601.074		3693.848

		5.02		3527.832		3598.633		3681.641

		5.03		3522.949		3596.191		3688.965

		5.04		3542.480		3596.191		3681.641

		5.05		3515.625		3596.191		3698.730

		5.06		3535.156		3601.074		3681.641

		5.07		3515.625		3603.516		3693.848

		5.08		3537.598		3598.633		3684.082

		5.09		3520.508		3596.191		3693.848

		5.10		3525.391		3598.633		3688.965

		5.11		3525.391		3601.074		3686.523

		5.12		3525.391		3603.516		3688.965

		5.13		3535.156		3603.516		3688.965

		5.14		3515.625		3596.191		3696.289

		5.15		3537.598		3596.191		3679.199

		5.16		3520.508		3593.750		3696.289

		5.17		3535.156		3598.633		3688.965

		5.18		3522.949		3598.633		3698.730

		5.19		3527.832		3596.191		3684.082

		5.20		3527.832		3596.191		3688.965

		5.21		3527.832		3598.633		3693.848

		5.22		3535.156		3598.633		3691.406

		5.23		3518.066		3598.633		3691.406

		5.24		3535.156		3596.191		3684.082

		5.25		3522.949		3596.191		3696.289

		5.26		3537.598		3596.191		3688.965

		5.27		3522.949		3603.516		3688.965

		5.28		3532.715		3598.633		3688.965

		5.29		3522.949		3598.633		3691.406

		5.30		3532.715		3596.191		3693.848

		5.31		3522.949		3596.191		3686.523

		5.32		3520.508		3603.516		3696.289

		5.33		3535.156		3603.516		3679.199

		5.34		3520.508		3596.191		3696.289

		5.35		3535.156		3596.191		3688.965

		5.36		3515.625		3598.633		3688.965

		5.37		3537.598		3603.516		3688.965

		5.38		3518.066		3598.633		3688.965

		5.39		3537.598		3598.633		3688.965

		5.40		3525.391		3596.191		3681.641

		5.41		3520.508		3598.633		3696.289

		5.42		3540.039		3601.074		3684.082

		5.43		3515.625		3601.074		3701.172

		5.44		3542.480		3598.633		3681.641

		5.45		3513.184		3593.750		3693.848

		5.46		3542.480		3596.191		3688.965

		5.47		3522.949		3596.191		3688.965

		5.48		3535.156		3598.633		3691.406

		5.49		3527.832		3601.074		3681.641

		5.50		3513.184		3601.074		3696.289

		5.51		3540.039		3596.191		3684.082

		5.52		3513.184		3598.633		3693.848

		5.53		3542.480		3603.516		3681.641

		5.54		3515.625		3601.074		3696.289

		5.55		3535.156		3596.191		3691.406

		5.56		3522.949		3598.633		3691.406

		5.57		3522.949		3598.633		3686.523

		5.58		3527.832		3601.074		3686.523

		5.59		3520.508		3598.633		3696.289

		5.60		3540.039		3598.633		3684.082

		5.61		3518.066		3598.633		3691.406

		5.62		3537.598		3598.633		3686.523

		5.63		3522.949		3603.516		3693.848

		5.64		3532.715		3601.074		3686.523

		5.65		3535.156		3598.633		3686.523

		5.66		3527.832		3598.633		3693.848

		5.67		3532.715		3598.633		3686.523

		5.68		3532.715		3601.074		3696.289

		5.69		3532.715		3601.074		3679.199

		5.70		3527.832		3598.633		3696.289

		5.71		3530.273		3598.633		3688.965

		5.72		3525.391		3598.633		3693.848

		5.73		3537.598		3598.633		3688.965

		5.74		3525.391		3601.074		3684.082

		5.75		3532.715		3598.633		3696.289

		5.76		3520.508		3596.191		3684.082

		5.77		3540.039		3596.191		3696.289

		5.78		3522.949		3598.633		3679.199

		5.79		3527.832		3601.074		3698.730

		5.80		3530.273		3601.074		3684.082

		5.81		3525.391		3596.191		3693.848

		5.82		3542.480		3598.633		3686.523

		5.83		3515.625		3601.074		3691.406

		5.84		3542.480		3601.074		3693.848

		5.85		3518.066		3601.074		3686.523

		5.86		3540.039		3598.633		3693.848

		5.87		3525.391		3601.074		3684.082

		5.88		3522.949		3598.633		3701.172

		5.89		3540.039		3601.074		3684.082

		5.90		3522.949		3601.074		3693.848

		5.91		3549.805		3603.516		3686.523

		5.92		3513.184		3598.633		3696.289

		5.93		3537.598		3598.633		3693.848

		5.94		3520.508		3601.074		3686.523

		5.95		3537.598		3603.516		3691.406

		5.96		3530.273		3601.074		3686.523

		5.97		3518.066		3598.633		3693.848

		5.98		3537.598		3598.633		3686.523

		5.99		3515.625		3598.633		3688.965

		6.00		3542.480		3601.074		3691.406

		6.01		3515.625		3601.074		3691.406

		6.02		3532.715		3598.633		3696.289

		6.03		3527.832		3603.516		3686.523

		6.04		3535.156		3598.633		3696.289

		6.05		3535.156		3601.074		3684.082

		6.06		3520.508		3598.633		3696.289

		6.07		3540.039		3596.191		3684.082

		6.08		3522.949		3596.191		3691.406

		6.09		3542.480		3598.633		3691.406

		6.10		3518.066		3601.074		3691.406

		6.11		3535.156		3601.074		3691.406

		6.12		3532.715		3598.633		3684.082

		6.13		3535.156		3598.633		3701.172

		6.14		3532.715		3598.633		3681.641

		6.15		3527.832		3596.191		3696.289

		6.16		3532.715		3601.074		3684.082

		6.17		3527.832		3601.074		3696.289

		6.18		3530.273		3601.074		3688.965

		6.19		3525.391		3598.633		3688.965

		6.20		3537.598		3601.074		3693.848

		6.21		3525.391		3601.074		3684.082

		6.22		3535.156		3601.074		3696.289

		6.23		3518.066		3601.074		3681.641

		6.24		3535.156		3601.074		3693.848

		6.25		3527.832		3603.516		3684.082

		6.26		3527.832		3598.633		3701.172

		6.27		3532.715		3601.074		3686.523

		6.28		3520.508		3598.633		3693.848

		6.29		3544.922		3601.074		3684.082

		6.30		3522.949		3601.074		3691.406

		6.31		3542.480		3603.516		3688.965

		6.32		3522.949		3598.633		3684.082

		6.33		3537.598		3598.633		3693.848

		6.34		3535.156		3603.516		3691.406

		6.35		3530.273		3601.074		3693.848

		6.36		3540.039		3603.516		3686.523

		6.37		3518.066		3603.516		3688.965

		6.38		3542.480		3601.074		3696.289

		6.39		3520.508		3601.074		3688.965

		6.40		3537.598		3601.074		3693.848

		6.41		3522.949		3601.074		3681.641

		6.42		3532.715		3601.074		3698.730

		6.43		3532.715		3596.191		3681.641

		6.44		3520.508		3598.633		3691.406

		6.45		3532.715		3601.074		3691.406

		6.46		3525.391		3603.516		3688.965

		6.47		3540.039		3601.074		3696.289

		6.48		3522.949		3601.074		3684.082

		6.49		3540.039		3598.633		3696.289

		6.50		3525.391		3598.633		3684.082

		6.51		3537.598		3601.074		3701.172

		6.52		3530.273		3601.074		3681.641

		6.53		3530.273		3601.074		3693.848

		6.54		3537.598		3603.516		3686.523

		6.55		3530.273		3598.633		3696.289

		6.56		3547.363		3603.516		3693.848

		6.57		3525.391		3598.633		3686.523

		6.58		3540.039		3596.191		3693.848

		6.59		3522.949		3601.074		3686.523

		6.60		3537.598		3598.633		3698.730

		6.61		3525.391		3605.957		3681.641

		6.62		3532.715		3601.074		3696.289

		6.63		3527.832		3598.633		3688.965

		6.64		3525.391		3598.633		3696.289

		6.65		3542.480		3598.633		3691.406

		6.66		3518.066		3601.074		3688.965

		6.67		3540.039		3601.074		3693.848

		6.68		3518.066		3603.516		3688.965

		6.69		3544.922		3601.074		3696.289

		6.70		3525.391		3598.633		3679.199

		6.71		3535.156		3596.191		3693.848

		6.72		3532.715		3603.516		3691.406

		6.73		3530.273		3603.516		3696.289

		6.74		3544.922		3603.516		3691.406

		6.75		3520.508		3598.633		3684.082

		6.76		3544.922		3603.516		3696.289

		6.77		3525.391		3603.516		3684.082

		6.78		3542.480		3603.516		3696.289

		6.79		3530.273		3598.633		3684.082

		6.80		3535.156		3598.633		3696.289

		6.81		3535.156		3601.074		3691.406

		6.82		3530.273		3603.516		3688.965

		6.83		3537.598		3601.074		3691.406

		6.84		3520.508		3603.516		3691.406

		6.85		3537.598		3601.074		3701.172

		6.86		3525.391		3596.191		3679.199

		6.87		3537.598		3601.074		3696.289

		6.88		3525.391		3603.516		3681.641

		6.89		3535.156		3601.074		3698.730

		6.90		3530.273		3598.633		3686.523

		6.91		3535.156		3601.074		3693.848

		6.92		3535.156		3605.957		3691.406

		6.93		3527.832		3601.074		3691.406

		6.94		3537.598		3603.516		3693.848

		6.95		3527.832		3598.633		3681.641

		6.96		3537.598		3598.633		3698.730

		6.97		3525.391		3598.633		3688.965

		6.98		3542.480		3601.074		3701.172

		6.99		3527.832		3601.074		3686.523

		7.00		3542.480		3598.633		3693.848

		7.01		3527.832		3598.633		3691.406

		7.02		3527.832		3593.750		3693.848

		7.03		3537.598		3601.074		3691.406

		7.04		3520.508		3603.516		3684.082

		7.05		3542.480		3601.074		3698.730

		7.06		3520.508		3598.633		3688.965

		7.07		3544.922		3598.633		3693.848

		7.08		3522.949		3601.074		3684.082

		7.09		3540.039		3601.074		3691.406

		7.10		3525.391		3598.633		3696.289

		7.11		3527.832		3598.633		3688.965

		7.12		3537.598		3598.633		3696.289

		7.13		3522.949		3603.516		3684.082

		7.14		3547.363		3603.516		3698.730

		7.15		3515.625		3598.633		3686.523

		7.16		3544.922		3601.074		3693.848

		7.17		3527.832		3598.633		3688.965

		7.18		3542.480		3601.074		3693.848

		7.19		3535.156		3601.074		3693.848

		7.20		3530.273		3601.074		3686.523

		7.21		3542.480		3598.633		3696.289

		7.22		3525.391		3601.074		3684.082

		7.23		3542.480		3603.516		3706.055

		7.24		3520.508		3603.516		3681.641

		7.25		3537.598		3601.074		3696.289

		7.26		3527.832		3598.633		3686.523

		7.27		3535.156		3598.633		3696.289

		7.28		3532.715		3601.074		3691.406

		7.29		3527.832		3601.074		3691.406

		7.30		3535.156		3601.074		3693.848

		7.31		3530.273		3601.074		3691.406

		7.32		3537.598		3601.074		3698.730

		7.33		3525.391		3598.633		3681.641

		7.34		3537.598		3603.516		3696.289

		7.35		3532.715		3603.516		3686.523

		7.36		3542.480		3601.074		3696.289

		7.37		3530.273		3596.191		3686.523

		7.38		3542.480		3598.633		3688.965

		7.39		3532.715		3596.191		3696.289

		7.40		3537.598		3601.074		3688.965

		7.41		3535.156		3601.074		3691.406

		7.42		3527.832		3596.191		3686.523

		7.43		3544.922		3603.516		3693.848

		7.44		3522.949		3603.516		3688.965

		7.45		3544.922		3601.074		3693.848

		7.46		3520.508		3596.191		3688.965

		7.47		3547.363		3601.074		3691.406

		7.48		3530.273		3598.633		3698.730

		7.49		3537.598		3605.957		3686.523

		7.50		3532.715		3601.074		3696.289

		7.51		3520.508		3598.633		3686.523

		7.52		3549.805		3601.074		3698.730

		7.53		3515.625		3596.191		3684.082

		7.54		3547.363		3605.957		3688.965

		7.55		3518.066		3601.074		3691.406

		7.56		3544.922		3598.633		3688.965

		7.57		3532.715		3603.516		3696.289

		7.58		3532.715		3598.633		3681.641

		7.59		3537.598		3603.516		3698.730

		7.60		3522.949		3601.074		3684.082

		7.61		3547.363		3596.191		3701.172

		7.62		3515.625		3596.191		3684.082

		7.63		3547.363		3598.633		3693.848

		7.64		3527.832		3598.633		3688.965

		7.65		3535.156		3603.516		3693.848

		7.66		3532.715		3603.516		3691.406

		7.67		3527.832		3596.191		3684.082

		7.68		3535.156		3596.191		3696.289

		7.69		3522.949		3598.633		3686.523

		7.70		3540.039		3603.516		3696.289

		7.71		3522.949		3601.074		3681.641

		7.72		3537.598		3603.516		3696.289

		7.73		3525.391		3598.633		3691.406

		7.74		3537.598		3598.633		3691.406

		7.75		3537.598		3603.516		3688.965

		7.76		3535.156		3601.074		3688.965

		7.77		3532.715		3598.633		3698.730

		7.78		3540.039		3598.633		3688.965

		7.79		3537.598		3601.074		3696.289

		7.80		3532.715		3601.074		3684.082

		7.81		3540.039		3603.516		3693.848

		7.82		3530.273		3603.516		3696.289

		7.83		3547.363		3596.191		3686.523

		7.84		3527.832		3596.191		3693.848

		7.85		3547.363		3603.516		3688.965

		7.86		3522.949		3603.516		3698.730

		7.87		3540.039		3601.074		3681.641

		7.88		3527.832		3598.633		3698.730

		7.89		3530.273		3601.074		3686.523

		7.90		3537.598		3603.516		3696.289

		7.91		3518.066		3601.074		3688.965

		7.92		3549.805		3601.074		3686.523

		7.93		3515.625		3601.074		3698.730

		7.94		3542.480		3598.633		3684.082

		7.95		3520.508		3601.074		3703.613

		7.96		3542.480		3601.074		3681.641

		7.97		3535.156		3601.074		3696.289

		7.98		3525.391		3598.633		3686.523

		7.99		3542.480		3598.633		3696.289

		8.00		3520.508		3603.516		3691.406

		8.01		3552.246		3603.516		3691.406

		8.02		3518.066		3601.074		3693.848

		8.03		3544.922		3598.633		3686.523

		8.04		3530.273		3598.633		3696.289

		8.05		3540.039		3601.074		3684.082

		8.06		3540.039		3601.074		3698.730

		8.07		3527.832		3601.074		3691.406

		8.08		3542.480		3601.074		3693.848

		8.09		3525.391		3598.633		3688.965

		8.10		3547.363		3601.074		3688.965

		8.11		3520.508		3601.074		3698.730

		8.12		3535.156		3601.074		3688.965

		8.13		3527.832		3598.633		3691.406

		8.14		3532.715		3601.074		3684.082

		8.15		3537.598		3598.633		3693.848

		8.16		3525.391		3603.516		3688.965

		8.17		3535.156		3603.516		3688.965

		8.18		3530.273		3601.074		3691.406

		8.19		3540.039		3598.633		3691.406

		8.20		3527.832		3598.633		3698.730

		8.21		3537.598		3603.516		3684.082

		8.22		3532.715		3601.074		3696.289

		8.23		3542.480		3601.074		3684.082

		8.24		3530.273		3598.633		3698.730

		8.25		3540.039		3598.633		3684.082

		8.26		3530.273		3598.633		3691.406

		8.27		3537.598		3601.074		3691.406

		8.28		3542.480		3598.633		3688.965

		8.29		3525.391		3596.191		3693.848

		8.30		3540.039		3601.074		3681.641

		8.31		3520.508		3601.074		3701.172

		8.32		3544.922		3601.074		3684.082

		8.33		3520.508		3603.516		3698.730

		8.34		3537.598		3598.633		3679.199

		8.35		3530.273		3603.516		3696.289

		8.36		3535.156		3603.516		3693.848

		8.37		3540.039		3601.074		3691.406

		8.38		3520.508		3603.516		3691.406

		8.39		3544.922		3601.074		3686.523

		8.40		3522.949		3598.633		3698.730

		8.41		3544.922		3598.633		3688.965

		8.42		3518.066		3601.074		3696.289

		8.43		3540.039		3605.957		3684.082

		8.44		3530.273		3598.633		3693.848

		8.45		3535.156		3601.074		3691.406

		8.46		3537.598		3603.516		3688.965

		8.47		3522.949		3603.516		3691.406

		8.48		3544.922		3601.074		3693.848

		8.49		3527.832		3598.633		3698.730

		8.50		3547.363		3596.191		3684.082

		8.51		3520.508		3598.633		3693.848

		8.52		3535.156		3601.074		3688.965

		8.53		3530.273		3601.074		3693.848

		8.54		3532.715		3598.633		3691.406

		8.55		3535.156		3601.074		3686.523

		8.56		3525.391		3598.633		3693.848

		8.57		3537.598		3601.074		3688.965

		8.58		3522.949		3601.074		3698.730

		8.59		3544.922		3598.633		3681.641

		8.60		3522.949		3601.074		3698.730

		8.61		3535.156		3601.074		3688.965

		8.62		3530.273		3601.074		3693.848

		8.63		3535.156		3601.074		3686.523

		8.64		3532.715		3603.516		3686.523

		8.65		3532.715		3598.633		3701.172

		8.66		3537.598		3596.191		3686.523

		8.67		3532.715		3598.633		3698.730

		8.68		3549.805		3601.074		3681.641

		8.69		3522.949		3603.516		3701.172

		8.70		3542.480		3598.633		3691.406

		8.71		3527.832		3601.074		3696.289

		8.72		3544.922		3598.633		3691.406

		8.73		3527.832		3603.516		3693.848

		8.74		3532.715		3601.074		3693.848

		8.75		3537.598		3598.633		3686.523

		8.76		3525.391		3603.516		3696.289

		8.77		3544.922		3598.633		3686.523

		8.78		3518.066		3601.074		3698.730

		8.79		3542.480		3601.074		3691.406

		8.80		3527.832		3601.074		3693.848

		8.81		3542.480		3601.074		3688.965

		8.82		3527.832		3598.633		3691.406

		8.83		3535.156		3603.516		3696.289

		8.84		3537.598		3603.516		3686.523

		8.85		3530.273		3601.074		3696.289

		8.86		3544.922		3603.516		3688.965

		8.87		3520.508		3598.633		3701.172

		8.88		3544.922		3603.516		3688.965

		8.89		3530.273		3601.074		3688.965

		8.90		3542.480		3598.633		3696.289

		8.91		3535.156		3601.074		3688.965

		8.92		3532.715		3596.191		3696.289

		8.93		3540.039		3598.633		3681.641

		8.94		3530.273		3601.074		3701.172

		8.95		3532.715		3601.074		3688.965

		8.96		3525.391		3603.516		3698.730

		8.97		3540.039		3601.074		3686.523

		8.98		3527.832		3601.074		3696.289

		8.99		3537.598		3603.516		3691.406

		9.00		3525.391		3603.516		3688.965

		9.01		3540.039		3603.516		3693.848

		9.02		3522.949		3601.074		3686.523

		9.03		3537.598		3601.074		3703.613

		9.04		3535.156		3598.633		3686.523

		9.05		3527.832		3603.516		3698.730

		9.06		3544.922		3603.516		3686.523

		9.07		3525.391		3598.633		3696.289

		9.08		3547.363		3601.074		3688.965

		9.09		3525.391		3603.516		3688.965

		9.10		3544.922		3603.516		3691.406

		9.11		3522.949		3598.633		3691.406

		9.12		3540.039		3598.633		3696.289

		9.13		3537.598		3601.074		3688.965

		9.14		3525.391		3601.074		3693.848

		9.15		3544.922		3598.633		3688.965

		9.16		3522.949		3603.516		3696.289

		9.17		3547.363		3598.633		3688.965

		9.18		3518.066		3603.516		3688.965

		9.19		3544.922		3605.957		3698.730

		9.20		3530.273		3603.516		3688.965

		9.21		3535.156		3598.633		3696.289

		9.22		3535.156		3598.633		3684.082

		9.23		3522.949		3598.633		3696.289

		9.24		3547.363		3596.191		3693.848

		9.25		3522.949		3603.516		3693.848

		9.26		3549.805		3603.516		3688.965

		9.27		3525.391		3598.633		3688.965

		9.28		3542.480		3601.074		3698.730

		9.29		3537.598		3598.633		3688.965

		9.30		3535.156		3601.074		3696.289

		9.31		3540.039		3601.074		3684.082

		9.32		3527.832		3601.074		3701.172

		9.33		3544.922		3598.633		3688.965

		9.34		3532.715		3601.074		3693.848

		9.35		3535.156		3603.516		3691.406

		9.36		3530.273		3603.516		3691.406

		9.37		3537.598		3603.516		3693.848

		9.38		3535.156		3598.633		3686.523

		9.39		3532.715		3596.191		3693.848

		9.40		3527.832		3603.516		3684.082

		9.41		3535.156		3605.957		3701.172

		9.42		3532.715		3598.633		3684.082

		9.43		3535.156		3598.633		3696.289

		9.44		3535.156		3596.191		3691.406

		9.45		3527.832		3603.516		3693.848

		9.46		3540.039		3601.074		3693.848

		9.47		3527.832		3601.074		3686.523

		9.48		3544.922		3601.074		3696.289

		9.49		3530.273		3601.074		3688.965

		9.50		3544.922		3605.957		3698.730

		9.51		3530.273		3601.074		3686.523

		9.52		3535.156		3598.633		3693.848

		9.53		3544.922		3598.633		3696.289

		9.54		3527.832		3598.633		3691.406

		9.55		3554.688		3601.074		3691.406

		9.56		3522.949		3603.516		3684.082

		9.57		3544.922		3603.516		3693.848

		9.58		3522.949		3601.074		3693.848

		9.59		3544.922		3598.633		3696.289

		9.60		3537.598		3598.633		3688.965

		9.61		3527.832		3603.516		3693.848

		9.62		3540.039		3601.074		3693.848

		9.63		3520.508		3601.074		3686.523

		9.64		3549.805		3601.074		3693.848

		9.65		3518.066		3603.516		3684.082

		9.66		3547.363		3601.074		3703.613

		9.67		3525.391		3603.516		3684.082

		9.68		3540.039		3601.074		3701.172

		9.69		3535.156		3596.191		3679.199

		9.70		3527.832		3603.516		3698.730

		9.71		3544.922		3601.074		3691.406

		9.72		3525.391		3601.074		3688.965

		9.73		3549.805		3598.633		3693.848

		9.74		3527.832		3598.633		3691.406

		9.75		3542.480		3598.633		3696.289

		9.76		3535.156		3601.074		3681.641

		9.77		3532.715		3603.516		3701.172

		9.78		3535.156		3601.074		3686.523

		9.79		3527.832		3598.633		3703.613

		9.80		3532.715		3601.074		3686.523

		9.81		3527.832		3605.957		3691.406

		9.82		3540.039		3601.074		3693.848

		9.83		3525.391		3601.074		3693.848

		9.84		3537.598		3598.633		3693.848

		9.85		3522.949		3598.633		3684.082

		9.86		3542.480		3601.074		3696.289

		9.87		3527.832		3603.516		3688.965

		9.88		3537.598		3603.516		3691.406

		9.89		3535.156		3598.633		3691.406

		9.90		3535.156		3596.191		3688.965

		9.91		3542.480		3601.074		3696.289

		9.92		3525.391		3601.074		3684.082

		9.93		3544.922		3603.516		3696.289

		9.94		3527.832		3601.074		3679.199

		9.95		3549.805		3598.633		3698.730

		9.96		3530.273		3603.516		3688.965

		9.97		3537.598		3603.516		3688.965

		9.98		3530.273		3601.074		3693.848

		9.99		3535.156		3601.074		3691.406

		10.00		3540.039		3603.516		3698.730

								3684.082

								3698.730

								3684.082

								3701.172

								3684.082

								3696.289

								3688.965

								3693.848

								3693.848

								3684.082

								3701.172

								3688.965

								3696.289

								3686.523

								3696.289

								3688.965

								3688.965

								3688.965

								3691.406

								3701.172

								3688.965

								3696.289

								3686.523

								3696.289

								3696.289

								3686.523





FO134

		



Time/sec.

Sensor Output, mV



10ºC


_940833721.xls
Chart1

		68.5

		67.4

		67.7

		64.5

		63.7

		68.4

		70.8

		69

		66.4

		66

		66

		67

		67.2

		66.4

		64.7

		66.7

		64.6

		65.8

		64.8

		66.6

		66.6

		65

		66.9

		66.9

		64.3

		64.3

		67.4

		67.5

		66.5

		65.8

		68.1

		65.3

		63.6

		64.1

		63

		66.3

		64.1

		63.1

		63.8

		62.8

		62.3

		62.3

		63.2

		64.1

		63.8

		61.4

		64.5

		59



Sensor Output in Oxygen, mV

T10-90, ms

490

500

500

510

510

510

510

510

510

510

520

520

530

530

530

530

530

530

530

530

540

540

540

540

540

540

540

540

540

550

550

560

560

560

560

570

570

570

570

570

570

580

580

590

590

600

610

620



Sheet1

		68.5		490				490		1		2.0833333333

		67.4		500				510		9		18.75

		67.7		500				530		10		20.8333333333

		64.5		510				550		11		22.9166666667

		63.7		510				570		10		20.8333333333

		68.4		510				590		4		8.3333333333

		70.8		510				610		2		4.1666666667

		69		510				630		1		2.0833333333

		66.4		510						48

		66		510

		66		520

		67		520

		67.2		530

		66.4		530

		64.7		530

		66.7		530

		64.6		530

		65.8		530

		64.8		530

		66.6		530

		66.6		540

		65		540

		66.9		540

		66.9		540

		64.3		540

		64.3		540

		67.4		540

		67.5		540

		66.5		540

		65.8		550

		68.1		550

		65.3		560

		63.6		560

		64.1		560

		63		560

		66.3		570

		64.1		570

		63.1		570

		63.8		570

		62.8		570

		62.3		570

		62.3		580

		63.2		580

		64.1		590

		63.8		590

		61.4		600

		64.5		610

		59		620





Sheet1

		



Sensor Output in Oxygen, mV

T10-90, ms



Sheet2

		





Sheet3

		






_940760657.xls
Chart3

		0		0		0		0		0		0		0		0

		21		21		21		21		21		21		21		21

		40		40		40		40		40		40		40		40

		60		60		60		60		60		60		60		60

		80		80		80		80		80		80		80		80

		100		100		100		100		100		100		100		100



Oxygen Concentration, %

Sensor Output, mV

at 50ºC

0.2441

0.2441

0.2441

0

0

0.2441

0.2441

0.2441

24.5117

26.2354

23.5229

26.8652

27.0825

27.5513

22.7808

35.8984

46.0352

49.4019

44.1357

50.4443

50.9717

51.7285

42.8931

53.2324

68.1543

73.2666

65.5151

74.7412

75.4053

76.6577

63.5718

75.6934

90.6616

97.1924

86.7212

98.8672

99.4434

101.2524

84.3433

98.0981

112.1606

120.1318

107.146

121.7871

122.5928

124.6069

104.4653

119.6948



Sheet1

		0		0.2441		0.2441		0.2441		0		0		0.2441		0.2441		0.2441

		21		24.5117		26.2354		23.5229		26.8652		27.0825		27.5513		22.7808		35.8984

		40		46.0352		49.4019		44.1357		50.4443		50.9717		51.7285		42.8931		53.2324

		60		68.1543		73.2666		65.5151		74.7412		75.4053		76.6577		63.5718		75.6934

		80		90.6616		97.1924		86.7212		98.8672		99.4434		101.2524		84.3433		98.0981

		100		112.1606		120.1318		107.146		121.7871		122.5928		124.6069		104.4653		119.6948

						Non-Linearity Error When Calibrated at 100% O2

		0		0.2176343564		0.2031934925		0.2278199839		0		0		0.1958960539		0.2336661073		0.2039353422

		21		0.854109197		0.8388469997		0.9540626808		1.0591507639		1.0914278816		1.110573331		0.8070498051		8.9916119999

		40		1.0440029743		1.123083147		1.1921116981		1.4200682995		1.5780535235		1.5133511868		1.0596628737		4.4734441262

		60		0.764921015		0.9885142818		1.1456330614		1.3703750233		1.5087509218		1.5196269227		0.8544655498		3.2386703516

		80		0.8319499004		0.9048062212		0.9374125026		1.180354898		1.1168355727		1.2574584554		0.7381015514		1.9568602813

		100		0		0		0		0		0		0		0		0





Sheet1

		



Oxygen Concentration, %

Sensor Output, mV

at 50ºC



Sheet2

		





Sheet3

		






_940831945.xls
Chart4

		320

		350

		380

		410

		440

		470

		500

		530

		560



T10-90/ms

Percentage

1.4492753623

5.7971014493

14.4927536232

24.6376811594

17.3913043478

24.6376811594

8.6956521739

1.4492753623

1.4492753623



Sheet1

		320		1						320		1		1.4492753623				490		1		2.0833333333		380

		330		2						350		4		5.7971014493				510		9		18.75		370

		330		3						380		10		14.4927536232				530		10		20.8333333333		420

		340		4						410		17		24.6376811594				550		11		22.9166666667		390

		350		5						440		12		17.3913043478				570		10		20.8333333333		400

		360		6						470		17		24.6376811594				590		4		8.3333333333		370

		360		7						500		6		8.6956521739				610		2		4.1666666667		390

		370		8						530		1		1.4492753623				630		1		2.0833333333		410

		370		9						560		1		1.4492753623										330

		370		10								69												360

		370		11																				350

		370		12																				380

		380		13																				330

		380		14																				360

		380		15																				320

		390		16																				340

		390		17																				400

		390		18																				670

		390		19																				390

		390		20																				390

		390		21																				460

		390		22																				380

		400		23																				410

		400		24																				370

		400		25																				470

		400		26																				490

		400		27																				490

		410		28																				490

		410		29																				400

		410		30																				460

		410		31

		410		32

		420		33

		420		34

		420		35

		420		36

		420		37

		420		38

		420		39

		430		40

		430		41

		430		42

		440		43

		440		44

		450		45

		450		46

		450		47

		450		48

		450		49

		450		50

		450		51

		460		52

		460		53

		460		54

		460		55

		460		56

		460		57

		460		58

		460		59

		470		60

		470		61

		480		62

		490		63

		490		64

		490		65

		490		66

		490		67

		510		68

		570		69

		670		70





Sheet1

		



T10-90/ms

Percentage



Sheet2

		



T10-90/ms

Percentage



Sheet3

		





		






_940760592.xls
Chart3

		0		0		0		0		0		0		0		0

		21		21		21		21		21		21		21		21

		40		40		40		40		40		40		40		40

		50		50		50		50		50		50		50		50

		60		60		60		60		60		60		60		60

		80		80		80		80		80		80		80		80

		100		100		100		100		100		100		100		100



Oxygen Concentration, %

Sensor Output, mV

at 26ºC

0.2515

0.2417

0.2417

0.2466

0.2637

0.249

0.2759

0.249

23.2715

24.7974

22.0825

25.52

26.0645

26.3745

21.9995

25.4517

43.7451

46.3989

41.2207

47.6099

48.6304

49.0771

41.0718

47.6294

54.1895

57.5488

51.0303

59.0015

60.2783

60.7959

50.8911

59.0576

64.6631

68.6035

60.8813

70.437

71.875

72.5122

60.7153

70.354

85.4834

90.9009

80.3174

92.9932

94.7754

95.6812

80.3564

92.7588

105.9619

112.5171

99.1406

114.7144

116.7993

117.9517

99.5996

114.4678



Sheet1

		0		0.2515		0.2417		0.2417		0.2466		0.2637		0.249		0.2759		0.249

		21		23.2715		24.7974		22.0825		25.52		26.0645		26.3745		21.9995		25.4517

		40		43.7451		46.3989		41.2207		47.6099		48.6304		49.0771		41.0718		47.6294

		50		54.1895		57.5488		51.0303		59.0015		60.2783		60.7959		50.8911		59.0576

		60		64.6631		68.6035		60.8813		70.437		71.875		72.5122		60.7153		70.354

		80		85.4834		90.9009		80.3174		92.9932		94.7754		95.6812		80.3564		92.7588

		100		105.9619		112.5171		99.1406		114.7144		116.7993		117.9517		99.5996		114.4678

		a=		0.9553708774		0.9005958753		1.0252004656		0.8857058291		0.8706692471		0.8626601381		1.0180399252		0.8874799053

		b=		-1.2329133733		-1.3324361574		-1.6389892821		-1.6032127578		-1.6935585905		-1.7522298127		-1.3963693346		-1.5878723062

						Non-Linearity Error When Calibrated with air and 100% O2

		0		-0.9926375976		-1.1147621344		-1.3911983296		-1.3847977003		-1.4639631101		-1.5374274383		-1.1154921192		-1.3668898098

		21		0		0		0		0		0		0		0		0

		40		0.5598811954		0.4542217997		0.6204915512		0.5651531935		0.6474351627		0.5846280515		0.4163628655		0.6822630962

		50		0.5381567872		0.4957757493		0.6772980388		0.6547597159		0.7889034858		0.693969678		0.4128023031		0.8245609502

		60		0.5443292087		0.4515929717		0.7765478256		0.7832487241		0.8857935434		0.8011546542		0.4142301363		0.8498889527

		80		0.4354374873		0.532539441		0.7024465955		0.7614065457		0.8244675692		0.7881273942		0.4096541113		0.7336987354

		100		0		0		0		0		0		0		0		0





Sheet1

		



Oxygen Concentration, %

Sensor Output, mV

at 26ºC



Sheet2

		





Sheet3

		






_940315645

_940419531.xls
Chart4

		23		23		23		23		23		23		23		23		23

		30		30		30		30		30		30		30		30		30

		35		35		35		35		35		35		35		35		35

		40		40		40		40		40		40		40		40		40

		45		45		45		45		45		45		45		45		45



S-11

S-12

S-13

S-14

S-15

S-16

S-17

S-18

S-19

Temperature, ºC

Output in Air, mV

19.22

21.79

18.42

20.57

17.57

22.93

20.94

23.11

22.333

22.07

25.11

21.33

23.92

20.08

26.23

23.92

26.31

25.63

24.58

28.05

23.8

26.85

22.63

29.79

27.14

29.88

29.36

27.79

31.5

27.21

30.63

25.45

33.36

30.31

33.39

32.93

30.98

35.54

30.09

34.28

28.6

37.54

34.1

37.6

37.3



Sheet1

												Sensor output

		T/ºC		S-11		S-12		S-13		S-14		S-15		S-16		S-17		S-18		S-19

		23		19.22		21.79		18.42		20.57		17.57		22.93		20.94		23.11		22.333

		30		22.07		25.11		21.33		23.92		20.08		26.23		23.92		26.31		25.63

		35		24.58		28.05		23.8		26.85		22.63		29.79		27.14		29.88		29.36

		40		27.79		31.5		27.21		30.63		25.45		33.36		30.31		33.39		32.93

		45		30.98		35.54		30.09		34.28		28.6		37.54		34.1		37.6		37.3

		23		1		1		1		1		1		1		1		1		1

		30		1.1482830385		1.1523634695		1.157980456		1.1628585318		1.1428571429		1.1439162669		1.1423113658		1.1384681956		1.1476290691

		35		1.2788761707		1.2872877467		1.2920738328		1.3052989791		1.2879908936		1.2991713912		1.2960840497		1.2929467763		1.3146464873

		40		1.4458896982		1.4456172556		1.4771986971		1.4890617404		1.4484917473		1.4548626254		1.4474689589		1.4448290783		1.4744996194

		45		1.6118626431		1.6310234052		1.6335504886		1.6665046184		1.6277746158		1.6371565635		1.6284622732		1.6270012981		1.6701741817





Sheet1

		



S-11

S-12

S-13

S-14

S-15

S-16

S-17

S-18

S-19

Temperature, ºC

Output in Air, mV



Sheet2

		



Temperature, ºC

Sensor Output in Air, mV



Sheet3

		





		






_940674998.xls
Chart4

		0		0

		23		23

		30		30

		35		35

		40		40

		45		45

		50		50



Without DC Temperature Compensation

With DC Temperature Compensation

Operation Temperature, ºC

Output in Air, mV

23.11

19.86

26.31

20.01

29.88

20.06

33.39

20.2

37.6

20.65

21.36



Sheet1

		

						Without DC Temperature Compensation		With DC Temperature Compensation

		23				23.11		19.86

		30				26.31		20.01

		35				29.88		20.06

		40				33.39		20.2

		45				37.6		20.65

		50						21.36





Sheet1

		



Without DC Temperature Compensation

With DC Temperature Compensation

Operation Temperature, ºC

Output in Air, mV



Sheet2

		





Sheet3

		






_940682039.xls
Chart3

		100		100		100

		150		150		150

		200		200		200

		250		250		250

		300		300		300



Cap-1

Cap-2

Cap-3

Flow Rate, cc/min

T10-90%, ms

S-11

625

530

480

625

490

470

505

450

470

470

450

460

450

440

450



Sheet1

				S-11		S-18		S19		S-11		S-18		S19		S-11		S-18		S19

		cc/min		Cap-1		Cap-1		Cap-1		Cap-2		Cap-2		Cap-2		Cap-3		Cap-3		Cap-3

		100		625				580		530		425		460		480		385		430

		150		625				560		490		370		410		470		370		410

		200		505				425		450		360		395		470		355		380

		250		470				390		450		340		380		460		350		360

		300		450				370		440		335		360		450		350		350





Sheet1

		



Cap-1

Cap-2

Cap-3

Flow Rate, cc/min

T10-90%, ms

S-11



Sheet2

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0



Cap-1

Cap-2

Cap-3

Flow Rate, cc/min

T10-90%, ms

Sensor-19

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0



Sheet3

		cc/min		S-11		S-18		S19

		100		530		425		460

		150		490		370		410

		200		450		360		395

		250		450		340		380

		300		440		335		360





Sheet3

		100		100		100

		150		150		150

		200		200		200

		250		250		250

		300		300		300



S-11

S-18

S19

Flow Rate, cc/min

T10-90%, ms

530

425

460

490

370

410

450

360

395

450

340

380

440

335

360



		cc/min		S-11		S-18		S19

		100		480		385		430

		150		470		370		410

		200		470		355		380

		250		460		350		360

		300		450		350		350





		100		100		100

		150		150		150

		200		200		200

		250		250		250

		300		300		300



S-11

S-18

S19

Flow Rate, cc/min

T10-90%, ms

480

385

430

470

370

410

470

355

380

460

350

360

450

350

350




_940424538

_940319708.xls
Chart1

		0

		0.002

		0.004

		0.006

		0.008

		0.01

		0.012

		0.014

		0.016

		0.018

		0.02

		0.022

		0.024

		0.026

		0.028

		0.03

		0.034

		0.036

		0.038

		0.04

		0.042

		0.044

		0.046

		0.048

		0.05

		0.052

		0.054

		0.056

		0.058

		0.06

		0.062

		0.064

		0.066

		0.068

		0.07

		0.072

		0.074

		0.076

		0.082

		0.087

		0.09

		0.094

		0.096

		0.098

		0.1

		0.102

		0.104

		0.106

		0.108

		0.11

		0.112

		0.114

		0.116

		0.118

		0.12

		0.122

		0.124

		0.126

		0.128

		0.13

		0.132

		0.134

		0.136

		0.138

		0.14

		0.144

		0.146

		0.148

		0.15

		0.152

		0.154

		0.156

		0.158

		0.16

		0.162

		0.164

		0.166

		0.168

		0.17

		0.172

		0.174

		0.176

		0.178

		0.18

		0.182

		0.184

		0.186

		0.188

		0.19

		0.192

		0.194

		0.197

		0.2

		0.204

		0.206

		0.208

		0.21

		0.212

		0.214

		0.216

		0.218

		0.22

		0.222

		0.224

		0.226

		0.228

		0.23

		0.232

		0.234

		0.236

		0.238

		0.24

		0.242

		0.244

		0.246

		0.248

		0.25

		0.254

		0.256

		0.258

		0.26

		0.262

		0.264

		0.266

		0.268

		0.27

		0.272

		0.274

		0.276

		0.278

		0.28

		0.282

		0.284

		0.286

		0.288

		0.29

		0.292

		0.294

		0.296

		0.298

		0.3

		0.302

		0.304

		0.307

		0.31

		0.314

		0.316

		0.318

		0.32

		0.322

		0.324

		0.326

		0.328

		0.333

		0.337

		0.339

		0.341

		0.343

		0.345

		0.347

		0.349

		0.351

		0.353

		0.355

		0.357

		0.359

		0.362

		0.365

		0.369

		0.371

		0.373

		0.375

		0.377

		0.379

		0.381

		0.383

		0.385

		0.387

		0.389

		0.391

		0.393

		0.395

		0.397

		0.399

		0.401

		0.403

		0.405

		0.407

		0.409

		0.411

		0.413

		0.415

		0.419

		0.421

		0.423

		0.425

		0.427

		0.429

		0.431

		0.433

		0.435

		0.437

		0.439

		0.441

		0.443

		0.445

		0.447

		0.449

		0.451

		0.453

		0.455

		0.457

		0.459

		0.461

		0.463

		0.465

		0.467

		0.469

		0.472

		0.475

		0.479

		0.481

		0.483

		0.485

		0.487

		0.489

		0.491

		0.493

		0.495

		0.497

		0.499

		0.501

		0.503

		0.505

		0.507

		0.509

		0.511

		0.513

		0.515

		0.517

		0.519

		0.521

		0.523

		0.525

		0.529

		0.531

		0.533

		0.535

		0.537

		0.539

		0.541

		0.543

		0.545

		0.547

		0.549

		0.551

		0.553

		0.555

		0.557

		0.559

		0.561

		0.563

		0.565

		0.567

		0.569

		0.571

		0.573

		0.575

		0.577

		0.579

		0.582

		0.59

		0.594

		0.596

		0.598

		0.6

		0.602

		0.604

		0.606

		0.608

		0.61

		0.612

		0.614

		0.616

		0.618

		0.62

		0.622

		0.624

		0.626

		0.628

		0.63

		0.632

		0.634

		0.638

		0.64

		0.642

		0.644

		0.646

		0.648

		0.65

		0.652

		0.654

		0.656

		0.658

		0.66

		0.662

		0.664

		0.666

		0.668

		0.67

		0.672

		0.674

		0.676

		0.678

		0.68

		0.682

		0.684

		0.686

		0.688

		0.691

		0.694

		0.698

		0.7

		0.702

		0.704

		0.706

		0.708

		0.71

		0.712

		0.714

		0.716

		0.718

		0.72

		0.722

		0.724

		0.726

		0.728

		0.73

		0.732

		0.734

		0.736

		0.738

		0.74

		0.742

		0.744

		0.748

		0.75

		0.752

		0.754

		0.756

		0.758

		0.76

		0.762

		0.764

		0.766

		0.768

		0.77

		0.772

		0.774

		0.776

		0.778

		0.78

		0.782

		0.784

		0.786

		0.788

		0.79

		0.792

		0.794

		0.796

		0.798

		0.801

		0.804

		0.808

		0.81

		0.812

		0.814

		0.816

		0.818

		0.82

		0.822

		0.824

		0.826

		0.828

		0.83

		0.832

		0.834

		0.84

		0.843

		0.847

		0.849

		0.853

		0.856

		0.86

		0.862

		0.864

		0.866

		0.868

		0.87

		0.872

		0.874

		0.876

		0.878

		0.88

		0.882

		0.884

		0.886

		0.888

		0.89

		0.892

		0.894

		0.896

		0.898

		0.9



Time, d^2/D

Oxygen Flux, Dco/d

0.0027319221

0.0005876319

0.0001243089

0.0000261263

0.0000068108

0.0000088143

0.0000296005

0.0000874019

0.0002158228

0.0004624286

0.0008862017

0.0015535374

0.0038936222

0.0056954744

0.007992521

0.010828058

0.0142344728

0.0182331977

0.0228352258

0.0280420054

0.0338465579

0.0402346987

0.0471862703

0.0546763268

0.0626762295

0.0711546327

0.0800783482

0.0894130891

0.0991240974

0.1091766636

0.1195365485

0.1301703196

0.1410456113

0.1521313213

0.1863615445

0.2156138765

0.2333223692

0.2569808965

0.2687923285

0.2805717851

0.2923057768

0.3039819942

0.3155892396

0.3271173605

0.338557184

0.3499004529

0.3611397641

0.3722685104

0.3832808233

0.3941715205

0.4049360544

0.4155704652

0.4260713352

0.4364357469

0.4466612432

0.4567457905

0.4666877439

0.476485815

0.4861390418

0.4956467609

0.5050085809

0.5232941788

0.5322183279

0.540997281

0.5496316806

0.5581223213

0.5664701343

0.5746761736

0.5827416032

0.5906676851

0.5984557685

0.6061072794

0.613623712

0.6210066195

0.6282576069

0.6353783232

0.6423704556

0.6492357229

0.6559758704

0.6625926644

0.6690878881

0.6754633372

0.6817208161

0.6878621344

0.6938891036

0.6998035347

0.7056072351

0.7141091047

0.7223719276

0.7330275213

0.7382043344

0.7432828029

0.7482646395

0.7531515389

0.7579451778

0.7626472131

0.7672592822

0.7717830016

0.7762199672

0.7805717534

0.7848399129

0.7890259764

0.7931314526

0.7971578279

0.8011065661

0.8049791088

0.8087768749

0.8125012608

0.8161536406

0.819735366

0.8232477663

0.8266921488

0.8300697987

0.8366299325

0.8398148785

0.8429380163

0.8460005238

0.8490035582

0.851948256

0.8548357336

0.8576670871

0.860443393

0.863165708

0.8658350698

0.8684524969

0.8710189892

0.8735355282

0.876003077

0.8784225811

0.8807949682

0.8831211487

0.8854020159

0.8876384466

0.8898313008

0.8919814224

0.8940896393

0.8961567638

0.8981835927

0.9001709079

0.9030794656

0.9059033658

0.9095411091

0.9113069422

0.9130383313

0.9147359463

0.9164004441

0.9180324691

0.9196326532

0.9212016161

0.9249913907

0.9278916114

0.9292994129

0.93067974

0.9320331282

0.9333601023

0.9346611773

0.935936858

0.9371876394

0.938414007

0.9396164368

0.9407953955

0.941951341

0.9436430886

0.9452855423

0.9474012756

0.948428267

0.9494352093

0.9504224938

0.9513905041

0.9523396161

0.9532701987

0.9541826135

0.955077215

0.9559543508

0.9568143618

0.9576575821

0.9584843396

0.9592949555

0.9600897448

0.9608690165

0.9616330735

0.9623822127

0.9631167253

0.9638368969

0.9645430074

0.9652353312

0.9659141375

0.9665796902

0.9678720646

0.9684993887

0.9691144642

0.9697175303

0.9703088213

0.9708885672

0.9714569934

0.9720143209

0.9725607663

0.9730965421

0.9736218566

0.9741369141

0.9746419148

0.9751370551

0.9756225275

0.9760985207

0.9765652198

0.9770228064

0.9774714582

0.9779113498

0.9783426521

0.978765533

0.9791801568

0.9795866848

0.9799852749

0.9803760823

0.9809480322

0.9815033124

0.9822185931

0.982565793

0.9829062135

0.983239987

0.9835672433

0.9838881095

0.9842027106

0.9845111688

0.9848136041

0.985110134

0.9854008739

0.9856859368

0.9859654336

0.986239473

0.9865081614

0.9867716035

0.9870299016

0.9872831561

0.9875314657

0.9877749267

0.9880136339

0.9882476801

0.9884771564

0.9887021518

0.9891390488

0.9893511201

0.9895590506

0.989762921

0.9899628106

0.9901587972

0.9903509569

0.9905393645

0.9907240933

0.9909052151

0.9910828002

0.9912569179

0.9914276357

0.99159502

0.991759136

0.9919200475

0.992077817

0.9922325059

0.9923841744

0.9925328813

0.9926786846

0.9928216409

0.9929618059

0.993099234

0.9932339787

0.9933660923

0.9935594417

0.9940479337

0.994278105

0.994389831

0.9944993755

0.9946067809

0.9947120892

0.9948153412

0.9949165771

0.9950158362

0.9951131572

0.995208578

0.9953021355

0.9953938662

0.9954838058

0.9955719892

0.9956584508

0.9957432241

0.9958263421

0.9959078371

0.9959877409

0.9960660845

0.9961428983

0.9962920556

0.9963644571

0.9964354449

0.9965050466

0.9965732892

0.9966401993

0.9967058029

0.9967701256

0.9968331923

0.9968950275

0.9969556554

0.9970150994

0.9970733827

0.997130528

0.9971865574

0.9972414928

0.9972953556

0.9973481666

0.9973999464

0.9974507152

0.9975004926

0.9975492981

0.9975971507

0.9976440688

0.9976900708

0.9977351746

0.9978011844

0.9978652702

0.997947822

0.9979878929

0.9980271814

0.9980657028

0.998103472

0.9981405037

0.9981768123

0.9982124119

0.9982473165

0.9982815395

0.9983150942

0.9983479937

0.9983802509

0.9984118782

0.9984428879

0.9984732922

0.9985031028

0.9985323312

0.998560989

0.9985890872

0.9986166368

0.9986436484

0.9986701326

0.9986960996

0.9987465225

0.998770998

0.9987949956

0.9988185245

0.9988415941

0.9988642132

0.9988863906

0.998908135

0.9989294548

0.9989503583

0.9989708537

0.9989909488

0.9990106516

0.9990299697

0.9990489105

0.9990674815

0.9990856899

0.9991035428

0.9991210471

0.9991382096

0.9991550369

0.9991715357

0.9991877123

0.9992035731

0.9992191242

0.9992343716

0.9992566863

0.9992783507

0.9993062574

0.9993198035

0.999333085

0.9993461073

0.9993588752

0.9993713938

0.999383668

0.9993957026

0.9994075021

0.9994190712

0.9994304145

0.9994415362

0.9994524408

0.9994631325

0.9994939711

0.9995087197

0.9995277179

0.9995369397

0.9995548466

0.9995678209

0.9995845336

0.999592646

0.9996006

0.9996083987

0.9996160452

0.9996235423

0.999630893

0.9996381002

0.9996451667

0.9996520952

0.9996588884

0.9996655489

0.9996720795

0.9996784824

0.9996847604

0.9996909158

0.999696951

0.9997028683

0.9997086701

0.9997143587

0.9997199361



FOTXT

		

		Idea gas @ 25ºC, 1atm=		2.24E-02		m3/mol						Faradaic C=		9.65E+04		C/mol

		D		Dia, inch		A, m2		T, mil		m		I,  A		F, mol/sec		C, mol/m3

		1.00E+00		0.36		1		0.19		1.00E+00		0.00023		5.96E-10		9.38E+00						0.012

										D, m2/sec		d, m		C		1		2		3		4		5		6		7		8

		0								1.00E+00		1		1		-2		2		-2		2		-2		2		-2		2

		0.002														-1.9609479035		1.8483075464		-1.6747675517		1.4588376384		-1.22160697		0.9833953825		-0.7610215379		0.5661582652

		0.004														-1.9226583401		1.708120393		-1.4024231762		1.0641036276		-0.7461617946		0.4835332392		-0.2895768906		0.1602675906

		0.006														-1.8851164206		1.5785659062		-1.1743664146		0.7761772116		-0.4557582245		0.2377521774		-0.1101871253		0.0453684105

		0.008						0.0027319221								-1.8483075464		1.4588376384		-0.9833953825		0.5661582652		-0.2783787119		0.1169021967		-0.0419273878		0.0128428503

		0.01						0.0005876319								-1.812217404		1.348190308		-0.8234793386		0.4129664933		-0.1700346874		0.0574805402		-0.0159538226		0.0036355429

		0.012						0.0001243089								-1.7768319595		1.2459351601		-0.6895682379		0.3012255319		-0.1038577796		0.0282630489		-0.0060706013		0.0010291463

		0.014						0.0000261263								-1.7421374529		1.1514356793		-0.5774332548		0.2197195718		-0.0634366937		0.0138968759		-0.0023099292		0.0002913299

		0.016						0.0000068108								-1.708120393		1.0641036276		-0.4835332392		0.1602675906		-0.0387473536		0.0068330618		-0.0008789529		0.0000824694

		0.018						0.0000088143								-1.6747675517		0.9833953825		-0.4049028896		0.1169021967		-0.0236670186		0.0033598007		-0.0003344511		0.0000233454

		0.02						0.0000296005								-1.6420659597		0.9088085533		-0.3390591105		0.0852706623		-0.0144558975		0.0016520063		-0.0001272622		0.0000066086

		0.022						0.0000874019								-1.6100029005		0.8398788536		-0.2839225982		0.0621980258		-0.0088297125		0.0008122877		-0.0000484246		0.0000018708

		0.024						0.0002158228								-1.5785659062		0.7761772116		-0.2377521774		0.0453684105		-0.0053932192		0.0003994		-0.0000184261		0.0000005296

		0.026						0.0004624286								-1.5477427521		0.7173070987		-0.199089816		0.0330925724		-0.0032941971		0.000196384		-0.0000070113		0.0000001499

		0.028						0.0008862017								-1.5175214525		0.6629020618		-0.1667145819		0.0241383451		-0.0020121071		0.0000965616		-0.0000026679		0.0000000424

		0.03						0.0015535374								-1.4878902553		0.6126234417		-0.139604086		0.0176069632		-0.001229002		0.0000474791		-0.0000010152		0.000000012

		0.034						0.0038936222								-1.4303523043		0.523217297		-0.0978920029		0.0093678167		-0.0004585172		0.0000114789		-0.000000147		0.000000001

		0.036						0.0056954744								-1.4024231762		0.4835332392		-0.081973175		0.0068330618		-0.0002800639		0.0000056441		-0.0000000559		0.0000000003

		0.038						0.007992521								-1.3750393935		0.4468590674		-0.0686430068		0.0049841639		-0.000171064		0.0000027752		-0.0000000213		0.0000000001

		0.04						0.010828058								-1.348190308		0.4129664933		-0.0574805402		0.0036355429		-0.0001044865		0.0000013646		-0.0000000081		0

		0.042						0.0142344728								-1.321865479		0.3816445429		-0.0481332718		0.0026518334		-0.0000638207		0.000000671		-0.0000000031		0

		0.044						0.0182331977								-1.2960546698		0.3526982444		-0.0403060209		0.0019342972		-0.0000389819		0.0000003299		-0.0000000012		0

		0.046						0.0228352258								-1.2707478438		0.3259474133		-0.033751608		0.0014109128		-0.0000238103		0.0000001622		-0.0000000004		0

		0.048						0.0280420054								-1.2459351601		0.3012255319		-0.0282630489		0.0010291463		-0.0000145434		0.0000000798		-0.0000000002		0

		0.05						0.0338465579								-1.22160697		0.2783787119		-0.0236670186		0.0007506787		-0.0000088832		0.0000000392		-0.0000000001		0

		0.052						0.0402346987								-1.1977538134		0.2572647369		-0.0198183774		0.0005475592		-0.0000054259		0.0000000193		-0		0

		0.054						0.0471862703								-1.1743664146		0.2377521774		-0.0165955877		0.0003994		-0.0000033141		0.0000000095		-0		0

		0.056						0.0546763268								-1.1514356793		0.2197195718		-0.0138968759		0.0002913299		-0.0000020243		0.0000000047		-0		9.0043002471693E-16

		0.058						0.0626762295								-1.1289526907		0.2030546713		-0.0116370184		0.0002125015		-0.0000012364		0.0000000023		-0		2.54892950346206E-16

		0.06						0.0711546327								-1.106908706		0.1876537406		-0.0097446504		0.0001550026		-0.0000007552		0.0000000011		-0		7.21548752848607E-17

		0.062						0.0800783482								-1.0852951532		0.1734209125		-0.0081600122		0.0001130618		-0.0000004613		0.0000000006		-0		2.04255395070845E-17

		0.064						0.0894130891								-1.0641036276		0.1602675906		-0.0068330618		0.0000824694		-0.0000002818		0.0000000003		-0		5.78204400615193E-18

		0.066						0.0991240974								-1.0433258888		0.1481118986		-0.0057218951		0.0000601547		-0.0000001721		0.0000000001		-0		1.63677600180314E-18

		0.068						0.1091766636								-1.0229538572		0.1368781699		-0.0047914221		0.000043878		-0.0000001051		0.0000000001		-0		4.63337130818828E-19

		0.07						0.1195365485								-1.0029796108		0.1264964772		-0.0040122591		0.0000320054		-0.0000000642		0		-0		1.3116107308448E-19

		0.072						0.1301703196								-0.9833953825		0.1169021967		-0.0033598007		0.0000233454		-0.0000000392		0		-0		3.71289627970681E-20

		0.074						0.1410456113								-0.9641935568		0.1080356062		-0.0028134426		0.0000170285		-0.000000024		0		-5.95123682952985E-16		1.05104345822035E-20

		0.076						0.1521313213								-0.9453666669		0.0998415131		-0.0023559312		0.0000124209		-0.0000000146		0		-2.26450970215904E-16		2.9752847045731E-21

		0.082						0.1863615445								-0.8910631136		0.0788032043		-0.0013833632		0.0000048204		-0.0000000033		0		-1.24759907142066E-17		6.74919689591503E-23

		0.087						0.2156138765								-0.8482006807		0.0647000296		-0.0008876619		0.0000021904		-0.000000001		0		-1.11427509498457E-18		2.87754244574909E-24

		0.09						0.2333223692								-0.8234793386		0.0574805402		-0.0006801965		0.0000013646		-0.0000000005		0		-2.61542850267088E-19		4.33394626868791E-25

		0.094						0.2569808965								-0.7916347091		0.0490918415		-0.0004769617		0.000000726		-0.0000000002		0		-3.78683826644451E-20		3.47295563146923E-26

		0.096						0.2687923285								-0.7761772116		0.0453684105		-0.0003994		0.0000005296		-0.0000000001		0		-1.44093274062296E-20		9.83121267646484E-27

		0.098						0.2805717851								-0.7610215379		0.0419273878		-0.0003344511		0.0000003863		-0.0000000001		0		-5.48290425127826E-21		2.78301115666699E-27

		0.1						0.2923057768								-0.7461617946		0.0387473536		-0.0002800639		0.0000002818		-0		7.69021443823317E-16		-2.08630411268804E-21		7.87812384190841E-28

		0.102						0.3039819942								-0.7315922034		0.035808513		-0.000234521		0.0000002055		-0		3.78126068457094E-16		-7.93861182165329E-22		2.23013246352834E-28

		0.104						0.3155892396								-0.7173070987		0.0330925724		-0.000196384		0.0000001499		-0		1.85923714865444E-16		-3.02072728858759E-22		6.31303963314964E-29

		0.106						0.3271173605								-0.7033009257		0.0305826257		-0.0001644488		0.0000001093		-0		9.1418261349758E-17		-1.14941926334388E-22		1.78708978329763E-29

		0.108						0.338557184								-0.6895682379		0.0282630489		-0.0001377068		0.0000000798		-0		4.49501480446484E-17		-4.37366407731473E-23		5.0588782569917E-30

		0.11						0.3499004529								-0.6761036952		0.0261194033		-0.0001153134		0.0000000582		-0		2.21018840153334E-17		-1.66422628115207E-23		1.43206286881899E-30

		0.112						0.3611397641								-0.6629020618		0.0241383451		-0.0000965616		0.0000000424		-0		1.08674453428282E-17		-6.33256021934291E-24		4.05387114705865E-31

		0.114						0.3722685104								-0.6499582042		0.0223075427		-0.0000808591		0.000000031		-0		5.34349778495915E-18		-2.40960735843229E-24		1.14756632790202E-31

		0.116						0.3832808233								-0.6372670889		0.0206155998		-0.0000677101		0.0000000226		-0		2.62738552411551E-18		-9.16881548804841E-25		3.24852080680968E-32

		0.118						0.3941715205								-0.624823781		0.0190519843		-0.0000566993		0.0000000165		-0		1.29187939625665E-18		-3.48883303164262E-25		9.19588452161035E-33

		0.12						0.4049360544								-0.6126234417		0.0176069632		-0.0000474791		0.000000012		-0		6.35214116525317E-19		-1.32753853957981E-25		2.6031630136869E-33

		0.122						0.4155704652								-0.6006613268		0.0162715415		-0.0000397582		0.0000000088		-0		3.12333314551042E-19		-5.05142710495361E-26		7.36901127879843E-34

		0.124						0.4260713352								-0.5889327847		0.0150374064		-0.0000332929		0.0000000064		-0		1.53573569668224E-19		-1.92212241195924E-26		2.08601332077734E-34

		0.126						0.4364357469								-0.5774332548		0.0138968759		-0.0000278789		0.0000000047		-0		7.55117696443699E-20		-7.31388276974844E-27		5.90506841396791E-35

		0.128						0.4466612432								-0.5661582652		0.0128428503		-0.0000233454		0.0000000034		-0		3.71289627970678E-20		-2.78301115666701E-27		1.67160164445387E-35

		0.13						0.4567457905								-0.5551034316		0.0118687686		-0.000019549		0.0000000025		-0		1.8256225286184E-20		-1.05896571519691E-27		4.73195543531957E-36

		0.132						0.4666877439								-0.5442644552		0.0109685672		-0.00001637		0.0000000018		-0		8.97654382433285E-21		-4.02947858572559E-28		1.33951784003933E-36

		0.134						0.476485815								-0.5336371211		0.0101366428		-0.000013708		0.0000000013		-0		4.41374587391561E-21		-1.53325999508887E-28		3.79189548234286E-37

		0.136						0.4861390418								-0.523217297		0.0093678167		-0.0000114789		0.000000001		-0		2.17022865601115E-21		-5.83421939718937E-29		1.07340648397712E-37

		0.138						0.4956467609								-0.5130009308		0.0086573032		-0.0000096122		0.0000000007		-0		1.06709641966624E-21		-2.21998330899956E-29		3.0385897639041E-38

		0.14						0.5050085809								-0.5029840498		0.0080006794		-0.0000080491		0.0000000005		-0		5.24688845901337E-22		-8.44727555945342E-30		8.60161354633614E-39

		0.144						0.5232941788								-0.4835332392		0.0068330618		-0.0000056441		0.0000000003		-7.69021443823317E-16		1.26852248610186E-22		-1.22306789511063E-30		6.89279939193032E-40

		0.146						0.5322183279								-0.4740917458		0.0063147998		-0.0000047263		0.0000000002		-4.69720977937336E-16		6.23729577727248E-23		-4.65390505236553E-31		1.95120767291825E-40

		0.148						0.540997281								-0.4648346075		0.0058358461		-0.0000039577		0.0000000001		-2.86907210308038E-16		3.0668639333885E-23		-1.77086099005763E-31		5.52346175533901E-41

		0.15						0.5496316806								-0.4557582245		0.0053932192		-0.0000033141		0.0000000001		-1.7524392393163E-16		1.50796991545464E-23		-6.73831677015832E-32		1.56357676254233E-41

		0.152						0.5581223213								-0.4468590674		0.0049841639		-0.0000027752		0.0000000001		-1.07039599464867E-16		7.41465325918072E-24		-2.56400209558621E-32		4.42615953663332E-42

		0.154						0.5664701343								-0.4381336757		0.0046061338		-0.0000023239		0.0000000001		-6.53801603875815E-17		3.64576788903006E-24		-9.7563040895985E-33		1.25295340229258E-42

		0.156						0.5746761736								-0.4295786564		0.004256776		-0.000001946		0		-3.99344298154716E-17		1.79261565390468E-24		-3.71237877116292E-33		3.54684962284634E-43

		0.158						0.5827416032								-0.4211906829		0.0039339156		-0.0000016296		0		-2.4392088903314E-17		8.81424978340853E-25		-1.41260010081835E-33		1.00403911462841E-43

		0.16						0.5906676851								-0.4129664933		0.0036355429		-0.0000013646		0		-1.48987729089014E-17		4.33394626868791E-25		-5.37509550569622E-34		2.84222521645794E-44

		0.162						0.5984557685								-0.4049028896		0.0033598007		-0.0000011427		0		-9.10022241518063E-18		2.13099137435726E-25		-2.04528172400798E-34		8.04574648873016E-45

		0.164						0.6061072794								-0.3969967362		0.0031049725		-0.0000009568		0		-5.5584475655896E-18		1.04780353886571E-25		-7.78251721504841E-35		2.27758293699146E-45

		0.166						0.613623712								-0.3892449588		0.0028694721		-0.0000008012		0		-3.3951191443265E-18		5.15202580953969E-26		-2.96133160980072E-35		6.4473620218344E-46

		0.168						0.6210066195								-0.3816445429		0.0026518334		-0.000000671		0		-2.07375060539043E-18		2.53323919586083E-26		-1.12681856793686E-35		1.82511364857261E-46

		0.17						0.6282576069								-0.3741925332		0.0024507019		-0.0000005618		0		-1.26665409682117E-18		1.24558786401325E-26		-4.28766599743533E-36		5.16651588498566E-47

		0.172						0.6353783232								-0.3668860317		0.0022648254		-0.0000004705		0		-7.73676736644907E-19		6.1245267699637E-27		-1.63150308564965E-36		1.46253283518457E-47

		0.174						0.6423704556								-0.3597221974		0.0020930469		-0.000000394		0		-4.72564447016525E-19		3.01141567284912E-27		-6.20804493651438E-37		4.14012526354382E-48

		0.176						0.6492357229								-0.3526982444		0.0019342972		-0.0000003299		0		-2.88644011131134E-19		1.48070613376306E-27		-2.36222795241819E-37		1.17198306837815E-48

		0.178						0.6559758704								-0.3458114414		0.0017875881		-0.0000002763		0		-1.7630476792715E-19		7.28059787405313E-28		-8.98853174590416E-38		3.3176395039539E-49

		0.18						0.6625926644								-0.3390591105		0.0016520063		-0.0000002313		0		-1.07687566674387E-19		3.5798531656619E-28		-3.42023312629161E-38		9.39154512993694E-50

		0.182						0.6690878881								-0.332438626		0.0015267078		-0.0000001937		0		-6.57759410173308E-20		1.76020553660453E-28		-1.30143553684536E-38		2.65855044897214E-50

		0.184						0.6754633372								-0.3259474133		0.0014109128		-0.0000001622		0		-4.0176174003423E-20		8.65488998490917E-29		-4.95210236853188E-39		7.52580155016242E-51

		0.186						0.6817208161								-0.3195829484		0.0013039004		-0.0000001358		0		-2.45397470957966E-20		4.25558942368618E-29		-1.88432828012846E-39		2.13039737479211E-51

		0.188						0.6878621344								-0.3133427563		0.0012050044		-0.0000001137		0		-1.49889630474606E-20		2.09246349457551E-29		-7.17007202810424E-40		6.03071040907691E-52

		0.19						0.6938891036								-0.3072244105		0.0011136094		-0.0000000952		0		-9.15531086612635E-21		1.0288594693279E-29		-2.72828962078187E-40		1.70716827144502E-52

		0.192						0.6998035347								-0.3012255319		0.0010291463		-0.0000000798		0		-5.59209578341123E-21		5.05887825699163E-30		-1.03814358149959E-40		4.83263713449415E-53

		0.194						0.7056072351								-0.2953437876		0.0009510895		-0.0000000668		0		-3.41567159303647E-21		2.48743875932647E-30		-3.95024812468379E-41		1.36801872810839E-53

		0.197						0.7141091047								-0.2867358072		0.0008449631		-0.0000000512		0		-1.63052839897179E-21		8.57629381061911E-31		-9.27202948753263E-42		2.06041084500391E-54

		0.2						0.7223719276								-0.2783787119		0.0007506787		-0.0000000392		0		-7.78360210411813E-22		2.95696990530049E-31		-2.17633242530952E-42		3.10324176342228E-55

		0.204						0.7330275213								-0.267613576		0.0006411248		-0.0000000275		0		-2.90391325726594E-22		7.14896618234094E-32		-3.15107788274467E-43		2.4867454024415E-56

		0.206						0.7382043344								-0.2623881404		0.0005924979		-0.000000023		0		-1.77372033771473E-22		3.51513016675041E-32		-1.1990190681586E-43		7.03945731471393E-57

		0.208						0.7432828029								-0.2572647369		0.0005475592		-0.0000000193		0		-1.08339456371534E-22		1.72838138746835E-32		-4.56239667600887E-44		1.99272347048591E-57

		0.21						0.7482646395								-0.2522413733		0.0005060289		-0.0000000161		0		-6.61741175161927E-23		8.49841137834354E-33		-1.73604106740554E-44		5.64098431497736E-58

		0.212						0.7531515389								-0.247316096		0.0004676485		-0.0000000135		0		-4.04193815965786E-23		4.17864925410633E-33		-6.60582321472995E-45		1.59684494678338E-58

		0.214						0.7579451778								-0.24248699		0.0004321791		-0.0000000113		0		-2.46882991412776E-23		2.054632190828E-33		-2.51358687093033E-45		4.52033482400819E-59

		0.216						0.7626472131								-0.2377521774		0.0003994		-0.0000000095		0		-1.50796991545463E-23		1.0102579046179E-33		-9.56446873059667E-46		1.27961246095317E-59

		0.218						0.7672592822								-0.2331098169		0.000369107		-0.0000000079		0		-9.21073279655089E-24		4.96741479277442E-34		-3.63938335119897E-46		3.62231585485672E-60

		0.22						0.7717830016								-0.2285581033		0.0003411116		-0.0000000066		0		-5.62594769166071E-24		2.44246638513626E-34		-1.38482455743862E-46		1.02540203012504E-60

		0.222						0.7762199672								-0.2240952668		0.0003152396		-0.0000000056		0		-3.43634845656959E-24		1.20095508255082E-34		-5.26940657200432E-47		2.90269917233967E-61

		0.224						0.7805717534								-0.2197195718		0.0002913299		-0.0000000047		9.0043002471693E-16		-2.09893361299322E-24		5.90506841396791E-35		-2.00506594658021E-47		8.21693563847732E-62

		0.226						0.7848399129								-0.2154293168		0.0002692336		-0.0000000039		6.56790605406585E-16		-1.28203596562807E-24		2.90350850587835E-35		-7.62949185112219E-48		2.32604301300906E-62

		0.228						0.7890259764								-0.2112228336		0.0002488132		-0.0000000033		4.79075427861213E-16		-7.83072035718174E-25		1.42764842889316E-35		-2.90310381090552E-48		6.58454238467264E-63

		0.23						0.7931314526								-0.2070984863		0.0002299417		-0.0000000027		3.49446632901099E-16		-4.7830312843322E-25		7.01971436417247E-36		-1.10466226340548E-48		1.86394654669195E-63

		0.232						0.7971578279								-0.2030546713		0.0002125015		-0.0000000023		2.54892950346206E-16		-2.92149217739834E-25		3.45157734616586E-36		-4.20335887269382E-49		5.27644371613771E-64

		0.234						0.8011065661								-0.199089816		0.000196384		-0.0000000019		1.85923714865444E-16		-1.78445760339436E-25		1.69713261231389E-36		-1.59942331678699E-49		1.49365111027352E-64

		0.236						0.8049791088								-0.1952023786		0.0001814891		-0.0000000016		1.35616256559532E-16		-1.08995292301199E-25		8.34476187236225E-37		-6.08597796134031E-50		4.22821460673964E-65

		0.238						0.8087768749								-0.1913908476		0.0001677238		-0.0000000013		9.89210497247814E-17		-6.6574704387629E-26		4.10310014675214E-37		-2.31578015383234E-50		1.1969193232396E-65

		0.24						0.8125012608								-0.1876537406		0.0001550026		-0.0000000011		7.21548752848607E-17		-4.06640614537022E-26		2.01748486916519E-37		-8.81179287034849E-51		3.38822883791373E-66

		0.242						0.8161536406								-0.1839896046		0.0001432462		-0.0000000009		5.26311239302339E-17		-2.48377504507215E-26		9.91992652319819E-38		-3.35298208085202E-51		9.59136880420483E-67

		0.244						0.819735366								-0.1803970147		0.0001323815		-0.0000000008		3.83901322706719E-17		-1.51709845351944E-26		4.87760496892199E-38		-1.275845789833E-51		2.71511636135292E-67

		0.246						0.8232477663								-0.1768745739		0.0001223409		-0.0000000007		2.80024849500326E-17		-9.26649022518148E-27		2.39830710209558E-38		-4.8547306254049E-52		7.6859278442659E-68

		0.248						0.8266921488								-0.1734209125		0.0001130618		-0.0000000006		2.04255395070845E-17		-5.6600045233837E-27		1.17924206503204E-38		-1.84727728327801E-52		2.17572578723026E-68

		0.25						0.8300697987								-0.1700346874		0.0001044865		-0.0000000005		1.48987729089015E-17		-3.4571504880745E-27		5.79830600812525E-39		-7.02908899508801E-53		6.1590256858193E-69

		0.254						0.8366299325								-0.1634593049		0.0000892377		-0.0000000003		7.92691914975067E-18		-1.28979680620768E-27		1.40183684293328E-39		-1.01773086733822E-53		4.93546103570076E-70

		0.256						0.8398148785								-0.1602675906		0.0000824694		-0.0000000003		5.78204400615193E-18		-7.87812384190835E-28		6.89279939193022E-40		-3.87257554906652E-54		1.39712602887045E-70

		0.258						0.8429380163								-0.1571381979		0.0000762144		-0.0000000002		4.21753171156404E-18		-4.81198549800496E-28		3.38917354732807E-40		-1.47355669996042E-54		3.95497224358132E-71

		0.26						0.8460005238								-0.1540699099		0.0000704338		-0.0000000002		3.07634700101259E-18		-2.93917751201548E-28		1.66644880849952E-40		-5.60704192981256E-55		1.11957011209252E-71

		0.262						0.8490035582								-0.1510615334		0.0000650917		-0.0000000002		2.24394539694629E-18		-1.79525986741214E-28		8.1938903174162E-41		-2.13353983620176E-55		3.16926936194082E-72

		0.264						0.851948256								-0.1481118986		0.0000601547		-0.0000000001		1.63677600180314E-18		-1.09655098352008E-28		4.02891695150841E-41		-8.11834883641055E-56		8.97154021891815E-73

		0.266						0.8548357336								-0.1452198585		0.0000555922		-0.0000000001		1.19389521853984E-18		-6.69777162228964E-29		1.98100916333359E-41		-3.08911915828015E-56		2.53965582308144E-73

		0.268						0.8576670871								-0.1423842885		0.0000513758		-0.0000000001		8.70849640562923E-19		-4.09102224871852E-29		9.74057631975351E-42		-1.17544310626995E-56		7.18923567450611E-74

		0.27						0.860443393								-0.139604086		0.0000474791		-0.0000000001		6.35214116525312E-19		-2.49881064678472E-29		4.7894188879613E-42		-4.47268760214103E-57		2.03512259865519E-74

		0.272						0.863165708								-0.1368781699		0.000043878		-0.0000000001		4.63337130818828E-19		-1.52628225144475E-29		2.35494620968597E-42		-1.70190579872704E-57		5.76100739922061E-75

		0.274						0.8658350698								-0.1342054802		0.00004055		-0.0000000001		3.37966822856131E-19		-9.32258518296588E-30		1.15792161434329E-42		-6.47593484140093E-58		1.63082097736058E-75

		0.276						0.8684524969								-0.1315849775		0.0000374744		-0		2.46519360858497E-19		-5.69426751909014E-30		5.69347384432252E-43		-2.46416294611831E-58		4.61651387665139E-76

		0.278						0.8710189892								-0.1290156429		0.0000346322		-0		1.79815861108803E-19		-3.478078445259E-30		2.79946794450063E-43		-9.37640537425338E-59		1.30683874374777E-76

		0.28						0.8735355282								-0.1264964772		0.0000320054		-0		1.3116107308448E-19		-2.12442243551426E-30		1.37649192506638E-43		-3.56782321886696E-59		3.69938778002567E-77

		0.282						0.876003077								-0.1240265009		0.0000295779		-0		9.56713550550641E-20		-1.29760462725283E-30		6.76817901592712E-44		-1.35759515645943E-59		1.04721948384819E-77

		0.284						0.8784225811								-0.1216047534		0.0000273346		-0		6.97844868360957E-20		-7.92581428495574E-31		3.32789799616343E-44		-5.16579576896009E-60		2.96445983109031E-78

		0.286						0.8807949682								-0.1192302932		0.0000252613		-0		5.09021179868764E-20		-4.84111498682014E-31		1.6363197614612E-44		-1.96564092024323E-60		8.39176717554457E-79

		0.288						0.8831211487								-0.1169021967		0.0000233454		-0		3.71289627970681E-20		-2.95696990530049E-31		8.04574648873027E-45		-7.47947538025174E-61		2.37553417286975E-79

		0.29						0.8854020159								-0.1146195588		0.0000215747		-0		2.70825642017781E-20		-1.80612752323735E-31		3.95607497296918E-45		-2.84602092822068E-61		6.72464153071045E-80

		0.292						0.8876384466								-0.1123814917		0.0000199383		-0		1.97545320011295E-20		-1.10318898557204E-31		1.94519293065906E-45		-1.08294161182167E-61		1.90360569142753E-80

		0.294						0.8898313008								-0.1101871253		0.0000184261		-0		1.44093274062296E-20		-6.73831677015823E-32		9.56446873059667E-46		-4.1207094543331E-62		5.388710479044E-81

		0.296						0.8919814224								-0.1080356062		0.0000170285		-0		1.05104345822035E-20		-4.11578736633704E-32		4.70282719295958E-46		-1.56797432305394E-62		1.52543148813464E-81

		0.298						0.8940896393								-0.1059260977		0.000015737		-0		7.66650878229203E-21		-2.51393726693288E-32		2.31236927317136E-46		-5.96631115346215E-63		4.31817822471967E-82

		0.3						0.8961567638								-0.1038577796		0.0000145434		-0		5.59209578341131E-21		-1.53552164374774E-32		1.13698663295813E-46		-2.27024564474984E-63		1.22238614618118E-82

		0.302						0.8981835927								-0.1018298476		0.0000134403		-0		4.07897990322228E-21		-9.37901971314662E-33		5.59053702418614E-47		-8.63852915970483E-64		3.46032009939244E-83

		0.304						0.9001709079								-0.0998415131		0.0000124209		-0		2.9752847045731E-21		-5.72873792679955E-33		2.74885414769397E-47		-3.28705337308524E-64		9.79544412186503E-84

		0.307						0.9030794656								-0.0969315694		0.000011035		-0		1.85350591235239E-21		-2.73470959531259E-33		9.47761255418628E-48		-7.71537757638442E-65		1.47531893282105E-84

		0.31						0.9059033658								-0.0941064379		0.0000098036		-0		1.15467409281701E-21		-1.30545971316107E-33		3.26773029418915E-48		-1.81095480936177E-65		2.22201865118274E-85

		0.314						0.9095411091								-0.0904672638		0.0000083729		-0		6.14346445444742E-22		-4.87042081181068E-34		7.90028106958202E-49		-2.62205331321029E-66		1.78058787752135E-86

		0.316						0.9113069422								-0.0887007957		0.0000077378		-0		4.48115858819661E-22		-2.97487000533587E-34		3.88454996221586E-49		-9.97719522416008E-67		5.04047271851903E-87

		0.318						0.9130383313								-0.0869688196		0.000007151		-0		3.26864140607686E-22		-1.81706096672105E-34		1.91002424800433E-49		-3.79643022662737E-67		1.42685264495349E-87

		0.32						0.9147359463								-0.0852706623		0.0000066086		-0		2.38420855482818E-22		-1.10986717095508E-34		9.39154512993694E-50		-1.44458258476778E-67		4.0391220905347E-88

		0.322						0.9164004441								-0.0836056632		0.0000061073		-0		1.73908658880344E-22		-6.77910735921349E-35		4.61780105775088E-50		-5.49679230129904E-68		1.14339117777486E-88

		0.324						0.9180324691								-0.081973175		0.0000056441		-0		1.26852248610186E-22		-4.1407024002878E-35		2.27056211879252E-50		-2.09158866528054E-68		3.23670182804568E-89

		0.326						0.9196326532								-0.0803725628		0.000005216		-0		9.2528417394858E-23		-2.52915545649923E-35		1.11643015167192E-50		-7.95872011335801E-69		9.16242745904537E-90

		0.328						0.9212016161								-0.0788032043		0.0000048204		-0		6.74919689591503E-23		-1.54481696697045E-35		5.48946128029752E-51		-3.02837871012624E-69		2.59369201743732E-90

		0.333						0.9249913907								-0.075012567		0.0000039577		-0		3.0668639333885E-23		-4.50433617988253E-36		9.30625929508837E-52		-2.7047527143737E-70		1.105829180342E-91

		0.337						0.9278916114								-0.0721117688		0.0000033801		-0		1.63173051847323E-23		-1.68048178375189E-36		2.24994285080226E-52		-3.91616940382829E-71		8.86142891770177E-93

		0.339						0.9292994129								-0.0707037109		0.0000031238		-0		1.19021494804701E-23		-1.02644413001723E-36		1.10629170520564E-52		-1.49014463115696E-71		2.50848561144581E-93

		0.341						0.93067974								-0.0693231468		0.0000028868		-0		8.68165182005694E-24		-6.26955651786094E-37		5.43961077309325E-53		-5.67016079436012E-72		7.1009993097847E-94

		0.343						0.9320331282								-0.0679695397		0.0000026679		-0		6.33256021934291E-24		-3.82946697059814E-37		2.67464405848112E-53		-2.15755724388548E-72		2.0101447250678E-94

		0.345						0.9333601023								-0.0666423632		0.0000024655		-0		4.61908859774355E-24		-2.33905177138517E-37		1.31511630849655E-53		-8.20973765906789E-73		5.69030025133258E-95

		0.347						0.9346611773								-0.0653411012		0.0000022785		-0		3.36925015077365E-24		-1.42870097358888E-37		6.46639652625676E-54		-3.12389358946218E-73		1.61080525926921E-95

		0.349						0.935936858								-0.0640652477		0.0000021057		-0		2.45759446658669E-24		-8.72655533710159E-38		3.17951224272992E-54		-1.18867515181857E-73		4.55985355549845E-96

		0.351						0.9371876394								-0.0628143066		0.000001946		-0		1.79261565390468E-24		-5.33021041206414E-38		1.56335882908212E-54		-4.52303696040469E-74		1.29079938918411E-96

		0.353						0.938414007								-0.0615877914		0.0000017984		-0		1.30756759356121E-24		-3.25571109554361E-38		7.68699926870092E-55		-1.72106427175566E-74		3.65398371425536E-97

		0.355						0.9396164368								-0.0603852252		0.000001662		-0		9.5376441012735E-25		-1.98859968335493E-38		3.77967979313492E-55		-6.54883489443966E-75		1.03436654029428E-97

		0.357						0.9407953955								-0.0592061404		0.0000015359		-0		6.95693709835727E-25		-1.21464361688978E-38		1.85845982798522E-55		-2.49190220135604E-75		2.92807582996685E-98

		0.359						0.941951341								-0.0580500784		0.0000014195		-0		5.07452084357362E-25		-7.41908554245588E-39		9.13800406719156E-56		-9.48195622765694E-76		8.288771660766E-99

		0.362						0.9436430886								-0.0563581724		0.0000012611		-0		3.16126196980826E-25		-3.54162551693717E-39		3.15063868121472E-56		-2.2256064674258E-76		1.2483948260868E-99

		0.365						0.9452855423								-0.054715578		0.0000011204		-0		1.9693637192193E-25		-1.69065462723702E-39		1.08629018181395E-56		-5.22394749450509E-77		1.88024197744188E-100

		0.369						0.9474012756								-0.0525996812		0.0000009568		-0		1.04780353886572E-25		-6.30750945361673E-40		2.6262870515105E-57		-7.56367235957202E-78		1.50670925739433E-101

		0.371						0.948428267								-0.0515726173		0.0000008843		-0		7.64287620079215E-26		-3.85264875602421E-40		1.29133927981287E-57		-2.87805878556441E-78		4.26517949634018E-102

		0.373						0.9494352093								-0.0505656079		0.0000008172		-0		5.57485773371888E-26		-2.35321128671222E-40		6.34948542516887E-58		-1.09513236155478E-78		1.20738331212365E-102

		0.375						0.9504224938								-0.0495782614		0.0000007552		-0		4.06640614537025E-26		-1.43734965489674E-40		3.12202732424243E-58		-4.16709656988233E-79		3.41785020688001E-103

		0.377						0.9513905041								-0.0486101939		0.0000006979		-0		2.96611316896769E-26		-8.77938178393771E-41		1.53509362737959E-58		-1.58562512006056E-79		9.67522071853301E-104

		0.379						0.9523396161								-0.0476610289		0.000000645		-0		2.16353876533955E-26		-5.36247698989259E-41		7.54801992449948E-59		-6.03347433687659E-80		2.73885308852615E-104

		0.381						0.9532701987								-0.0467303973		0.0000005961		-0		1.57812589152021E-26		-3.27541963373069E-41		3.71134397045835E-59		-2.29580195931539E-80		7.75312156565084E-105

		0.383						0.9541826135								-0.0458179373		0.0000005509		-0		1.15111472435097E-26		-2.00063772716409E-41		1.8248592617448E-59		-8.73577368877135E-81		2.19474692759476E-105

		0.385						0.955077215								-0.0449232941		0.0000005091		-0		8.3964474300664E-27		-1.22199649600116E-41		8.97279085873733E-60		-3.32405596361573E-81		6.21287056497002E-106

		0.387						0.9559543508								-0.0440461197		0.0000004705		-0		6.12452676996366E-27		-7.463997184317E-42		4.41190054939685E-60		-1.26483909072095E-81		1.75873401036595E-106

		0.389						0.9568143618								-0.043186073		0.0000004348		-0		4.46734508474197E-27		-4.5590354923112E-42		2.16932131420568E-60		-4.81284894997848E-82		4.97860898094026E-107

		0.391						0.9576575821								-0.0423428197		0.0000004018		-0		3.25856557669788E-27		-2.78467476700361E-42		1.06665028179528E-60		-1.83134085475693E-82		1.40934031178151E-107

		0.393						0.9584843396								-0.0415160317		0.0000003713		-0		2.37685905526003E-27		-1.70088905231475E-42		5.24469480940212E-61		-6.9684491683805E-83		3.98954832969676E-108

		0.395						0.9592949555								-0.0407053877		0.0000003432		-0		1.73372572550668E-27		-1.03890896077502E-42		2.5788043291445E-61		-2.65156995139222E-83		1.12935788055811E-108

		0.397						0.9600897448								-0.0399105723		0.0000003171		-0		1.26461217152602E-27		-6.34569213853058E-43		1.26799213485078E-61		-1.00895092110716E-83		3.19697649201217E-109

		0.399						0.9608690165								-0.0391312766		0.0000002931		-8.40380910935804E-16		9.22431916907951E-28		-3.87597087303676E-43		6.23468805241525E-62		-3.83916690815001E-84		9.04997332238704E-110

		0.401						0.9616330735								-0.0383671974		0.0000002709		-7.03721340364037E-16		6.72839199628856E-28		-2.36745651706137E-43		3.06558172109679E-62		-1.46084435231598E-84		2.561858597977E-110

		0.403						0.9623822127								-0.0376180376		0.0000002503		-5.892848331506E-16		4.90781574508736E-28		-1.44605069123885E-43		1.50733945463108E-62		-5.55867007804022E-85		7.25208709708924E-111

		0.405						0.9631167253								-0.036883506		0.0000002313		-4.93457558643448E-16		3.57985316566187E-28		-8.83252801714757E-44		7.41155329786705E-63		-2.11513382568873E-85		2.05291452484147E-111

		0.407						0.9638368969								-0.0361633169		0.0000002138		-4.13213353685826E-16		2.6112122690277E-28		-5.39493889435251E-44		3.64424364520919E-63		-8.04831198427602E-86		5.81137262953823E-112

		0.409						0.9645430074								-0.0354571902		0.0000001976		-3.46018158347152E-16		1.90466736997021E-28		-3.29524747810498E-44		1.79186618673706E-63		-3.06246938182018E-86		1.64507832307111E-112

		0.411						0.9652353312								-0.0347648514		0.0000001826		-2.89749991954485E-16		1.38930022398372E-28		-2.01274864360821E-44		8.81056467064849E-64		-1.16530257933708E-86		4.65687344721773E-113

		0.413						0.9659141375								-0.0340860312		0.0000001687		-2.42631942319612E-16		1.01338172890074E-28		-1.22939388597194E-44		4.33213430725183E-64		-4.43410180513387E-87		1.31826369597505E-113

		0.415						0.9665796902								-0.0334204658		0.0000001559		-2.03176052004981E-16		7.39179704099634E-29		-7.50918070005864E-45		2.13010043710246E-64		-1.68722348743761E-87		3.73172943568836E-114

		0.419						0.9678720646								-0.0321280686		0.0000001332		-1.42469402086525E-16		3.93281902299055E-29		-2.80153187359983E-45		5.14987182065608E-65		-2.44290465587115E-88		2.9903764272364E-115

		0.421						0.9684993887								-0.0315007344		0.0000001231		-1.193015658644E-16		2.86867220790037E-29		-1.71118543177096E-45		2.5321800844998E-65		-9.29551529056977E-89		8.46513165110926E-116

		0.423						0.9691144642								-0.0308856495		0.0000001137		-9.99011956901014E-17		2.09246349457551E-29		-1.0451980252299E-45		1.24506710140222E-65		-3.53704367092474E-89		2.39630212497468E-116

		0.425						0.9697175303								-0.0302825748		0.0000001051		-8.36556404603752E-17		1.52628225144477E-29		-6.38410596339571E-46		6.12196619222836E-66		-1.34588320700426E-89		6.7834312693838E-117

		0.427						0.9703088213								-0.0296912758		0.0000000971		-7.00518760810881E-17		1.11329899762381E-29		-3.89943417113713E-46		3.01015664268832E-66		-5.12123054003014E-90		1.92024783965587E-117

		0.429						0.9708885672								-0.0291115225		0.0000000898		-5.86603044991875E-17		8.12061240269899E-30		-2.38178798131099E-46		1.48008707153979E-66		-1.94868337071498E-90		5.43582092789151E-118

		0.431						0.9714569934								-0.0285430896		0.000000083		-4.91211872749022E-17		5.92332751000569E-30		-1.45480439954784E-46		7.27755395939374E-67		-7.41495007814809E-91		1.53876747312986E-118

		0.433						0.9720143209								-0.0279857558		0.0000000767		-4.11332852752145E-17		4.32058655811832E-30		-8.88599597256679E-47		3.57835647985154E-67		-2.8214683559011E-91		4.35592961536521E-119

		0.435						0.9725607663								-0.0274393046		0.0000000709		-3.44443457374787E-17		3.15151714549966E-30		-5.42759730785904E-47		1.75946961964701E-67		-1.07359909364882E-91		1.23307277709882E-119

		0.437						0.9730965421								-0.0269035234		0.0000000655		-2.8843136290839E-17		2.29877591497759E-30		-3.31519535088977E-47		8.65127149821957E-68		-4.08516016659505E-92		3.49057172149627E-120

		0.439						0.9736218566								-0.0263782039		0.0000000605		-2.41527743750027E-17		1.67677041352198E-30		-2.02493287382377E-47		4.25381022213499E-68		-1.5544474362414E-92		9.88108015130805E-121

		0.441						0.9741369141								-0.0258631418		0.0000000559		-2.02251414037479E-17		1.22306789511063E-30		-1.23683605624982E-47		2.09158866528054E-68		-5.91483989243805E-93		2.7971275981896E-121

		0.443						0.9746419148								-0.0253581369		0.0000000517		-1.69362052760677E-17		8.9212873986026E-31		-7.55463773547649E-48		1.02842931778329E-68		-2.2506602756391E-93		7.91808454212138E-122

		0.445						0.9751370551								-0.0248629927		0.0000000478		-1.41821035228932E-17		6.50735492008541E-31		-4.61439905684295E-48		5.05676321177799E-69		-8.56400472110143E-94		2.24144450388192E-122

		0.447						0.9756225275								-0.0243775167		0.0000000441		-1.18758633977043E-17		4.74658714196257E-31		-2.81849102516233E-48		2.48639879647797E-69		-3.2586960216474E-94		6.34506165885978E-123

		0.449						0.9760985207								-0.0239015201		0.0000000408		-9.94465533362296E-18		3.46224998834837E-31		-1.7215441406482E-48		1.22255654777903E-69		-1.23996892894456E-94		1.79615455055912E-123

		0.451						0.9765652198								-0.0234348179		0.0000000377		-8.3274930329422E-18		2.52543029829679E-31		-1.05152516071236E-48		6.01128231977335E-70		-4.71821530616606E-95		5.08453872152597E-124

		0.453						0.9770228064								-0.0229772285		0.0000000348		-6.97330755941205E-18		1.84209638617052E-31		-6.42275232742419E-49		2.95573363814252E-70		-1.79533173418211E-95		1.43932681086098E-124

		0.455						0.9774714582								-0.022528574		0.0000000322		-5.83933461436631E-18		1.343659770864E-31		-3.92303950496877E-49		1.45332740585321E-70		-6.83143058679226E-96		4.07443385118226E-125

		0.457						0.9779113498								-0.02208868		0.0000000298		-4.88976406791545E-18		9.80090723478123E-32		-2.39620620148026E-49		7.14597730102399E-71		-2.5994329055527E-96		1.15338720034885E-125

		0.459						0.9783426521								-0.0216573754		0.0000000275		-4.09460909827857E-18		7.14896618234094E-32		-1.46361109867492E-49		3.51366054070747E-71		-9.89112213704412E-97		3.26499848204079E-126

		0.461						0.978765533								-0.0212344924		0.0000000254		-3.42875922740643E-18		5.21459047126622E-32		-8.93978759774964E-50		1.727658775736E-71		-3.76367849006369E-97		9.24252938172399E-127

		0.463						0.9791801568								-0.0208198667		0.0000000235		-2.87118734837652E-18		3.80362042419347E-32		-5.46045341998526E-50		8.49485831313891E-72		-1.4321201963038E-97		2.6163672002272E-127

		0.465						0.9795866848								-0.020413337		0.0000000217		-2.40428570299846E-18		2.77443231852271E-32		-3.33526397868051E-50		4.17690222014826E-72		-5.44937157112606E-98		7.40638957552059E-128

		0.467						0.9799852749								-0.0200147452		0.0000000201		-2.01330984023436E-18		2.02372314574359E-32		-2.03719086161779E-50		2.05377317826427E-72		-2.07354456677864E-98		2.09659433659082E-128

		0.469						0.9803760823								-0.0196239363		0.0000000185		-1.685912996001E-18		1.47614174736801E-32		-1.24432327791366E-50		1.00983553012356E-72		-7.89006037540004E-99		5.93502106175122E-129

		0.472						0.9809480322								-0.0190519843		0.0000000165		-1.29187939625665E-18		9.19588452161035E-33		-5.93998687191189E-51		3.48175253532154E-73		-1.851956387296E-99		8.93889938032323E-130

		0.475						0.9815033124								-0.0184967023		0.0000000146		-9.89939800233587E-19		5.72873792679963E-33		-2.83555283942324E-51		1.20045298027237E-73		-4.3469153558569E-100		1.3463123601447E-130

		0.479						0.9822185931								-0.0177814194		0.0000000125		-6.94157259426796E-19		3.04798540480596E-33		-1.05789059768648E-51		2.90229459017294E-74		-6.29383116141174E-101		1.07885119081009E-131

		0.481						0.982565793								-0.0174342186		0.0000000115		-5.81276026942697E-19		2.22325791493046E-33		-6.46163263829835E-52		1.42705154934341E-74		-2.3948705348185E-101		3.05400259277843E-132

		0.483						0.9829062135								-0.0170937972		0.0000000107		-4.86751114260905E-19		1.6216861630988E-33		-3.94678773434911E-52		7.01677952120682E-75		-9.11274028719789E-102		8.64524404861981E-133

		0.485						0.983239987								-0.0167600229		0.0000000099		-4.07597485966143E-19		1.18288840621012E-33		-2.41071170274867E-52		3.45013429065539E-75		-3.46749581384847E-102		2.44728818622908E-133

		0.487						0.9835672433								-0.0164327659		0.0000000091		-3.41315521831287E-19		8.62821064510161E-34		-1.47247110940153E-52		1.69642306525106E-75		-1.31941949842987E-102		6.92776216931965E-134

		0.489						0.9838881095								-0.0161118989		0.0000000084		-2.85812080432311E-19		6.29357922061002E-34		-8.99390485204048E-53		8.34127304583567E-76		-5.02053327904325E-103		1.96110490561425E-134

		0.491						0.9842027106								-0.0157972972		0.0000000078		-2.39334399100153E-19		4.59065512367631E-34		-5.49350842750546E-53		4.10138469880387E-76		-1.91036697850652E-103		5.55147875580431E-135

		0.493						0.9845111688								-0.0154888384		0.0000000072		-2.00414742812793E-19		3.3485102396936E-34		-3.35545409247094E-53		2.01664138736952E-76		-7.26915207952516E-104		1.57150779073155E-135

		0.495						0.9848136041								-0.0151864026		0.0000000066		-1.67824054075529E-19		2.44246638513626E-34		-2.0495230534838E-53		9.91577914268282E-77		-2.76599064733483E-104		4.44861062243614E-136

		0.497						0.985110134								-0.0148898721		0.0000000061		-1.40533140082687E-19		1.78158094659625E-34		-1.25185582368326E-53		4.87556571149775E-77		-1.05248922810314E-104		1.25930883618713E-136

		0.499						0.9854008739								-0.0145991318		0.0000000057		-1.17680171476563E-19		1.29951867038601E-34		-7.64637899840084E-54		2.39730440392823E-77		-4.00483485488463E-105		3.56484052998675E-137

		0.501						0.9856859368								-0.0143140684		0.0000000052		-9.85434663354487E-20		9.47893374089007E-35		-4.67043493995675E-54		1.1787490406581E-77		-1.52388279011702E-105		1.00913196501601E-137

		0.503						0.9859654336								-0.0140345712		0.0000000048		-8.25186999267737E-20		6.91411265661124E-35		-2.85271793785404E-54		5.79588181866135E-78		-5.79853812244567E-106		2.85664201315878E-138

		0.505						0.986239473								-0.0137605315		0.0000000045		-6.90998205241273E-20		5.04328388984212E-35		-1.74245005820155E-54		2.84982170904957E-78		-2.20640619971001E-106		8.08655743177687E-139

		0.507						0.9865081614								-0.0134918427		0.0000000041		-5.78630686220666E-20		3.6786661798489E-35		-1.06429456801135E-54		1.40125075484135E-78		-8.39561319649558E-107		2.28913566334891E-139

		0.509						0.9867716035								-0.0132284003		0.0000000038		-4.84535948858639E-20		2.68328834115816E-35		-6.50074831222187E-55		6.88991761031334E-79		-3.19462123313635E-107		6.48006537939646E-140

		0.511						0.9870299016								-0.012970102		0.0000000035		-4.05742542397432E-20		1.95724101339661E-35		-3.97067972430025E-55		3.38775658196038E-79		-1.21558778189844E-107		1.83437128666377E-140

		0.513						0.9872831561								-0.0127168471		0.0000000033		-3.39762222181701E-20		1.42764842889316E-35		-2.42530501347476E-55		1.66575208989961E-79		-4.62544241606404E-108		5.19272232659155E-141

		0.515						0.9875314657								-0.0124685374		0.000000003		-2.84511372506693E-20		1.04135363124445E-35		-1.48138475445091E-55		8.19046456814598E-80		-1.76003065043243E-108		1.46995133194236E-141

		0.517						0.9877749267								-0.0122250761		0.0000000028		-2.38245207386101E-20		7.59582936078127E-36		-9.04834970664371E-56		4.02723251850136E-80		-6.6971061615716E-109		4.16112547981634E-142

		0.519						0.9880136339								-0.0119863687		0.0000000026		-1.99502671342641E-20		5.54054088323077E-36		-5.52676353443816E-56		1.98018093151296E-80		-2.54832101522432E-109		1.17792779138457E-142

		0.521						0.9882476801								-0.0117523223		0.0000000024		-1.67060300223993E-20		4.04137478880835E-36		-3.37576642767993E-56		9.73650392301309E-81		-9.69663589013494E-110		3.3344677742748E-143

		0.523						0.9884771564								-0.0115228458		0.0000000022		-1.39893584998658E-20		2.94785482642114E-36		-2.06192989861824E-56		4.78741649988621E-81		-3.68967437870378E-110		9.43918245158922E-144

		0.525						0.9887021518								-0.0112978502		0.000000002		-1.17144618425425E-20		2.15022078667859E-36		-1.25943396792938E-56		2.35396164009248E-81		-1.40396083498535E-110		2.67203558066392E-144

		0.529						0.9891390488								-0.010860953		0.0000000017		-8.21431639212835E-21		1.14402886964811E-36		-4.69870754850913E-57		5.69109348369874E-82		-2.03277306532468E-111		2.1412035226183E-145

		0.531						0.9893511201								-0.0106488815		0.0000000016		-6.87853527658767E-21		8.34476187236214E-37		-2.8699869456883E-57		2.79829752667751E-82		-7.7349204217085E-112		6.06130035860864E-146

		0.533						0.9895590506								-0.0104409509		0.0000000015		-5.75997384232751E-21		6.08682635149302E-37		-1.75299802836947E-57		1.37591643332497E-82		-2.94322051736767E-112		1.71582764782414E-146

		0.535						0.989762921								-0.0102370804		0.0000000014		-4.82330864497007E-21		4.43984569001755E-37		-1.07073730494971E-57		6.76534933632099E-83		-1.11992710222882E-112		4.85715002203511E-147

		0.537						0.9899628106								-0.0100371907		0.0000000013		-4.03896040528587E-21		3.23850700066912E-37		-6.54010077397174E-58		3.32650664923414E-83		-4.26144322827834E-113		1.37495781505052E-147

		0.539						0.9901587972								-0.009841204		0.0000000012		-3.38215991474627E-21		2.36222795241819E-37		-3.99471634508087E-58		1.6356356393883E-83		-1.62152503959387E-113		3.89221865619131E-148

		0.541						0.9903509569								-0.0096490442		0.0000000011		-2.83216583998846E-21		1.72305352374811E-37		-2.43998666521551E-58		8.04238267629255E-84		-6.17007739674227E-114		1.10180588100785E-148

		0.543						0.9905393645								-0.0094606365		0.000000001		-2.37160972496453E-21		1.25682766672095E-37		-1.49035235849994E-58		3.95442099415966E-84		-2.34778089466474E-114		3.11898253067675E-149

		0.545						0.9907240933								-0.0092759076		0.0000000009		-1.98594750636824E-21		9.16753752604923E-38		-9.10312414469722E-59		1.94437967309706E-84		-8.93355913535667E-115		8.82918869317316E-150

		0.547						0.9909052151								-0.0090947858		0.0000000009		-1.6630002215517E-21		6.68697439727684E-38		-5.56021995209098E-59		9.56046996194017E-85		-3.3993154559815E-115		2.49935907665205E-150

		0.549						0.9910828002								-0.0089172006		0.0000000008		-1.3925694047868E-21		4.87760496892193E-38		-3.39620172417829E-59		4.70086100764336E-85		-1.29347613803122E-115		7.07516399425589E-151

		0.551						0.9912569179								-0.0087430829		0.0000000007		-1.16611502633401E-21		3.55781685698399E-38		-2.07441184893672E-59		2.31140250439076E-85		-4.92181599890086E-116		2.00283128635797E-151

		0.553						0.9914276357								-0.008572365		0.0000000007		-9.76485803844138E-22		2.5951385707722E-38		-1.26705798668367E-59		1.13651127498074E-85		-1.87280399033176E-116		5.66959743247082E-152

		0.555						0.99159502								-0.0084049806		0.0000000006		-8.17693369501273E-22		1.89294291196841E-38		-7.73923433980488E-60		5.58819969998594E-86		-7.12622086438344E-117		1.60494472326385E-152

		0.557						0.991759136								-0.0082408646		0.0000000006		-6.84723161252504E-22		1.38074818367222E-38		-4.72715130608845E-60		2.74770489078104E-86		-2.71160378075555E-117		4.54326360101005E-153

		0.559						0.9919200475								-0.008079953		0.0000000005		-5.7337606619163E-22		1.00714370975492E-38		-2.88736049194454E-60		1.35104015105989E-86		-1.03179443968075E-117		1.28610311925799E-153

		0.561						0.992077817								-0.0079221835		0.0000000005		-4.80135815297994E-22		7.34629575540108E-39		-1.76360985096982E-60		6.64303323075257E-87		-3.92608895632761E-118		3.64068955408486E-154

		0.563						0.9922325059								-0.0077674945		0.0000000005		-4.02057941942112E-22		5.35852637544223E-39		-1.07721904317638E-60		3.26636410252236E-87		-1.49391912770616E-118		1.03060324096489E-154

		0.565						0.9923841744								-0.0076158261		0.0000000004		-3.36676797539873E-22		3.90860998145888E-39		-6.57969145694969E-61		1.60606368802365E-87		-5.68452316020317E-119		2.9174227148688E-155

		0.567						0.9925328813								-0.0074671191		0.0000000004		-2.81927687969973E-22		2.85101367741223E-39		-4.01889847221786E-61		7.89697807417211E-88		-2.16302227875618E-119		8.25861491495389E-156

		0.569						0.9926786846								-0.0073213158		0.0000000004		-2.36081671873099E-22		2.07958303011797E-39		-2.45475719274583E-61		3.88292588699851E-88		-8.23053270527732E-120		2.33784154644062E-156

		0.571						0.9928216409								-0.0071783594		0.0000000003		-1.97690961805546E-22		1.51688699826931E-39		-1.49937424819094E-61		1.90922569396956E-88		-3.13180632848543E-120		6.61794157076773E-157

		0.573						0.9929618059								-0.0070381944		0.0000000003		-1.65543204051051E-22		1.10644592314642E-39		-9.15823016134393E-62		9.38761865818436E-89		-1.19168603422874E-120		1.87340115932064E-157

		0.575						0.993099234								-0.0069007663		0.0000000003		-1.38623193277009E-22		8.07062478776672E-40		-5.593878899104E-62		4.61587042065541E-89		-4.53449369221569E-121		5.30320775155584E-158

		0.577						0.9932339787								-0.0067660216		0.0000000003		-1.16080813008723E-22		5.88686560294687E-40		-3.41676072631572E-62		2.26961282898998E-89		-1.72542368158658E-121		1.50122745020403E-158

		0.579						0.9933660923								-0.0066339079		0.0000000002		-9.72041895026883E-23		4.29399055692549E-40		-2.08696935909067E-62		1.11596338806769E-89		-6.56542291830878E-122		4.24966164409629E-159

		0.582						0.9935594417								-0.0064405585		0.0000000002		-7.44855101931224E-23		2.67501690598333E-40		-9.96249995086961E-63		3.84766453528849E-90		-1.54103724564275E-122		6.40053294533499E-160

		0.59						0.9940479337								-0.0059520664		0.0000000002		-3.6624353394256E-23		7.57241465383643E-41		-1.38667395164973E-63		2.24900218162311E-91		-3.23058330857986E-124		4.1100543225874E-162

		0.594						0.994278105								-0.0057218951		0.0000000001		-2.56814210060412E-23		4.02891695150847E-41		-5.17341562150318E-64		5.43733654901731E-92		-4.67751134598962E-125		3.29354251768235E-163

		0.596						0.994389831								-0.0056101691		0.0000000001		-2.15052052916096E-23		2.93876784544804E-41		-3.15994029103698E-64		2.67352582775353E-92		-1.77984343899634E-125		9.32333159022157E-164

		0.598						0.9944993755								-0.0055006247		0.0000000001		-1.80081100078335E-23		2.14359257174705E-41		-1.93010254219974E-64		1.31456647702949E-92		-6.77249595568494E-126		2.63924061931928E-164

		0.6						0.9946067809								-0.0053932192		0.0000000001		-1.50796991545464E-23		1.56357676254233E-41		-1.17891335920889E-64		6.46369301762736E-93		-2.57700764375282E-126		7.47113945187833E-165

		0.602						0.9947120892								-0.0052879109		0.0000000001		-1.26274954169377E-23		1.14050231587145E-41		-7.20084388333659E-65		3.17818293377859E-93		-9.80579160093193E-127		2.11492367542485E-165

		0.604						0.9948153412								-0.0051846589		0.0000000001		-1.05740597919495E-23		8.31903852544564E-42		-4.39830053897964E-65		1.5627052109398E-93		-3.73120930215353E-127		5.98690759512965E-166

		0.606						0.9949165771								-0.005083423		0.0000000001		-8.85454611480168E-24		6.06806325815916E-42		-2.68649729734514E-65		7.68378544338526E-94		-1.41976532064511E-127		1.69476860886825E-166

		0.608						0.9950158362								-0.0049841639		0.0000000001		-7.41465325918072E-24		4.42615953663332E-42		-1.64092191169758E-65		3.77809956264716E-94		-5.40235993875553E-128		4.79753627722901E-167

		0.61						0.9951131572								-0.0048868428		0.0000000001		-6.2089103429002E-24		3.22852406282812E-42		-1.00228082229978E-65		1.8576828323027E-94		-2.05565613439618E-128		1.35808240788094E-167

		0.612						0.995208578								-0.0047914221		0.0000000001		-5.19924078694314E-24		2.354946209686E-42		-6.12196619222819E-66		9.13418359735955E-95		-7.82199296378946E-129		3.84444789995626E-168

		0.614						0.9953021355								-0.0046978646		0.0000000001		-4.35375988180016E-24		1.71774208356263E-42		-3.73931828534738E-66		4.49125698636347E-95		-2.97635255730878E-129		1.08828297677012E-168

		0.616						0.9953938662								-0.0046061338		0.0000000001		-3.64576788903006E-24		1.25295340229258E-42		-2.28398864026457E-66		2.20834069104876E-95		-1.13253420022342E-129		3.08070201066091E-169

		0.618						0.9954838058								-0.0045161943		0.0000000001		-3.05290688084224E-24		9.13927791220303E-43		-1.39506822120469E-66		1.08583601930342E-95		-4.30941459379861E-130		8.72082452916557E-170

		0.62						0.9955719892								-0.0044280108		0		-2.55645469124299E-24		6.66636130311393E-43		-8.52112531343287E-67		5.33903063778075E-96		-1.63977866077527E-130		2.46868344310198E-170

		0.622						0.9956584508								-0.0043415493		0		-2.14073368218006E-24		4.86256939011743E-43		-5.20473303768096E-67		2.62518903816134E-96		-6.23953439105931E-131		6.98832767688881E-171

		0.624						0.9957432241								-0.004256776		0		-1.79261565390468E-24		3.54684962284634E-43		-3.17906907798008E-67		1.29079938918411E-96		-2.37421002898057E-131		1.97824973696122E-171

		0.626						0.9958263421								-0.004173658		0		-1.50110726494085E-24		2.58713886379752E-43		-1.94178647192861E-67		6.34683079541195E-97		-9.03412483756701E-132		5.60001219566337E-172

		0.628						0.9959078371								-0.0040921629		0		-1.2570028694941E-24		1.88710777515295E-43		-1.18604994420593E-67		3.1207220489198E-97		-3.43758178865853E-132		1.58524659478799E-172

		0.63						0.9959877409								-0.0040122591		0		-1.05259380912972E-24		1.37649192506638E-43		-7.24443439320949E-68		1.53445182651695E-97		-1.30803688970274E-132		4.48750230978603E-173

		0.632						0.9960660845								-0.0039339156		0		-8.81424978340853E-25		1.00403911462841E-43		-4.42492577432352E-68		7.54486420447563E-98		-4.97722122705014E-133		1.27031826130673E-173

		0.634						0.9961428983								-0.0038571018		0		-7.38090976504517E-25		7.32365025428819E-44		-2.70276008388249E-68		3.70979231020825E-98		-1.89388627629813E-133		3.59600591511739E-174

		0.638						0.9962920556								-0.0037079444		0		-5.17557945783294E-25		3.89656140151704E-44		-1.00834815729381E-68		8.96903946220315E-99		-2.74212849483604E-134		2.88161601908475E-175

		0.64						0.9963644571								-0.0036355429		0		-4.33394626868791E-25		2.84222521645794E-44		-6.15902568581912E-69		4.41005599638212E-99		-1.04340942210719E-134		8.1572536311227E-176

		0.642						0.9964354449								-0.0035645551		0		-3.62917629086858E-25		2.07317256130606E-44		-3.76195435318556E-69		2.16841435174687E-99		-3.97028521527111E-135		2.30914828213646E-176

		0.644						0.9965050466								-0.0034949535		0		-3.03901334572617E-25		1.51221108167799E-44		-2.29781482938713E-69		1.06620433044826E-99		-1.51073628018102E-135		6.53671692706687E-177

		0.646						0.9965732892								-0.0034267108		0		-2.54482047034737E-25		1.10303522158767E-44		-1.4035133057052E-69		5.2425020769245E-100		-5.74851423639903E-136		1.85040815763765E-177

		0.648						0.9966401993								-0.0033598007		0		-2.13099137435726E-25		8.04574648873016E-45		-8.57270818387458E-70		2.57772616764778E-100		-2.18737157236501E-136		5.23811936183714E-178

		0.65						0.9967058029								-0.0032941971		0		-1.78445760339438E-25		5.86871890343977E-45		-5.23624003471361E-70		1.26746200533204E-100		-8.32318438961968E-137		1.48280228530141E-178

		0.652						0.9967701256								-0.0032298744		0		-1.49427584580085E-25		4.28075401279853E-45		-3.19831366156992E-70		6.2320814178109E-101		-3.16706129213847E-137		4.19750384711343E-179

		0.654						0.9968331923								-0.0031668077		0		-1.25128234994013E-25		3.12246253732653E-45		-1.95354113065292E-70		3.06430004487988E-101		-1.20510092755738E-137		1.18822574804372E-179

		0.656						0.9968950275								-0.0031049725		0		-1.04780353886571E-25		2.27758293699146E-45		-1.19322973071989E-70		1.50670925739428E-101		-4.58553880597293E-138		3.36361914065903E-180

		0.658						0.9969556554								-0.0030443447		0		-8.7741368373976E-26		1.66131185654387E-45		-7.28828877945324E-71		7.40845463260333E-102		-1.74484689707316E-138		9.52170388668556E-181

		0.66						0.9970150994								-0.0029849006		0		-7.34731983485349E-26		1.21179213273303E-45		-4.45171218627433E-71		3.64272003865293E-102		-6.63933034491002E-139		2.69539567693028E-181

		0.662						0.9970733827								-0.0029266173		0		-6.1525264257952E-26		8.8390398658116E-46		-2.71912131765312E-71		1.79111703291096E-102		-2.52633669480011E-139		7.63010270186359E-182

		0.664						0.997130528								-0.0028694721		0		-5.15202580953969E-26		6.4473620218344E-46		-1.66084877699679E-71		8.80688109858156E-103		-9.61298318344243E-140		2.15992285434284E-182

		0.666						0.9971865574								-0.0028134426		0		-4.31422282574472E-26		4.70282719295958E-46		-1.01445222106994E-71		4.33032310337094E-103		-3.65784362295594E-140		6.11429088047933E-183

		0.668						0.9972414928								-0.0027585072		0		-3.61266019974375E-26		3.43033065801809E-46		-6.19630952009116E-72		2.12920987233582E-103		-1.39184888963978E-140		1.73082815879028E-183

		0.67						0.9972953556								-0.0027046444		0		-3.02518303897754E-26		2.50214773805741E-46		-3.78472744909358E-72		1.04692757843481E-103		-5.29613491247605E-141		4.89961333836039E-184

		0.672						0.9973481666								-0.0026518334		0		-2.53323919586083E-26		1.82511364857261E-46		-2.31172471573236E-72		5.14771873232452E-104		-2.01523636796559E-141		1.38697829380237E-184

		0.674						0.9973999464								-0.0026000536		0		-2.1212934029984E-26		1.33127224245845E-46		-1.41200951276097E-72		2.5311214159381E-104		-7.66819139974043E-142		3.92624612317406E-185

		0.676						0.9974507152								-0.0025492848		0		-1.77633667951966E-26		9.71055027135668E-47		-8.62460331266516E-73		1.24454655651462E-104		-2.91782940590803E-142		1.11143834684523E-185

		0.678						0.9975004926								-0.0024995074		0		-1.48747551590313E-26		7.08305811277277E-47		-5.26793776023412E-73		6.11940668503406E-105		-1.11026551088296E-142		3.14625003141532E-186

		0.68						0.9975492981								-0.0024507019		0		-1.24558786401325E-26		5.16651588498566E-47		-3.21767474278485E-73		3.00889813891012E-105		-4.22467983275621E-143		8.90637729774076E-187

		0.682						0.9975971507								-0.0024028493		0		-1.04303506873865E-26		3.7685539162349E-47		-1.96536694653263E-73		1.47946826813761E-105		-1.60753617169493E-143		2.52120955987814E-187

		0.684						0.9976440688								-0.0023559312		0		-8.73420644219663E-27		2.74885414769393E-47		-1.20045298027234E-73		7.27451131735177E-106		-6.11684824793339E-144		7.13701815263691E-188

		0.686						0.9976900708								-0.0023099292		0		-7.31388276974839E-27		2.00506594658021E-47		-7.33240863945103E-74		3.57686041978417E-106		-2.32752663032136E-144		2.02034090785896E-188

		0.688						0.9977351746								-0.0022648254		0		-6.1245267699637E-27		1.46253283518457E-47		-4.47866075051928E-74		1.758734010366E-106		-8.85648947836068E-145		5.71916351713426E-189

		0.691						0.9978011844								-0.0021988156		0		-4.69309505573899E-27		9.11110541071075E-48		-2.13796418776615E-74		6.06383556221157E-107		-2.07879679978219E-145		8.61379036188189E-190

		0.694						0.9978652702								-0.0021347298		0		-3.5962192720291E-27		5.67592328924415E-48		-1.02059323596691E-74		2.09071420173927E-107		-4.87935557914146E-146		1.29734679164467E-190

		0.698						0.997947822								-0.002052178		0		-2.52171062682851E-27		3.01988528109875E-48		-3.80763840257481E-75		5.05464905083473E-108		-7.06474308275456E-147		1.03961322237959E-191

		0.7						0.9979878929								-0.0020121071		0		-2.11163956633179E-27		2.20276115587641E-48		-2.32571880596367E-75		2.48535926842329E-108		-2.68821082279496E-147		2.9429280921032E-192

		0.702						0.9980271814								-0.0019728186		0		-1.7682527133202E-27		1.60673544131205E-48		-1.42055715184385E-75		1.22204541423451E-108		-1.02289316725934E-147		8.33081531559037E-193

		0.704						0.9980657028								-0.0019342972		0		-1.48070613376306E-27		1.17198306837815E-48		-8.67681259006976E-76		6.00876908793569E-109		-3.89221865619131E-148		2.35827997322476E-193

		0.706						0.998103472								-0.001896528		0		-1.23991929323653E-27		8.54866505865724E-49		-5.29982736882288E-76		2.95449788785042E-109		-1.48103111375693E-148		6.67579849202171E-194

		0.708						0.9981405037								-0.0018594963		0		-1.03828829953789E-27		6.23555717287272E-49		-3.23715302684658E-76		1.45271979028757E-109		-5.6354828792236E-149		1.88977924640298E-194

		0.71						0.9981768123								-0.0018231877		0		-8.69445776703173E-28		4.54833275012786E-49		-1.97726435031967E-76		7.14298966931622E-110		-2.14436192373153E-149		5.34957069840181E-195

		0.712						0.9982124119								-0.0017875881		0		-7.28059787405313E-28		3.3176395039539E-49		-1.20771995596804E-76		3.51219152909433E-110		-8.15952804488573E-150		1.51435183298106E-195

		0.714						0.9982473165								-0.0017526835		0		-6.0966545383271E-28		2.41994868952499E-49		-7.37679558025546E-77		1.72693646611743E-110		-3.10478829056155E-150		4.28681403301803E-196

		0.716						0.9982815395								-0.0017184605		0		-5.10523959744696E-28		1.76515611565224E-49		-4.50577244865159E-77		8.49130673342043E-111		-1.18140537984297E-150		1.21350759800019E-196

		0.718						0.9983150942								-0.0016849058		0		-4.27504481080411E-28		1.28753808959316E-49		-2.75214151431389E-77		4.17515591660047E-111		-4.49537469516011E-151		3.43518678221587E-197

		0.72						0.9983479937								-0.0016520063		0		-3.5798531656619E-28		9.39154512993694E-50		-1.68101762819307E-77		2.05291452484152E-111		-1.71053848193683E-151		9.72429694561293E-198

		0.722						0.9983802509								-0.0016197491		0		-2.99771096090313E-28		6.85036975919776E-50		-1.02677142566931E-77		1.00941333221798E-111		-6.50878313066263E-152		2.75274554431825E-198

		0.724						0.9984118782								-0.0015881218		0		-2.51023452339207E-28		4.99678862087801E-50		-6.27155565111024E-78		4.96326204978321E-112		-2.47666207392377E-152		7.79244820900145E-199

		0.726						0.9984428879								-0.0015571121		0		-2.10202966350396E-28		3.64475165566244E-50		-3.83068804815371E-78		2.44042449099505E-112		-9.42396590157642E-153		2.20587947968175E-199

		0.728						0.9984732922								-0.0015267078		0		-1.76020553660453E-28		2.65855044897214E-50		-2.33979760981158E-78		1.19995108791577E-112		-3.58592051169003E-153		6.24438449685193E-200

		0.73						0.9985031028								-0.0014968972		0		-1.47396755853984E-28		1.93919672928864E-50		-1.42915653429897E-78		5.90013179552693E-113		-1.3644813712673E-153		1.76765494686728E-200

		0.732						0.9985323312								-0.0014676688		0		-1.23427651967217E-28		1.41448658848571E-50		-8.7293379177953E-79		2.90108118198841E-113		-5.19199855787635E-154		5.00386229060627E-201

		0.734						0.998560989								-0.001439011		0		-1.03356313250421E-28		1.03175313715587E-50		-5.33191002204926E-79		1.42645491934052E-113		-1.97561136360218E-154		1.41648899677655E-201

		0.736						0.9985890872								-0.0014109128		0		-8.65488998490917E-29		7.52580155016242E-51		-3.25674922324564E-79		7.01384590525707E-114		-7.51741399094399E-155		4.00978476517173E-202

		0.738						0.9986166368								-0.0013833632		0		-7.24746445525675E-29		5.48946128029759E-51		-1.98923377537322E-79		3.44869183846586E-114		-2.86045697814768E-155		1.13508639315888E-202

		0.74						0.9986436484								-0.0013563516		0		-6.06890915098813E-29		4.00411636515126E-51		-1.21503092250436E-79		1.69571381483976E-114		-1.08843468427988E-155		3.21319271579202E-203

		0.742						0.9986701326								-0.0013298674		0		-5.08200606023422E-29		2.92067783103176E-51		-7.42145121864745E-80		8.33778567793806E-115		-4.14161118658224E-156		9.09587806800034E-204

		0.744						0.9986960996								-0.0013039004		0		-4.25558942368618E-29		2.13039737479211E-51		-4.53304826820704E-80		4.09966996805993E-115		-1.57592765726432E-156		2.57485327354653E-204

		0.748						0.9987465225								-0.0012534775		0		-2.98406861800547E-29		1.13348178739797E-51		-1.69119371542749E-80		9.91163349612841E-116		-2.28176115370096E-157		2.06332765156263E-205

		0.75						0.998770998								-0.001229002		0		-2.49881064678476E-29		8.26782946955246E-52		-1.03298701521589E-80		4.87352730665816E-116		-8.68234691136254E-158		5.84085001833485E-206

		0.752						0.9987949956								-0.0012050044		0		-2.09246349457551E-29		6.03071040907691E-52		-6.3095206886749E-81		2.39630212497468E-116		-3.30372649946191E-158		1.65342275672239E-206

		0.754						0.9988185245								-0.0011814755		0		-1.7521949819467E-29		4.39891366556176E-52		-3.85387722540687E-81		1.17825622241072E-116		-1.25710351068354E-158		4.68049479761652E-207

		0.756						0.9988415941								-0.0011584059		0		-1.46725965003369E-29		3.20865041172287E-52		-2.35396164009241E-81		5.79345864271667E-117		-4.78341423489614E-159		1.32495040735631E-207

		0.758						0.9988642132								-0.0011357868		0		-1.22865942591912E-29		2.3404499945635E-52		-1.43780797335638E-81		2.84863023903206E-117		-1.82014062868742E-159		3.75065812026461E-208

		0.76						0.9988863906								-0.0011136094		0		-1.0288594693279E-29		1.70716827144502E-52		-8.78218120906149E-82		1.40066491178449E-117		-6.92583110204E-160		1.0617330472901E-208

		0.762						0.998908135								-0.001091865		0		-8.61550227259949E-30		1.24524066474323E-52		-5.36418688851423E-82		6.88703703352876E-118		-2.63535331819819E-160		3.005544700588E-209

		0.764						0.9989294548								-0.0010705452		0		-7.21448182400019E-30		9.0830197530421E-53		-3.27646404576775E-82		3.38634020900581E-118		-1.00278031754083E-160		8.50806986774042E-210

		0.766						0.9989503583								-0.0010496417		0		-6.04129003068847E-30		6.62532554308807E-53		-2.00127565767615E-82		1.66505566258788E-118		-3.81568709707149E-161		2.40845703809332E-210

		0.768						0.9989708537								-0.0010291463		0		-5.05887825699163E-30		4.83263713449415E-53		-1.22238614618111E-82		8.18704025113269E-119		-1.45191003134793E-161		6.81783929200593E-211

		0.77						0.9989909488								-0.0010090512		0		-4.2362225764795E-30		3.52501647229343E-53		-7.46637718119558E-83		4.02554878973201E-119		-5.52467402462464E-162		1.92998803285354E-211

		0.772						0.9990106516								-0.0009893484		0		-3.54734405649566E-30		2.57121335290171E-53		-4.56048920269902E-83		1.97935304596431E-119		-2.10219796126264E-162		5.46339338230715E-212

		0.774						0.9990299697								-0.0009700303		0		-2.97048836031951E-30		1.87549140780007E-53		-2.78556269837461E-83		9.7324332288841E-120		-7.99908962707905E-163		1.54657265960895E-212

		0.776						0.9990489105								-0.0009510895		0		-2.48743875932647E-30		1.36801872810839E-53		-1.70143140389063E-83		4.78541494898102E-120		-3.04373974483395E-163		4.37802446954651E-213

		0.778						0.9990674815								-0.0009325185		0		-2.08294086051701E-30		9.97858605308413E-54		-1.03924023100766E-83		2.35297748213352E-120		-1.15817575076545E-163		1.23932736925495E-213

		0.78						0.9990856899								-0.0009143101		0		-1.74422088268259E-30		7.27856845618534E-54		-6.34771554865644E-84		1.156953595551E-120		-4.40698345493467E-164		3.50827716671829E-214

		0.782						0.9991035428								-0.0008964572		0		-1.46058226868279E-30		5.30912480882021E-54		-3.8772067789933E-84		5.68871411826979E-121		-1.67690466315092E-164		9.93120057206203E-215

		0.784						0.9991210471								-0.0008789529		0		-1.22306789511061E-30		3.87257554906647E-54		-2.36821141272674E-84		2.79712759818944E-121		-6.38080282817627E-165		2.81131564342104E-215

		0.786						0.9991382096								-0.0008617904		0		-1.02417721214662E-30		2.8247294842858E-54		-1.44651178414208E-84		1.37534118218848E-121		-2.42796419061515E-165		7.9582479374931E-216

		0.788						0.9991550369								-0.0008449631		0		-8.57629381061911E-31		2.06041084500391E-54		-8.83534438866951E-85		6.76252083976394E-122		-9.23866521134094E-166		2.25281392300523E-216

		0.79						0.9991715357								-0.0008284643		0		-7.18164929406122E-31		1.50290244564184E-54		-5.39665914389321E-85		3.32511588400721E-122		-3.51541160356308E-166		6.37724611189348E-217

		0.792						0.9991877123								-0.0008122877		0		-6.01379660279698E-31		1.09624532728177E-54		-3.29629821251948E-85		1.63495180333714E-122		-1.33765197241859E-166		1.80526529760647E-217

		0.794						0.9992035731								-0.0007964269		0		-5.03585570653247E-31		7.99621972185878E-55		-2.01339043585044E-85		8.03902027021719E-123		-5.08990980601497E-167		5.11032934523921E-218

		0.796						0.9992191242								-0.0007808758		0		-4.21694386624591E-31		5.83259314762937E-55		-1.22978589491016E-85		3.95276770685533E-123		-1.9367654941311E-167		1.44662759825043E-218

		0.798						0.9992343716								-0.0007656284		0		-3.53120037732645E-31		4.25440320665238E-55		-7.51157510431467E-86		1.94356675554647E-123		-7.36960127431704E-168		4.09510085679586E-219

		0.801						0.9992566863								-0.0007433137		0		-2.70588400608016E-31		2.65035527018331E-55		-3.58577696800959E-86		6.70110951420253E-124		-1.72979413368611E-168		6.16774466852062E-220

		0.804						0.9992783507								-0.0007216493		0		-2.07346156320469E-31		1.65108540892525E-55		-1.71173106648726E-86		2.31043613980253E-124		-4.06017589494644E-169		9.28941084147834E-221

		0.808						0.9993062574								-0.0006937426		0		-1.45393527556054E-31		8.7846298657681E-56		-6.38614162229127E-87		5.58586335298822E-125		-5.87866555395316E-170		7.44395746851495E-222

		0.81						0.9993198035								-0.0006801965		0		-1.21750181091176E-31		6.40767434384967E-56		-3.90067755868484E-87		2.7465561143562E-125		-2.23689555028632E-170		2.10722902315253E-222

		0.812						0.999333085								-0.000666915		0		-1.01951626354339E-31		4.67387825374693E-56		-2.38254744675832E-87		1.35047530034402E-125		-8.51162845881908E-171		5.96512564022242E-223

		0.814						0.9993461073								-0.0006538927		0		-8.53726378321437E-32		3.40921475696039E-56		-1.45526828368987E-87		6.64025587282378E-126		-3.23876628981205E-171		1.68860259196731E-223

		0.816						0.9993588752								-0.0006411248		0		-7.14896618234094E-32		2.4867454024415E-56		-8.88882939307265E-88		3.26499848204079E-126		-1.23238545135912E-171		4.78008157007132E-224

		0.818						0.9993713938								-0.0006286062		0		-5.98642829529773E-32		1.81387889511461E-56		-5.4293279709786E-88		1.60539221558581E-126		-4.68935935729305E-172		1.35314134451922E-224

		0.82						0.999383668								-0.000616332		0		-5.01293792985964E-32		1.32307740185707E-56		-3.31625244595463E-88		7.8936764596976E-127		-1.78435173488794E-172		3.83046078065187E-225

		0.822						0.9993957026								-0.0006042974		0		-4.1977528918812E-32		9.65077556180613E-57		-2.02557855117335E-88		3.88130249078397E-127		-6.78965050704824E-173		1.08432351517008E-225

		0.824						0.9994075021								-0.0005924979		0		-3.51513016675041E-32		7.03945731471393E-57		-1.23723043822514E-88		1.90842747379881E-127		-2.58353513528277E-173		3.06949360110848E-226

		0.826						0.9994190712								-0.0005809288		0		-2.94351297168938E-32		5.13471264234922E-57		-7.55704663432588E-89		9.38369382803417E-128		-9.83062940914573E-174		8.68909586062896E-227

		0.828						0.9994304145								-0.0005695855		0		-2.46485000653996E-32		3.74535603254405E-57		-4.61587042065527E-89		4.61394059073184E-128		-3.74066035565579E-174		2.45970171913457E-227

		0.83						0.9994415362								-0.0005584638		0		-2.06402540541654E-32		2.73193317476411E-57		-2.81938973930666E-89		2.26866393607204E-128		-1.42336154828056E-174		6.96290229059149E-228

		0.832						0.9994524408								-0.0005475592		0		-1.72838138746835E-32		1.99272347048591E-57		-1.72209307837957E-89		1.11549681961065E-128		-5.4160439721841E-175		1.97105234066254E-228

		0.834						0.9994631325								-0.0005368675		0		-1.44731853237253E-32		1.45353000084578E-57		-1.05186045379175E-89		5.48487210810073E-129		-2.0608630564783E-175		5.57963786865333E-229

		0.84						0.9994939711								-0.0005060289		0		-8.49841137834354E-33		5.64098431497736E-58		-2.39697026437298E-90		6.52020143099586E-130		-1.13540287910268E-176		1.26569650714642E-230

		0.843						0.9995087197								-0.0004912803		0		-6.51215251714488E-33		3.51415034777248E-58		-1.14423415164357E-90		2.24806190572437E-130		-2.66501967546969E-177		1.90630051784024E-231

		0.847						0.9995277179								-0.0004722821		0		-4.56639680837551E-33		1.8697100665396E-58		-4.26891904021687E-91		5.43506327576194E-131		-3.85864055453652E-178		1.52759095481595E-232

		0.849						0.9995369397								-0.0004630603		0		-3.82382660149401E-33		1.36380170899273E-58		-2.60747062700466E-91		2.67240806454241E-131		-1.46825428447657E-178		4.32429122426926E-233

		0.853						0.9995548466								-0.0004451534		0		-2.68131152378155E-33		7.25613172949168E-59		-9.72797481218137E-92		6.46099063929664E-132		-2.12586255126046E-179		3.46521867793844E-234

		0.856						0.9995678209								-0.0004321791		0		-2.054632190828E-33		4.52033482400819E-59		-4.64381272136383E-92		2.22764696507637E-132		-4.98982839547707E-180		5.21906169677021E-235

		0.86						0.9995845336								-0.0004154664		0		-1.44073190145369E-33		2.40505234215005E-59		-1.73251781700067E-92		5.38570676386595E-133		-7.22469495598466E-181		4.18223221679926E-236

		0.862						0.999592646								-0.000407354		0		-1.20644551964801E-33		1.75429043953951E-59		-1.05822792047187E-92		2.64813958159438E-133		-2.74907423307142E-181		1.18390266818123E-236

		0.864						0.9996006								-0.0003994		0		-1.0102579046179E-33		1.27961246095317E-59		-6.46369301762717E-93		1.30208411840333E-133		-1.04605235030406E-181		3.35138140368372E-237

		0.866						0.9996083987								-0.0003916013		0		-8.45973578766091E-34		9.33373410309637E-60		-3.94804622122366E-93		6.40231754844774E-134		-3.98034184167562E-182		9.48706140700787E-238

		0.868						0.9996160452								-0.0003839548		0		-7.08404549668916E-34		6.80820130826254E-60		-2.41148039091752E-93		3.1480047572829E-134		-1.51456293482707E-182		2.68558911383253E-238

		0.87						0.9996235423								-0.0003764577		0		-5.93206476641483E-34		4.96603015918917E-60		-1.47294062681393E-93		1.5478666712304E-134		-5.76307506941779E-183		7.60234236811097E-239

		0.872						0.999630893								-0.000369107		0		-4.96741479277442E-34		3.62231585485672E-60		-8.99677268075819E-94		7.61082468621793E-135		-2.19291212613341E-183		2.15206448314388E-239

		0.874						0.9996381002								-0.0003618998		0		-4.15963255545957E-34		2.64218535364053E-60		-5.49526010728058E-94		3.74222492679532E-135		-8.34426679336746E-184		6.09204547145472E-240

		0.876						0.9996451667								-0.0003548333		0		-3.48320881550044E-34		1.92725972077564E-60		-3.35652402458226E-94		1.84004335667957E-135		-3.17508337379967E-184		1.7245309476996E-240

		0.878						0.9996520952								-0.0003479048		0		-2.91678255004894E-34		1.40577950983133E-60		-2.0501765717461E-94		9.04745070297005E-136		-1.20815341601873E-184		4.8817872478287E-241

		0.88						0.9996588884								-0.0003411116		0		-2.44246638513626E-34		1.02540203012504E-60		-1.2522549949154E-94		4.44861062243614E-136		-4.59715385328934E-185		1.38193209955728E-241

		0.882						0.9996655489								-0.0003344511		0		-2.04528172400798E-34		7.47947538025174E-61		-7.64881715019806E-95		2.18737157236501E-136		-1.74926654765883E-185		3.91196140027695E-242

		0.884						0.9996720795								-0.0003279205		0		-1.71268573275685E-34		5.45567010014294E-61		-4.67192417157135E-95		1.07552555205887E-136		-6.65614759133768E-186		1.10739463987831E-242

		0.886						0.9996784824								-0.0003215176		0		-1.43417524576552E-34		3.97946844242328E-61		-2.85362756571427E-95		5.28833430838107E-137		-2.53273583816993E-186		3.13480314080908E-243

		0.888						0.9996847604								-0.0003152396		0		-1.20095508255082E-34		2.90269917233967E-61		-1.74300566206862E-95		2.60026177004032E-137		-9.63733261308461E-187		8.87397353910567E-244

		0.89						0.9996909158								-0.0003090842		0		-1.0056603016703E-34		2.11728340279799E-61		-1.06463393278966E-95		1.27854270899956E-137		-3.66710884315238E-187		2.51203673199123E-244

		0.892						0.999696951								-0.000303049		0		-8.421236206499E-35		1.54438635959324E-61		-6.50282116411478E-96		6.28656498191944E-138		-1.39537440570114E-187		7.11105179102197E-245

		0.894						0.9997028683								-0.0002971317		9.74329219716064E-16		-7.05180657204659E-35		1.12650447481222E-61		-3.9719458294603E-96		3.09108948756356E-138		-5.30954988075031E-188		2.01299037273662E-245

		0.896						0.9997086701								-0.0002913299		9.0043002471693E-16		-5.90506841396791E-35		8.21693563847732E-62		-2.42607835491891E-96		1.51988156451207E-138		-2.02034090785896E-188		5.69835568606991E-246

		0.898						0.9997143587								-0.0002856413		8.32135804823752E-16		-4.94480848522777E-35		5.99358749090913E-62		-1.48185711410008E-96		7.47322256258689E-139		-7.68761472373961E-189		1.61308558474712E-246

		0.9						0.9997199361								-0.0002800639		7.69021443823317E-16		-4.1407024002878E-35		4.3718355104264E-62		-9.05123489584277E-97		3.67456628029351E-139		-2.9252201898582E-189		4.56630868108172E-247





FOTXT

		



Time, d^2/D

Oxygen Flux, Dco/d




_940232509

_940244328.xls
Chart1

		



Frequency, Hz

Response Magnitude



FO134

		





FO134

		



Frequency, Hz

Response Magnitude




_940247611.xls
Chart2

		0

		0.01

		0.02

		0.03

		0.04

		0.05

		0.06

		0.07

		0.08

		0.095

		0.103

		0.112

		0.122

		0.131

		0.142

		0.152

		0.162

		0.172

		0.182

		0.192

		0.202

		0.212

		0.222

		0.232

		0.242

		0.252

		0.262

		0.272

		0.282

		0.292

		0.302

		0.312

		0.322

		0.332

		0.342

		0.352

		0.362

		0.372

		0.382

		0.392

		0.402

		0.412

		0.422

		0.432

		0.442

		0.452

		0.462

		0.472

		0.482

		0.492

		0.502

		0.512

		0.522

		0.532

		0.542

		0.552

		0.562

		0.572

		0.582

		0.592

		0.602

		0.612

		0.622

		0.632

		0.642

		0.652

		0.662

		0.672

		0.682

		0.692

		0.702

		0.712

		0.722

		0.732

		0.742

		0.752

		0.762

		0.772

		0.782

		0.792

		0.802

		0.812

		0.822

		0.832

		0.842

		0.854

		0.864

		0.874

		0.884

		0.895

		0.904

		0.914

		0.925

		0.934

		0.944

		0.954

		0.964

		0.974

		0.984

		0.994

		1.004

		1.014

		1.024

		1.035

		1.044

		1.054

		1.064

		1.074

		1.084

		1.094

		1.106

		1.114

		1.124

		1.134

		1.145

		1.154

		1.164

		1.174

		1.184

		1.194

		1.204

		1.214

		1.224

		1.234

		1.244

		1.255

		1.264

		1.274

		1.284

		1.294

		1.304

		1.314

		1.324

		1.334

		1.344

		1.354

		1.365

		1.374

		1.384

		1.394

		1.404

		1.414

		1.424

		1.436

		1.446

		1.456

		1.466

		1.478

		1.487

		1.497

		1.507

		1.517

		1.527

		1.537

		1.547

		1.557

		1.567

		1.577

		1.587

		1.597

		1.607

		1.618

		1.627

		1.637

		1.647

		1.657

		1.667

		1.677

		1.687

		1.697

		1.707

		1.717

		1.727

		1.739

		1.749

		1.758

		1.768

		1.778

		1.79

		1.799

		1.809

		1.819

		1.829

		1.839

		1.849

		1.859

		1.869

		1.88

		1.889

		1.899

		1.91

		1.919

		1.929

		1.939

		1.949

		1.959

		1.971

		1.98

		1.99

		2

		2.009

		2.019

		2.03

		2.04

		2.05

		2.06

		2.07

		2.08

		2.09

		2.1

		2.11

		2.12

		2.13

		2.14

		2.15

		2.16

		2.17

		2.18

		2.19

		2.2

		2.21

		2.22

		2.23

		2.24

		2.25

		2.259

		2.27

		2.28

		2.29

		2.3

		2.31

		2.32

		2.33

		2.34

		2.35

		2.36

		2.373

		2.383

		2.393

		2.403

		2.413

		2.423

		2.433

		2.443

		2.453

		2.463

		2.473

		2.483

		2.493

		2.503

		2.513

		2.523

		2.533

		2.543

		2.553

		2.563

		2.573

		2.583

		2.593

		2.603

		2.613

		2.627

		2.633

		2.643

		2.653

		2.663

		2.673

		2.683

		2.693

		2.703

		2.713

		2.723

		2.733

		2.743

		2.753

		2.763

		2.773

		2.783

		2.793

		2.803

		2.813

		2.823

		2.833

		2.843

		2.853

		2.863

		2.873

		2.883

		2.893

		2.903

		2.913

		2.923

		2.933

		2.943

		2.953

		2.963

		2.973

		2.983

		2.993

		3.003

		3.013

		3.023

		3.033

		3.043

		3.053

		3.063

		3.073

		3.083

		3.094

		3.103

		3.113

		3.123

		3.133

		3.143

		3.153

		3.163

		3.173

		3.183

		3.193

		3.203

		3.213

		3.223

		3.233

		3.243

		3.253

		3.263

		3.273

		3.283

		3.293

		3.303

		3.313

		3.323

		3.333

		3.343

		3.353

		3.363

		3.373

		3.386

		3.393

		3.403

		3.413

		3.423

		3.433

		3.443

		3.453

		3.463

		3.473

		3.483

		3.493

		3.503

		3.513

		3.523

		3.533

		3.543

		3.553

		3.563

		3.573

		3.583

		3.593

		3.603

		3.613

		3.623

		3.633

		3.643

		3.653

		3.663

		3.673

		3.683

		3.693

		3.703

		3.713

		3.723

		3.733

		3.743

		3.753

		3.763

		3.773

		3.783

		3.793

		3.803

		3.813

		3.823

		3.833

		3.843

		3.853

		3.863

		3.873

		3.883

		3.893

		3.903

		3.913

		3.923

		3.933

		3.943

		3.953

		3.963

		3.973

		3.983

		3.993

		4.003

		4.013

		4.023

		4.032

		4.043

		4.052

		4.063

		4.073

		4.083

		4.093

		4.103

		4.113

		4.123

		4.133

		4.144

		4.154

		4.164

		4.174

		4.184

		4.194

		4.203

		4.214

		4.223

		4.234

		4.244

		4.254

		4.264

		4.274

		4.284

		4.294

		4.304

		4.314

		4.324

		4.334

		4.344

		4.354

		4.364

		4.374

		4.384

		4.397

		4.404

		4.414

		4.424

		4.434

		4.444

		4.453

		4.464

		4.473

		4.484

		4.494

		4.504

		4.514

		4.524

		4.534

		4.544

		4.554

		4.564

		4.574

		4.584

		4.594

		4.604

		4.614

		4.624

		4.634

		4.644

		4.654

		4.664

		4.674

		4.684

		4.694

		4.703

		4.714

		4.724

		4.734

		4.744

		4.754

		4.764

		4.774

		4.784

		4.794

		4.804

		4.814

		4.825

		4.834

		4.844

		4.854

		4.864

		4.874

		4.884

		4.894

		4.904

		4.914

		4.924

		4.935

		4.944

		4.954

		4.964

		4.974

		4.984

		4.994

		5.004

		5.014

		5.024

		5.034

		5.045

		5.054

		5.064

		5.074

		5.084

		5.094

		5.104

		5.114

		5.124

		5.134

		5.144

		5.156

		5.164

		5.174

		5.184

		5.194

		5.204

		5.214

		5.224

		5.234

		5.244

		5.254

		5.264

		5.274

		5.284

		5.294

		5.304

		5.314

		5.324

		5.334

		5.344

		5.354

		5.364

		5.374

		5.384

		5.394

		5.407

		5.414

		5.424

		5.434

		5.444

		5.454

		5.464

		5.474

		5.484

		5.494

		5.504

		5.514

		5.524

		5.534

		5.544

		5.554

		5.564

		5.574

		5.584

		5.595

		5.604

		5.614

		5.624

		5.634

		5.644

		5.654

		5.664

		5.674

		5.684

		5.694

		5.705

		5.714

		5.724

		5.734

		5.744

		5.754

		5.764

		5.774

		5.784

		5.794

		5.804

		5.815

		5.824

		5.834

		5.844

		5.854

		5.864

		5.874

		5.884

		5.894

		5.904

		5.914

		5.925

		5.934

		5.944

		5.954

		5.964

		5.974

		5.984

		5.994

		6.004

		6.014

		6.024

		6.035

		6.044

		6.054

		6.064

		6.074

		6.084

		6.094

		6.104

		6.114

		6.124

		6.134

		6.145

		6.154

		6.165

		6.174

		6.184

		6.194

		6.204

		6.214

		6.224

		6.234

		6.244

		6.255

		6.264

		6.274

		6.284

		6.294

		6.304

		6.314

		6.324

		6.334

		6.344

		6.354

		6.365

		6.374

		6.384

		6.394

		6.404

		6.417

		6.424

		6.434

		6.444

		6.454

		6.464

		6.474

		6.484

		6.494

		6.504

		6.514

		6.524

		6.534

		6.544

		6.554

		6.564

		6.574

		6.584

		6.594

		6.604

		6.614

		6.624

		6.634

		6.644

		6.654

		6.664

		6.674

		6.684

		6.694

		6.704

		6.714

		6.724

		6.734

		6.744

		6.754

		6.764

		6.774

		6.784

		6.794

		6.804

		6.814

		6.824

		6.834

		6.844

		6.854

		6.864

		6.874

		6.884

		6.894

		6.904

		6.914

		6.924

		6.934

		6.944

		6.954

		6.964

		6.974

		6.984

		6.994

		7.004

		7.014

		7.024

		7.034

		7.044

		7.054

		7.064

		7.074

		7.084

		7.094

		7.104

		7.114

		7.124

		7.134

		7.144

		7.154

		7.164

		7.176

		7.185

		7.195

		7.205

		7.215

		7.225

		7.235

		7.245

		7.255

		7.265

		7.275

		7.285

		7.295

		7.305

		7.315

		7.325

		7.335

		7.345

		7.355

		7.365

		7.375

		7.385

		7.395

		7.405

		7.415

		7.428

		7.435

		7.445

		7.455

		7.465

		7.475

		7.485

		7.495

		7.505

		7.515

		7.525

		7.535

		7.545

		7.555

		7.565

		7.575

		7.585

		7.595

		7.605

		7.615

		7.625

		7.635

		7.645

		7.655

		7.665

		7.675

		7.685

		7.695

		7.705

		7.715

		7.725

		7.735

		7.745

		7.755

		7.765

		7.775

		7.785

		7.795

		7.805

		7.815

		7.825

		7.835

		7.845

		7.855

		7.865

		7.875

		7.885

		7.895

		7.905

		7.915

		7.925

		7.935

		7.945

		7.955

		7.965

		7.975

		7.985

		7.995

		8.005

		8.015

		8.025

		8.035

		8.045

		8.055

		8.065

		8.075

		8.085

		8.095

		8.105

		8.115

		8.125

		8.135

		8.145

		8.155

		8.164

		8.175

		8.185

		8.195

		8.204

		8.215

		8.225

		8.235

		8.245

		8.255

		8.265

		8.275

		8.285

		8.295

		8.305

		8.315

		8.325

		8.335

		8.345

		8.355

		8.365

		8.375

		8.385

		8.395

		8.405

		8.414

		8.425

		8.437

		8.447

		8.457

		8.467

		8.477

		8.487

		8.497

		8.506

		8.517

		8.527

		8.537

		8.546

		8.557

		8.567

		8.577

		8.587

		8.597

		8.607

		8.617

		8.627

		8.637

		8.647

		8.657

		8.667

		8.677

		8.69

		8.697

		8.707

		8.717

		8.727

		8.737

		8.747

		8.756

		8.767

		8.777

		8.787

		8.797

		8.807

		8.817

		8.827

		8.837

		8.847

		8.857

		8.867

		8.877

		8.887

		8.897

		8.907

		8.917

		8.927

		8.937

		8.947

		8.957

		8.967

		8.977

		8.987

		8.997

		9.006

		9.017

		9.027

		9.037

		9.047

		9.057

		9.067

		9.077

		9.087

		9.097

		9.107

		9.117

		9.127

		9.137

		9.147

		9.157

		9.167

		9.177

		9.187

		9.197

		9.207

		9.217

		9.227

		9.237

		9.247

		9.256

		9.267

		9.277

		9.287

		9.297

		9.307

		9.317

		9.327

		9.337

		9.347

		9.357

		9.367

		9.377

		9.387

		9.397

		9.407

		9.417

		9.427

		9.437

		9.447

		9.457

		9.467

		9.477

		9.487

		9.497

		9.507

		9.517

		9.527

		9.537

		9.547

		9.557

		9.567

		9.577

		9.587

		9.597

		9.607

		9.617

		9.627

		9.637

		9.647

		9.657

		9.667

		9.677

		9.687

		9.7

		9.707

		9.717

		9.727

		9.737

		9.747

		9.757

		9.767

		9.777

		9.787

		9.797

		9.807

		9.817

		9.827

		9.837

		9.847

		9.857

		9.867

		9.877

		9.887

		9.897

		9.907

		9.917

		9.927

		9.937

		9.947

		9.957

		9.967

		9.977

		9.987

		9.997

		10.007

		10.017

		10.027

		10.037

		10.047

		10.057

		10.067

		10.077

		10.087

		10.097

		10.107

		10.117

		10.127

		10.137

		10.147

		10.157

		10.167

		10.177

		10.187

		10.197

		10.207

		10.217

		10.227

		10.237

		10.247

		10.257

		10.267

		10.277

		10.287

		10.297

		10.307

		10.317

		10.327

		10.337

		10.347

		10.357

		10.367

		10.377

		10.387

		10.397

		10.407

		10.417

		10.427

		10.437

		10.447

		10.459

		10.468

		10.478

		10.488

		10.498

		10.508

		10.518

		10.528

		10.538

		10.548

		10.558

		10.568

		10.578

		10.588

		10.598

		10.608

		10.618

		10.628

		10.638

		10.649

		10.658

		10.668

		10.678

		10.688

		10.698

		10.711

		10.718

		10.728

		10.738

		10.748

		10.759

		10.768

		10.778

		10.788

		10.798

		10.808

		10.818

		10.828

		10.838

		10.848

		10.858

		10.869

		10.878

		10.888

		10.898

		10.908

		10.918

		10.928

		10.938

		10.948

		10.958

		10.968

		10.979

		10.988

		10.998

		11.008

		11.018

		11.028

		11.038

		11.048

		11.058

		11.068

		11.078

		11.089

		11.098

		11.108

		11.118

		11.128

		11.138

		11.148

		11.158

		11.168

		11.178

		11.188

		11.199

		11.208

		11.218

		11.228

		11.238

		11.248

		11.258

		11.268

		11.278

		11.288

		11.298

		11.309

		11.318

		11.328

		11.338

		11.348

		11.358

		11.368

		11.378

		11.388

		11.398

		11.408

		11.418

		11.428

		11.438

		11.448

		11.458

		11.469

		11.478

		11.488

		11.498

		11.508

		11.518

		11.528

		11.538

		11.548

		11.558

		11.568

		11.578

		11.588

		11.598

		11.608

		11.618

		11.628

		11.638

		11.648

		11.658

		11.668

		11.678

		11.688

		11.698

		11.708

		11.721

		11.728

		11.738

		11.748

		11.758

		11.768

		11.778

		11.788

		11.798

		11.808

		11.818

		11.828

		11.838

		11.848

		11.858

		11.868

		11.878

		11.888

		11.898

		11.908

		11.918

		11.928

		11.938

		11.948

		11.958

		11.968

		11.978

		11.988

		11.998

		12.008

		12.018

		12.028

		12.038

		12.048

		12.058

		12.068

		12.078

		12.088

		12.098

		12.108

		12.118

		12.128

		12.138

		12.148

		12.158

		12.168

		12.178

		12.188

		12.198

		12.208

		12.218

		12.228

		12.238

		12.248

		12.258

		12.268

		12.278

		12.288

		12.298

		12.308

		12.318

		12.328

		12.338

		12.348

		12.358

		12.368

		12.378

		12.388

		12.398

		12.408

		12.418

		12.428

		12.438

		12.448

		12.458

		12.468

		12.478

		12.488

		12.498

		12.508

		12.518

		12.528

		12.538

		12.548

		12.558

		12.568

		12.578

		12.588

		12.598

		12.608

		12.618

		12.628

		12.638

		12.648

		12.658

		12.668

		12.678

		12.688

		12.698

		12.708

		12.718

		12.731

		12.74

		12.75

		12.76

		12.77

		12.78

		12.79

		12.8

		12.81

		12.82

		12.83

		12.84

		12.85

		12.86

		12.87

		12.88

		12.89

		12.9

		12.91

		12.92

		12.93

		12.94

		12.95

		12.96

		12.97

		12.983

		12.99

		13

		13.01

		13.02

		13.03

		13.04

		13.05

		13.06

		13.07

		13.08

		13.09

		13.1

		13.11

		13.122

		13.131

		13.141

		13.151

		13.161

		13.171

		13.181

		13.191

		13.201

		13.211

		13.221

		13.231

		13.241

		13.251

		13.261

		13.271

		13.281

		13.291

		13.301

		13.311

		13.321

		13.331

		13.341

		13.351

		13.361

		13.371

		13.381

		13.391

		13.401

		13.411

		13.421

		13.431

		13.441

		13.451

		13.461

		13.471

		13.481

		13.491

		13.501

		13.511

		13.521

		13.531

		13.541

		13.551

		13.561

		13.571

		13.581

		13.591

		13.601

		13.611

		13.621

		13.631

		13.641

		13.651

		13.661

		13.671

		13.681

		13.691

		13.701

		13.711

		13.721

		13.731

		13.741

		13.751

		13.761

		13.771

		13.781

		13.791

		13.801

		13.811

		13.821

		13.831

		13.841

		13.851

		13.861

		13.871

		13.881

		13.891

		13.901

		13.911

		13.921

		13.931

		13.941

		13.951

		13.961

		13.971

		13.981

		13.991

		14.001

		14.011

		14.021

		14.031

		14.041

		14.051

		14.061

		14.071

		14.081

		14.091

		14.101

		14.111

		14.121

		14.131

		14.141

		14.151

		14.161

		14.171

		14.181

		14.191

		14.201

		14.211

		14.221

		14.231

		14.241

		14.251

		14.261

		14.271

		14.281

		14.291

		14.301

		14.311

		14.321

		14.331

		14.341

		14.351

		14.361

		14.371

		14.381

		14.391

		14.401

		14.411

		14.421

		14.431

		14.441

		14.451

		14.461

		14.471

		14.481

		14.491

		14.501

		14.511

		14.521

		14.531

		14.541

		14.551

		14.561

		14.571

		14.581

		14.591

		14.601

		14.611

		14.621

		14.631

		14.641

		14.651

		14.661

		14.671

		14.681

		14.691

		14.701

		14.711

		14.721

		14.731

		14.741

		14.751

		14.761

		14.771

		14.781

		14.791

		14.801

		14.811

		14.821

		14.831

		14.841

		14.851

		14.861

		14.871

		14.881

		14.891

		14.901

		14.911

		14.921

		14.931

		14.941

		14.951

		14.961

		14.971

		14.981

		14.991

		15.001

		15.011

		15.021

		15.031

		15.041

		15.051

		15.061

		15.071

		15.081

		15.091

		15.101

		15.111

		15.121

		15.131

		15.141

		15.151

		15.161

		15.171

		15.181

		15.191

		15.201

		15.211

		15.221

		15.231

		15.241

		15.251

		15.261

		15.271

		15.281

		15.291

		15.301

		15.311

		15.321

		15.331

		15.341

		15.351

		15.361

		15.371

		15.381

		15.391

		15.401

		15.411

		15.421

		15.431

		15.441

		15.451

		15.461

		15.471

		15.481

		15.491

		15.501

		15.511

		15.521

		15.531

		15.541

		15.551

		15.561

		15.571

		15.581

		15.591

		15.601

		15.611

		15.621

		15.631

		15.641

		15.651

		15.661

		15.671

		15.681

		15.691

		15.701

		15.711

		15.721

		15.731

		15.741

		15.751

		15.761

		15.771

		15.781

		15.791

		15.801

		15.811

		15.821

		15.831

		15.841

		15.851

		15.861

		15.871

		15.881

		15.891

		15.901

		15.911

		15.921

		15.931

		15.941

		15.951

		15.961

		15.971

		15.981

		15.991

		16.001

		16.011

		16.02

		16.031

		16.041

		16.051

		16.061

		16.071

		16.081

		16.091

		16.1

		16.111

		16.121

		16.131

		16.141

		16.151

		16.161

		16.171

		16.181

		16.19

		16.201

		16.211

		16.221

		16.231

		16.241

		16.251

		16.261

		16.27

		16.281

		16.291

		16.301

		16.311

		16.321

		16.331

		16.341

		16.35

		16.361

		16.371

		16.381

		16.391

		16.401

		16.411

		16.421

		16.431

		16.44

		16.451

		16.461

		16.472

		16.481

		16.491

		16.501

		16.511

		16.52

		16.531

		16.541

		16.551

		16.561

		16.571

		16.582

		16.591

		16.601

		16.611

		16.621

		16.631

		16.641

		16.651

		16.661

		16.671

		16.681

		16.692

		16.701

		16.711

		16.721

		16.731

		16.741

		16.751

		16.761

		16.77

		16.781

		16.791

		16.802

		16.811

		16.821

		16.831

		16.841

		16.851

		16.861

		16.871

		16.881

		16.891

		16.901

		16.912

		16.921

		16.931

		16.94

		16.951

		16.961

		16.971

		16.981

		16.991

		17.001

		17.011

		17.022

		17.031

		17.041

		17.051

		17.061

		17.071

		17.081

		17.091

		17.101

		17.111

		17.121

		17.131

		17.141

		17.151

		17.161

		17.171

		17.181

		17.19

		17.201

		17.211

		17.221

		17.231

		17.241

		17.251

		17.261

		17.27

		17.281

		17.291

		17.301

		17.311

		17.321

		17.331

		17.341

		17.351

		17.361

		17.371

		17.381

		17.391

		17.401

		17.411

		17.421

		17.431

		17.44

		17.451

		17.461

		17.471

		17.481

		17.491

		17.501

		17.511

		17.52

		17.531

		17.541

		17.551

		17.561

		17.571

		17.581

		17.591

		17.601

		17.611

		17.621

		17.631

		17.641

		17.651

		17.661

		17.671

		17.681

		17.69

		17.701

		17.711

		17.721

		17.731

		17.741

		17.751

		17.761

		17.77

		17.781

		17.791

		17.801

		17.811

		17.821

		17.831

		17.841

		17.851

		17.861

		17.871

		17.881

		17.891

		17.901

		17.911

		17.921

		17.931

		17.94

		17.951

		17.961

		17.971

		17.981

		17.991

		18.001

		18.011

		18.02

		18.031

		18.041

		18.051

		18.061

		18.071

		18.081

		18.091

		18.101

		18.111

		18.121

		18.131

		18.141

		18.151

		18.161

		18.171

		18.181

		18.19

		18.201

		18.211

		18.221

		18.231

		18.241

		18.251

		18.261

		18.271

		18.281

		18.291

		18.301

		18.311

		18.321

		18.331

		18.341

		18.351

		18.361

		18.371

		18.381

		18.391

		18.401

		18.411

		18.421

		18.431

		18.44

		18.451

		18.461

		18.471

		18.481

		18.491

		18.501

		18.511

		18.521

		18.531

		18.541

		18.551

		18.561

		18.571

		18.581

		18.591

		18.601

		18.611

		18.621

		18.631

		18.641

		18.651

		18.661

		18.671

		18.681

		18.69

		18.701

		18.711

		18.721

		18.731

		18.741

		18.751

		18.761

		18.771

		18.781

		18.791

		18.801

		18.811

		18.821

		18.831

		18.841

		18.851

		18.861

		18.871

		18.881

		18.891

		18.901

		18.911

		18.921

		18.931

		18.94

		18.951

		18.961

		18.971

		18.981

		18.991

		19.001

		19.011

		19.021

		19.031

		19.041

		19.051

		19.061

		19.071

		19.081

		19.091

		19.101

		19.111

		19.121

		19.131

		19.141

		19.151

		19.161

		19.171

		19.181

		19.19

		19.201

		19.211

		19.221

		19.231

		19.241

		19.251

		19.261

		19.271

		19.281

		19.291

		19.301

		19.311

		19.321

		19.331

		19.341

		19.351

		19.361

		19.371

		19.381

		19.391

		19.401

		19.411

		19.421

		19.431

		19.44

		19.451

		19.461

		19.471

		19.481

		19.491

		19.501

		19.511

		19.521

		19.531

		19.541

		19.551

		19.561

		19.571

		19.581

		19.591

		19.601

		19.611

		19.621

		19.631

		19.641

		19.651

		19.661

		19.671

		19.681

		19.69

		19.701

		19.711

		19.721

		19.731

		19.741

		19.751

		19.761

		19.771

		19.781

		19.791

		19.801

		19.811

		19.821

		19.831

		19.841

		19.851

		19.861

		19.871

		19.881

		19.891

		19.901

		19.911

		19.921

		19.931

		19.941

		19.951

		19.961

		19.971

		19.981

		19.991

		20.001

		20.011

		20.021

		20.031

		20.041

		20.051

		20.061

		20.071

		20.081

		20.091

		20.101

		20.111

		20.121

		20.131

		20.141

		20.151

		20.161

		20.171

		20.181

		20.191

		20.201

		20.211

		20.221

		20.231

		20.241

		20.251

		20.261

		20.271

		20.281

		20.291

		20.301

		20.311

		20.321

		20.331

		20.341

		20.351

		20.361

		20.371

		20.381

		20.391

		20.401

		20.411

		20.421

		20.431

		20.441

		20.451

		20.461

		20.471

		20.481

		20.493

		20.503

		20.513

		20.523

		20.533

		20.543

		20.553

		20.563

		20.573

		20.583

		20.593

		20.603

		20.613

		20.623

		20.633

		20.643

		20.653

		20.663

		20.673

		20.683

		20.693

		20.703

		20.713

		20.723

		20.733

		20.743

		20.753

		20.763

		20.773

		20.783

		20.793

		20.803

		20.813

		20.823

		20.833

		20.843

		20.853

		20.863

		20.873

		20.883

		20.893

		20.903

		20.913

		20.923

		20.933

		20.943

		20.953

		20.963

		20.973

		20.983

		20.993

		21.003

		21.013

		21.023

		21.033

		21.043

		21.053

		21.063

		21.073

		21.083

		21.093

		21.103

		21.113

		21.123

		21.133

		21.143

		21.153

		21.163

		21.173

		21.183

		21.193

		21.203

		21.213

		21.223

		21.233

		21.243

		21.253

		21.263

		21.273

		21.283

		21.293

		21.303

		21.313

		21.323

		21.333

		21.343

		21.353

		21.363

		21.373

		21.383

		21.393

		21.403

		21.413

		21.423

		21.433

		21.443

		21.453

		21.463

		21.473

		21.483

		21.493

		21.503

		21.513

		21.523

		21.533

		21.543

		21.553

		21.563

		21.573

		21.583

		21.593

		21.603

		21.613

		21.623

		21.633

		21.643

		21.653

		21.663

		21.673

		21.683

		21.693

		21.703

		21.713

		21.723

		21.733

		21.743

		21.753

		21.763

		21.773

		21.783

		21.793

		21.803

		21.813

		21.823

		21.833

		21.843

		21.853

		21.863

		21.873

		21.883

		21.893

		21.903

		21.913

		21.923

		21.933

		21.943

		21.953

		21.963

		21.973

		21.983

		21.993

		22.003

		22.013

		22.023

		22.033

		22.043

		22.053

		22.063

		22.073

		22.083

		22.093

		22.103

		22.113

		22.123

		22.133

		22.143

		22.153

		22.163

		22.173

		22.183

		22.193

		22.203

		22.213

		22.223

		22.233

		22.243

		22.253

		22.263

		22.273

		22.283

		22.293

		22.303

		22.313

		22.323

		22.333

		22.343

		22.353

		22.363

		22.373

		22.383

		22.393

		22.403

		22.413

		22.423

		22.433

		22.443

		22.453

		22.463

		22.473

		22.483

		22.493

		22.503

		22.513

		22.523

		22.533

		22.543

		22.553

		22.563

		22.573

		22.583

		22.593

		22.603

		22.613

		22.623

		22.633

		22.643

		22.653

		22.663

		22.673

		22.683

		22.693

		22.703

		22.713

		22.723

		22.733

		22.743

		22.753

		22.763

		22.773

		22.783

		22.793

		22.803

		22.813

		22.823

		22.833

		22.843

		22.853

		22.863

		22.873

		22.883

		22.893

		22.903

		22.913

		22.923

		22.933

		22.943

		22.953

		22.963

		22.973

		22.983

		22.997

		23.003

		23.013

		23.023

		23.033

		23.043

		23.053

		23.064

		23.073

		23.083

		23.093

		23.103

		23.113

		23.123

		23.133

		23.143

		23.153

		23.163

		23.174

		23.183

		23.193

		23.203

		23.213

		23.223

		23.233

		23.243

		23.253

		23.263

		23.273

		23.284

		23.293

		23.303

		23.313

		23.323

		23.333

		23.343

		23.353

		23.363

		23.373

		23.383

		23.394

		23.403

		23.413

		23.423

		23.433

		23.443

		23.453

		23.463

		23.473

		23.483

		23.493

		23.504

		23.513

		23.523

		23.533

		23.543

		23.553

		23.563

		23.573

		23.583

		23.593

		23.603

		23.614

		23.623

		23.633

		23.643

		23.653

		23.663

		23.673

		23.683

		23.693

		23.703

		23.713

		23.724

		23.733

		23.743

		23.753

		23.763

		23.773

		23.783

		23.793

		23.803

		23.813

		23.823

		23.833

		23.843

		23.853

		23.863

		23.873

		23.883

		23.893

		23.903

		23.913

		23.923

		23.933

		23.943

		23.953

		23.963

		23.973

		23.983

		23.993

		24.003

		24.013

		24.023

		24.033

		24.043

		24.053

		24.063

		24.073

		24.083

		24.093

		24.103

		24.113

		24.123

		24.133

		24.143

		24.153

		24.163

		24.173

		24.183

		24.193

		24.203

		24.213

		24.223

		24.233

		24.243

		24.253

		24.263

		24.273

		24.283

		24.293

		24.303

		24.313

		24.323

		24.333

		24.343

		24.353

		24.363

		24.373

		24.383

		24.393

		24.403

		24.413

		24.423

		24.433

		24.443

		24.453

		24.463

		24.473

		24.483

		24.493

		24.503

		24.513

		24.523

		24.533

		24.543

		24.553

		24.563

		24.573

		24.583

		24.593

		24.603

		24.613

		24.623

		24.633

		24.643

		24.653

		24.663

		24.673

		24.683

		24.693

		24.703

		24.713

		24.723

		24.733

		24.743

		24.753

		24.763

		24.773

		24.783

		24.793

		24.803

		24.813

		24.823

		24.833

		24.843

		24.853

		24.863

		24.873

		24.883

		24.893

		24.903

		24.913

		24.923

		24.933

		24.943

		24.953

		24.963

		24.973

		24.983

		24.993

		25.003

		25.013

		25.023

		25.033

		25.043

		25.053

		25.063

		25.073

		25.083

		25.093

		25.103

		25.113

		25.123

		25.133

		25.143

		25.153

		25.163

		25.173

		25.183

		25.193

		25.203

		25.213

		25.223

		25.233

		25.243

		25.253

		25.263

		25.273

		25.283

		25.293

		25.303

		25.313

		25.323

		25.333

		25.343

		25.353

		25.363

		25.373

		25.383

		25.393

		25.403

		25.413

		25.423

		25.433

		25.443

		25.453

		25.463

		25.473

		25.483

		25.493

		25.503

		25.513

		25.523

		25.533

		25.543

		25.553

		25.563

		25.573

		25.583

		25.593

		25.603

		25.613

		25.623

		25.633

		25.643

		25.653

		25.663

		25.673

		25.683

		25.693

		25.703

		25.713

		25.723

		25.733

		25.743

		25.753

		25.763

		25.773

		25.783

		25.793

		25.803

		25.813

		25.823

		25.833

		25.843

		25.853

		25.863

		25.873

		25.883

		25.893

		25.903

		25.913

		25.923

		25.933

		25.943

		25.953

		25.963

		25.973

		25.983

		25.993

		26.003

		26.013

		26.023

		26.033

		26.043

		26.053

		26.063

		26.073

		26.083

		26.093

		26.103

		26.113

		26.123

		26.133

		26.143

		26.153

		26.163

		26.173

		26.183

		26.193

		26.203

		26.213

		26.223

		26.233

		26.243

		26.253

		26.263

		26.273

		26.283

		26.293

		26.303

		26.313

		26.323

		26.333

		26.343

		26.353

		26.363

		26.373

		26.383

		26.393

		26.403

		26.413

		26.423

		26.433

		26.443

		26.453

		26.463

		26.473

		26.483

		26.493

		26.503

		26.513

		26.523

		26.533

		26.543

		26.553

		26.563

		26.573

		26.583

		26.593

		26.603

		26.613

		26.623

		26.633

		26.643

		26.653

		26.663

		26.673

		26.683

		26.693

		26.703

		26.713

		26.723

		26.733

		26.743

		26.753

		26.763

		26.773

		26.783

		26.793

		26.803

		26.813

		26.823

		26.833

		26.843

		26.853

		26.863

		26.873

		26.883

		26.893

		26.903

		26.913

		26.923

		26.933

		26.943

		26.953

		26.963

		26.973

		26.983

		26.993

		27.003

		27.013

		27.023

		27.033

		27.043

		27.053

		27.063

		27.073

		27.083

		27.093

		27.103

		27.113

		27.123

		27.133

		27.143

		27.153

		27.163

		27.173

		27.184

		27.193

		27.203

		27.213

		27.223

		27.233

		27.243

		27.253

		27.263

		27.273

		27.283

		27.294

		27.303

		27.313

		27.323

		27.333

		27.343

		27.353

		27.363

		27.373

		27.383

		27.393

		27.404

		27.413

		27.423

		27.433

		27.443

		27.453

		27.463

		27.473

		27.483

		27.493

		27.503

		27.514

		27.523

		27.533

		27.543

		27.553

		27.563

		27.573

		27.583

		27.593

		27.603

		27.613

		27.624

		27.633

		27.643

		27.653

		27.663

		27.673

		27.683

		27.693

		27.703

		27.713

		27.723

		27.734

		27.743

		27.753

		27.763

		27.773

		27.783

		27.793

		27.803

		27.813

		27.823

		27.833

		27.844

		27.853

		27.863

		27.873

		27.883

		27.893

		27.903

		27.913

		27.923

		27.933

		27.943

		27.954

		27.963

		27.973

		27.983

		27.993

		28.003

		28.013

		28.023

		28.033

		28.043

		28.053

		28.063

		28.073

		28.083

		28.093

		28.103

		28.113

		28.123

		28.133

		28.143

		28.153

		28.163

		28.173

		28.183

		28.193

		28.203

		28.213

		28.223

		28.233

		28.243

		28.253

		28.263

		28.273

		28.283

		28.293

		28.303

		28.313

		28.323

		28.333

		28.343

		28.353

		28.363

		28.373

		28.383

		28.393

		28.403

		28.414

		28.424

		28.434

		28.444

		28.454

		28.464

		28.474

		28.484

		28.494

		28.504

		28.514

		28.524

		28.534

		28.544

		28.554

		28.564

		28.574

		28.584

		28.594

		28.604

		28.614

		28.624

		28.634

		28.644

		28.654

		28.664

		28.674

		28.684

		28.694

		28.704

		28.714

		28.724

		28.734

		28.744

		28.754

		28.764

		28.774

		28.784

		28.794

		28.804

		28.814

		28.824

		28.834

		28.844

		28.854

		28.864

		28.874

		28.884

		28.894

		28.904

		28.914

		28.924

		28.934

		28.944

		28.954

		28.964

		28.974

		28.984

		28.994

		29.004

		29.014

		29.024

		29.034

		29.044

		29.054

		29.064

		29.074

		29.084

		29.094

		29.104

		29.114

		29.124

		29.134

		29.144

		29.154

		29.164

		29.174

		29.184

		29.194

		29.204

		29.214

		29.224

		29.234

		29.244

		29.254

		29.264

		29.274

		29.284

		29.294

		29.304

		29.314

		29.324

		29.334

		29.344

		29.354

		29.364

		29.374

		29.384

		29.394

		29.404

		29.414

		29.424

		29.434

		29.444

		29.454

		29.464

		29.474

		29.484

		29.494

		29.504

		29.514

		29.524

		29.534

		29.544

		29.554

		29.564

		29.574

		29.584

		29.594

		29.604

		29.614

		29.624

		29.634

		29.644

		29.654

		29.664

		29.674

		29.684

		29.694

		29.704

		29.714

		29.724

		29.734

		29.744

		29.754

		29.764

		29.774

		29.784

		29.794

		29.804

		29.814

		29.824

		29.834

		29.844

		29.854

		29.864

		29.874

		29.884

		29.894

		29.904

		29.914

		29.924

		29.934

		29.944

		29.954

		29.964

		29.974

		29.984

		29.994

		30.004

		30.014

		30.024

		30.034

		30.044

		30.054

		30.064

		30.074

		30.084

		30.094

		30.104

		30.114

		30.124

		30.134

		30.144

		30.154

		30.164

		30.174

		30.184

		30.194

		30.204

		30.214

		30.224

		30.234

		30.244

		30.254

		30.264

		30.274

		30.284

		30.294

		30.304

		30.314

		30.324

		30.334

		30.344

		30.354

		30.364

		30.374

		30.384

		30.394

		30.404

		30.414

		30.424

		30.434

		30.444

		30.454

		30.464

		30.474

		30.484

		30.494

		30.506

		30.516

		30.526

		30.536

		30.546

		30.556

		30.566

		30.576

		30.586

		30.596

		30.606

		30.616

		30.626

		30.636

		30.646

		30.656

		30.666

		30.676

		30.686

		30.696

		30.706

		30.716

		30.726

		30.736

		30.746

		30.756

		30.766

		30.776

		30.786

		30.796

		30.806

		30.816

		30.826

		30.836

		30.846

		30.856

		30.866

		30.876

		30.886

		30.896

		30.906

		30.916

		30.926

		30.936

		30.946

		30.956

		30.966

		30.976

		30.986

		30.996

		31.006

		31.016

		31.026

		31.036

		31.046

		31.056

		31.066

		31.076

		31.086

		31.096

		31.106

		31.116

		31.126

		31.136

		31.146

		31.156

		31.166

		31.176

		31.186

		31.196

		31.206

		31.216

		31.226

		31.236

		31.246

		31.256

		31.266

		31.276

		31.286

		31.296

		31.306

		31.316

		31.326

		31.336

		31.346

		31.356

		31.366

		31.376

		31.386

		31.396

		31.406

		31.416

		31.426

		31.436

		31.446

		31.456

		31.466

		31.476

		31.486

		31.496

		31.506

		31.516

		31.526

		31.536

		31.546

		31.556

		31.566

		31.576

		31.586

		31.596

		31.606

		31.616

		31.626

		31.636

		31.646

		31.656

		31.666

		31.676

		31.686

		31.696

		31.706

		31.716

		31.726

		31.736

		31.746

		31.756

		31.766

		31.776

		31.786

		31.796

		31.806

		31.816

		31.826

		31.836

		31.846

		31.856

		31.866

		31.876

		31.886

		31.896

		31.906

		31.916

		31.926

		31.936

		31.946

		31.956

		31.966

		31.976

		31.986

		31.996

		32.006

		32.016

		32.026

		32.036

		32.046

		32.055

		32.066

		32.076

		32.086

		32.096

		32.106

		32.116

		32.126

		32.136

		32.145

		32.156

		32.166

		32.176

		32.186

		32.196

		32.206

		32.216

		32.226

		32.236

		32.246

		32.256

		32.266

		32.276

		32.286

		32.296

		32.305

		32.316

		32.326

		32.336

		32.346

		32.356

		32.366

		32.376

		32.386

		32.395

		32.406

		32.416

		32.426

		32.436

		32.446

		32.456

		32.466

		32.476

		32.486

		32.496

		32.506

		32.516

		32.526

		32.536

		32.546

		32.555

		32.566

		32.576

		32.586

		32.596

		32.606

		32.616

		32.626

		32.636

		32.645

		32.656

		32.666

		32.676

		32.686

		32.696

		32.706

		32.716

		32.726

		32.736

		32.746

		32.756

		32.766

		32.776

		32.786

		32.796

		32.805

		32.816

		32.826

		32.836

		32.846

		32.856

		32.866

		32.876

		32.886

		32.895

		32.906

		32.916

		32.926

		32.936

		32.946

		32.956

		32.966

		32.976

		32.986

		32.996

		33.01

		33.016

		33.026

		33.036

		33.046

		33.055

		33.066

		33.076

		33.086

		33.096

		33.106

		33.117

		33.126

		33.136

		33.145

		33.156

		33.166

		33.176

		33.186

		33.196

		33.206

		33.216

		33.227

		33.236

		33.246

		33.256

		33.266

		33.276

		33.286

		33.296

		33.305

		33.316

		33.326

		33.337

		33.346

		33.356

		33.366

		33.376

		33.386

		33.395

		33.406

		33.416

		33.426

		33.436

		33.447

		33.456

		33.466

		33.476

		33.486

		33.496

		33.506

		33.516

		33.526

		33.536

		33.546

		33.557

		33.566

		33.576

		33.586

		33.596

		33.606

		33.616

		33.626

		33.636

		33.645

		33.656

		33.666

		33.676

		33.686

		33.696

		33.706

		33.716

		33.726

		33.736

		33.746

		33.756

		33.766

		33.777

		33.786

		33.796

		33.805

		33.816

		33.826

		33.836

		33.846

		33.856

		33.866

		33.876

		33.886

		33.895

		33.906

		33.916

		33.926

		33.936

		33.946

		33.956

		33.966

		33.976

		33.986

		33.996

		34.006

		34.016

		34.026

		34.036

		34.046

		34.055

		34.066

		34.076

		34.086

		34.096

		34.106

		34.116

		34.126

		34.136

		34.145

		34.156

		34.166

		34.176

		34.186

		34.196

		34.206

		34.216

		34.226

		34.236

		34.246

		34.256

		34.266

		34.276

		34.286

		34.296

		34.305

		34.316

		34.326

		34.336

		34.346

		34.356

		34.366

		34.376

		34.386

		34.395

		34.406

		34.416

		34.426

		34.436

		34.446

		34.456

		34.466

		34.476

		34.486

		34.496

		34.506

		34.516

		34.526

		34.536

		34.546

		34.555

		34.566

		34.576

		34.586

		34.596

		34.606

		34.616

		34.626

		34.636

		34.645

		34.656

		34.666

		34.676

		34.686

		34.696

		34.706

		34.716

		34.726

		34.736

		34.746

		34.756

		34.766

		34.776

		34.786

		34.796

		34.805

		34.816

		34.826

		34.836

		34.846

		34.856

		34.866

		34.876

		34.886

		34.895

		34.906

		34.916

		34.926

		34.936

		34.946

		34.956

		34.966

		34.976

		34.986

		34.996

		35.006

		35.016

		35.026

		35.036

		35.046

		35.056

		35.066

		35.076

		35.086

		35.096

		35.106

		35.116

		35.126

		35.136

		35.145

		35.156

		35.166

		35.176

		35.186

		35.196

		35.206

		35.216

		35.226

		35.236

		35.246

		35.256

		35.266

		35.276

		35.286

		35.296

		35.306

		35.316

		35.326

		35.336

		35.346

		35.356

		35.366

		35.376

		35.386

		35.395

		35.406

		35.416

		35.426

		35.436

		35.446

		35.456

		35.466

		35.476

		35.486

		35.496

		35.506

		35.516

		35.526

		35.536

		35.546

		35.556

		35.566

		35.576

		35.586

		35.596

		35.606

		35.616

		35.626

		35.636

		35.645

		35.656

		35.666

		35.676

		35.686

		35.696

		35.706

		35.716

		35.726

		35.736

		35.746

		35.756

		35.766

		35.776

		35.786

		35.796

		35.806

		35.816

		35.826

		35.836

		35.846

		35.856

		35.866

		35.876

		35.886

		35.895

		35.906

		35.916

		35.926

		35.936

		35.946

		35.956

		35.966

		35.976

		35.986

		35.996

		36.006

		36.016

		36.026

		36.036

		36.046

		36.056

		36.066

		36.076

		36.086

		36.096

		36.106

		36.116

		36.126

		36.136

		36.145

		36.156

		36.166

		36.176

		36.186

		36.196

		36.206

		36.216

		36.226

		36.236

		36.246

		36.256

		36.266

		36.276

		36.286

		36.296

		36.306

		36.316

		36.326

		36.336

		36.346

		36.356

		36.366

		36.376

		36.386

		36.395

		36.406

		36.416

		36.426

		36.436

		36.446

		36.456

		36.466

		36.476

		36.486

		36.496

		36.506

		36.516

		36.526

		36.536

		36.546

		36.556

		36.566

		36.576

		36.586

		36.596

		36.606

		36.616

		36.626

		36.636

		36.645

		36.656

		36.666

		36.676

		36.686

		36.696

		36.706

		36.716

		36.726

		36.736

		36.746

		36.756

		36.766

		36.776

		36.786

		36.796

		36.806

		36.816

		36.826

		36.836

		36.846

		36.856

		36.866

		36.876

		36.886

		36.895

		36.906

		36.916

		36.926

		36.936

		36.946

		36.956

		36.966

		36.976

		36.986

		36.996

		37.006

		37.016

		37.026

		37.036

		37.046

		37.056

		37.066

		37.076

		37.086

		37.096

		37.106

		37.116

		37.127

		37.136

		37.145

		37.156

		37.166

		37.176

		37.186

		37.196

		37.206

		37.216

		37.226

		37.237

		37.246

		37.256

		37.266

		37.276

		37.286

		37.296

		37.306

		37.316

		37.326

		37.336

		37.347

		37.356

		37.366

		37.376

		37.386

		37.395

		37.406

		37.416

		37.426

		37.436

		37.446

		37.457

		37.466

		37.476

		37.486

		37.496

		37.506

		37.516

		37.526

		37.536

		37.546

		37.556

		37.567

		37.576

		37.586

		37.596

		37.606

		37.616

		37.626

		37.636

		37.645

		37.656

		37.666

		37.677

		37.686

		37.696

		37.706

		37.716

		37.726

		37.736

		37.746

		37.756

		37.766

		37.776

		37.787

		37.796

		37.806

		37.816

		37.826

		37.836

		37.846

		37.856

		37.866

		37.876

		37.886

		37.897

		37.906

		37.916

		37.926

		37.936

		37.946

		37.956

		37.966

		37.976

		37.986

		37.996

		38.006

		38.016

		38.026

		38.036

		38.046

		38.056

		38.066

		38.076

		38.086

		38.096

		38.106

		38.116

		38.126

		38.136

		38.145

		38.156

		38.166

		38.176

		38.186

		38.196

		38.206

		38.216

		38.226

		38.236

		38.246

		38.256

		38.266

		38.276

		38.286

		38.296

		38.306

		38.316

		38.326

		38.336

		38.346

		38.356

		38.366

		38.376

		38.386

		38.396

		38.406

		38.416

		38.426

		38.436

		38.446

		38.456

		38.466

		38.476

		38.486

		38.496

		38.506

		38.516

		38.526

		38.536

		38.546

		38.556

		38.566

		38.576

		38.586

		38.596

		38.606

		38.616

		38.626

		38.636

		38.646

		38.656

		38.666

		38.676

		38.686

		38.696

		38.706

		38.716

		38.726

		38.736

		38.746

		38.756

		38.766

		38.776

		38.786

		38.796

		38.806

		38.816

		38.826

		38.836

		38.846

		38.856

		38.866

		38.876

		38.886

		38.896

		38.906

		38.916

		38.926

		38.936

		38.946

		38.956

		38.966

		38.976

		38.986

		38.996

		39.006

		39.016

		39.026

		39.036

		39.046

		39.056

		39.066

		39.076

		39.086

		39.096

		39.106

		39.116

		39.126

		39.136

		39.146

		39.156

		39.166

		39.176

		39.186

		39.196

		39.206

		39.216

		39.226

		39.236

		39.246

		39.256

		39.266

		39.276

		39.286

		39.296

		39.306

		39.316

		39.326

		39.336

		39.346

		39.356

		39.366

		39.376

		39.386

		39.396

		39.406

		39.416

		39.426

		39.436

		39.446

		39.456

		39.466

		39.476

		39.486

		39.496

		39.506

		39.516

		39.526

		39.536

		39.546

		39.556

		39.566

		39.576

		39.586

		39.596

		39.606

		39.616

		39.626

		39.636

		39.646

		39.656

		39.666

		39.676

		39.686

		39.696

		39.706

		39.716

		39.726

		39.736

		39.746

		39.756

		39.766

		39.776

		39.786

		39.796

		39.806

		39.816

		39.826

		39.836

		39.846

		39.856

		39.866

		39.876

		39.886

		39.896

		39.906

		39.916

		39.926

		39.936

		39.946

		39.956

		39.966

		39.976

		39.986

		39.996

		40.006

		40.016

		40.026

		40.036

		40.046

		40.056

		40.066

		40.076

		40.086

		40.096

		40.106

		40.116

		40.126

		40.136

		40.146

		40.156

		40.166

		40.176

		40.186

		40.196

		40.206

		40.216

		40.226

		40.236

		40.246

		40.256

		40.266

		40.276

		40.286

		40.296

		40.306

		40.316

		40.326

		40.336

		40.346

		40.356

		40.366

		40.376

		40.386

		40.396

		40.406

		40.416

		40.426

		40.436

		40.446

		40.456

		40.466

		40.476

		40.486

		40.496

		40.506

		40.52

		40.526

		40.536

		40.546

		40.556

		40.566

		40.576

		40.586

		40.596

		40.606

		40.616

		40.626

		40.636

		40.646

		40.656

		40.666

		40.676

		40.686

		40.696

		40.706

		40.716

		40.726

		40.736

		40.746

		40.756

		40.766

		40.776

		40.786

		40.796

		40.806

		40.816

		40.826

		40.836

		40.846

		40.856

		40.866

		40.876

		40.886

		40.896

		40.906

		40.916

		40.926

		40.936

		40.946

		40.956

		40.966

		40.976

		40.986

		40.996

		41.006

		41.016

		41.026

		41.036

		41.046

		41.056

		41.066

		41.076

		41.086

		41.096

		41.106

		41.116

		41.126

		41.136

		41.146

		41.156

		41.166

		41.176

		41.186

		41.196

		41.206

		41.216

		41.226

		41.236

		41.246

		41.256

		41.266

		41.276

		41.286

		41.296

		41.306

		41.316

		41.326

		41.336

		41.346

		41.357

		41.366

		41.376

		41.386

		41.396

		41.406

		41.416

		41.426

		41.436

		41.446

		41.456

		41.467

		41.476

		41.486

		41.496

		41.506

		41.516

		41.526

		41.536

		41.546

		41.556

		41.566

		41.577

		41.586

		41.596

		41.606

		41.616

		41.626

		41.636

		41.646

		41.656

		41.666

		41.676

		41.687

		41.696

		41.706

		41.716

		41.726

		41.736

		41.746

		41.756

		41.766

		41.776

		41.786

		41.797

		41.806

		41.816

		41.826

		41.836

		41.846

		41.856

		41.866

		41.876

		41.886

		41.896

		41.907

		41.916

		41.926

		41.936

		41.946

		41.956

		41.966

		41.976

		41.986

		41.996

		42.006

		42.017

		42.026

		42.036

		42.046

		42.056

		42.066

		42.076

		42.086

		42.096

		42.106

		42.116

		42.126

		42.136

		42.146

		42.156

		42.166

		42.176

		42.186

		42.196

		42.206

		42.216

		42.226

		42.236

		42.246

		42.256

		42.266

		42.276

		42.286

		42.296

		42.306

		42.316

		42.326

		42.336

		42.346

		42.356

		42.366

		42.376

		42.386

		42.396

		42.406

		42.416

		42.426

		42.436

		42.446

		42.456

		42.466

		42.476

		42.486

		42.496

		42.506

		42.516

		42.526

		42.536

		42.546

		42.556

		42.566

		42.576

		42.586

		42.596

		42.606

		42.616

		42.626

		42.636

		42.646

		42.656

		42.666

		42.676

		42.686

		42.696

		42.706

		42.716

		42.726

		42.736

		42.746

		42.756

		42.766

		42.776

		42.786

		42.796

		42.806

		42.816

		42.826

		42.836

		42.846

		42.856

		42.866

		42.876

		42.886

		42.896

		42.906

		42.916

		42.926

		42.936

		42.946

		42.956

		42.966

		42.976

		42.986

		42.996

		43.006

		43.016

		43.026

		43.036

		43.046

		43.056

		43.066

		43.076

		43.086

		43.096

		43.106

		43.116

		43.126

		43.136

		43.146

		43.156

		43.166

		43.176

		43.186

		43.196

		43.206

		43.216

		43.226

		43.236

		43.246

		43.256

		43.266

		43.276

		43.286

		43.296

		43.306

		43.316

		43.326

		43.336

		43.346

		43.356

		43.366

		43.376

		43.386

		43.396

		43.406

		43.416

		43.426

		43.436

		43.446

		43.456

		43.466

		43.476

		43.486

		43.496

		43.506

		43.516

		43.526

		43.536

		43.546

		43.556

		43.566

		43.576

		43.586

		43.596

		43.606

		43.616

		43.626

		43.636

		43.646

		43.656

		43.666

		43.676

		43.686

		43.696

		43.706

		43.716

		43.726

		43.736

		43.746

		43.756

		43.766

		43.776

		43.786

		43.796

		43.806

		43.816

		43.826

		43.836

		43.846

		43.856

		43.866

		43.876

		43.886

		43.896

		43.906

		43.916

		43.926

		43.936

		43.946

		43.956

		43.966

		43.976

		43.986

		43.996

		44.006

		44.016

		44.026

		44.036

		44.046

		44.056

		44.066

		44.076

		44.086

		44.096

		44.106

		44.116

		44.126

		44.136

		44.146

		44.156

		44.166

		44.176

		44.186

		44.196

		44.206

		44.216

		44.226

		44.236

		44.246

		44.256

		44.266

		44.276

		44.286

		44.296

		44.306

		44.316

		44.326

		44.336

		44.346

		44.356

		44.366

		44.376

		44.386

		44.396

		44.406

		44.416

		44.426

		44.436

		44.446

		44.456

		44.466

		44.476

		44.486

		44.496

		44.506

		44.516

		44.526

		44.536

		44.546

		44.556

		44.566

		44.576

		44.586

		44.596

		44.606

		44.616

		44.626

		44.636

		44.646

		44.656

		44.666

		44.676

		44.686

		44.696

		44.706

		44.716

		44.726

		44.736

		44.746

		44.756

		44.766

		44.776

		44.786

		44.796

		44.806

		44.816

		44.826

		44.836

		44.846

		44.856

		44.866

		44.876

		44.886

		44.896

		44.906

		44.916

		44.926

		44.936

		44.946

		44.956

		44.966

		44.976

		44.986

		44.996

		45.006

		45.016

		45.026

		45.036

		45.046

		45.056

		45.066

		45.076

		45.086

		45.096

		45.106

		45.116

		45.126

		45.136

		45.146

		45.156

		45.166

		45.176

		45.186

		45.196

		45.206

		45.216

		45.226

		45.236

		45.246

		45.256

		45.266

		45.276

		45.286

		45.296

		45.306

		45.316

		45.326

		45.336

		45.346

		45.356

		45.366

		45.376

		45.386

		45.396

		45.406

		45.416

		45.426

		45.436

		45.446

		45.456

		45.466

		45.477

		45.486

		45.496

		45.506

		45.516

		45.526

		45.536

		45.546

		45.556

		45.566

		45.576

		45.587

		45.596

		45.606

		45.616

		45.626

		45.636

		45.646

		45.656

		45.666

		45.676

		45.686

		45.697

		45.706

		45.716

		45.726

		45.736

		45.746

		45.756

		45.766

		45.776

		45.786

		45.796

		45.807

		45.816

		45.826

		45.836

		45.846

		45.856

		45.866

		45.876

		45.886

		45.896

		45.906

		45.917

		45.926

		45.936

		45.946

		45.956

		45.966

		45.976

		45.986

		45.996

		46.006

		46.016

		46.027

		46.036

		46.046

		46.056

		46.066

		46.076

		46.086

		46.096

		46.106

		46.116

		46.126

		46.136

		46.146

		46.156

		46.166

		46.176

		46.186

		46.196

		46.206

		46.216

		46.226

		46.236

		46.247

		46.256

		46.266

		46.276

		46.286

		46.296

		46.306

		46.316

		46.326

		46.336

		46.346

		46.357

		46.366

		46.376

		46.386

		46.396

		46.406

		46.416

		46.426

		46.436

		46.446

		46.456

		46.467

		46.476

		46.486

		46.496

		46.506

		46.516

		46.526

		46.536

		46.546

		46.556

		46.566

		46.577

		46.586

		46.596

		46.606

		46.616

		46.626

		46.636

		46.646

		46.656

		46.666

		46.676

		46.687

		46.696

		46.706

		46.716

		46.726

		46.736

		46.746

		46.756

		46.766

		46.776

		46.786

		46.797

		46.806

		46.816

		46.826

		46.836

		46.846

		46.856

		46.866

		46.876

		46.886

		46.896

		46.907

		46.916

		46.926

		46.936

		46.946

		46.956

		46.966

		46.976

		46.986

		46.996

		47.006

		47.017

		47.026

		47.036

		47.046

		47.056

		47.066

		47.076

		47.086

		47.096

		47.106

		47.116

		47.126

		47.136

		47.146

		47.156

		47.166

		47.176

		47.186

		47.196

		47.206

		47.216

		47.226

		47.236

		47.246

		47.256

		47.266

		47.276

		47.286

		47.296

		47.306

		47.316

		47.326

		47.336

		47.346

		47.356

		47.366

		47.376

		47.386

		47.396

		47.406

		47.416

		47.426

		47.436

		47.446

		47.456

		47.466

		47.476

		47.486

		47.496

		47.506

		47.516

		47.526

		47.536

		47.546

		47.556

		47.566

		47.576

		47.586

		47.596

		47.606

		47.616

		47.626

		47.636

		47.646

		47.656

		47.666

		47.676

		47.686

		47.696

		47.706

		47.716

		47.726

		47.736

		47.746

		47.756

		47.766

		47.776

		47.786

		47.796

		47.806

		47.816

		47.826

		47.836

		47.846

		47.856

		47.866

		47.876

		47.886

		47.896

		47.906

		47.916

		47.926

		47.936

		47.946

		47.956

		47.966

		47.976

		47.986

		47.996

		48.006

		48.016

		48.029

		48.039

		48.049

		48.06

		48.069

		48.079

		48.089

		48.099

		48.109

		48.119

		48.129

		48.139

		48.149

		48.159

		48.17

		48.179

		48.189

		48.199

		48.209

		48.219

		48.229

		48.239

		48.249

		48.259

		48.269

		48.28

		48.289

		48.299

		48.309

		48.319

		48.329

		48.339

		48.349

		48.359

		48.369

		48.379

		48.39

		48.399

		48.409

		48.419

		48.429

		48.439

		48.449

		48.459

		48.469

		48.479

		48.489

		48.5

		48.509

		48.519

		48.529

		48.539

		48.549

		48.559

		48.569

		48.579

		48.589

		48.599

		48.61

		48.619

		48.629

		48.639

		48.649

		48.659

		48.669

		48.679

		48.689

		48.699

		48.709

		48.72

		48.729

		48.739

		48.749

		48.759

		48.769

		48.779

		48.789

		48.799

		48.809

		48.819

		48.829

		48.839

		48.849

		48.859

		48.869

		48.879

		48.889

		48.899

		48.909

		48.919

		48.929

		48.939

		48.949

		48.959

		48.969

		48.979

		48.989

		48.999

		49.009

		49.019

		49.029

		49.039

		49.049

		49.059

		49.069

		49.079

		49.089

		49.099

		49.109

		49.119

		49.129

		49.139

		49.149

		49.159

		49.169

		49.179

		49.189

		49.199

		49.209

		49.219

		49.229

		49.239

		49.249

		49.259

		49.269

		49.279

		49.289

		49.299

		49.309

		49.319

		49.329

		49.339

		49.349

		49.359

		49.369

		49.379

		49.389

		49.399

		49.409

		49.419

		49.429

		49.439

		49.449

		49.459

		49.469

		49.479

		49.489

		49.499

		49.509

		49.519

		49.529

		49.539

		49.549

		49.559

		49.569

		49.579

		49.589

		49.599

		49.609

		49.619

		49.629

		49.639

		49.649

		49.659

		49.669

		49.679

		49.689

		49.699

		49.709

		49.719

		49.729

		49.739

		49.749

		49.759

		49.769

		49.779

		49.789

		49.799

		49.809

		49.819

		49.829

		49.839

		49.849

		49.859

		49.869

		49.879

		49.889

		49.899

		49.909

		49.919

		49.929

		49.939

		49.949

		49.959

		49.969

		49.979

		49.989

		49.999

		50.009

		50.019

		50.029

		50.039

		50.049

		50.059

		50.069

		50.079

		50.089

		50.099

		50.109

		50.119

		50.129

		50.139

		50.149

		50.159

		50.169

		50.179

		50.189

		50.199

		50.209

		50.219

		50.229

		50.239

		50.249

		50.259

		50.269

		50.279

		50.289

		50.299

		50.309

		50.319

		50.329

		50.339

		50.349

		50.359

		50.369

		50.379

		50.389

		50.399

		50.409

		50.419

		50.429

		50.439

		50.449

		50.459

		50.469

		50.479

		50.489

		50.499

		50.509

		50.519

		50.533

		50.539

		50.549

		50.559

		50.569

		50.579

		50.589

		50.599

		50.609

		50.619

		50.629

		50.639

		50.649

		50.659

		50.669

		50.679

		50.689

		50.699

		50.709

		50.719

		50.729

		50.739

		50.749

		50.759

		50.769

		50.779

		50.789

		50.799

		50.809

		50.819

		50.829

		50.839

		50.849

		50.859

		50.869

		50.879

		50.889

		50.899

		50.909

		50.919

		50.929

		50.939

		50.949

		50.959

		50.969

		50.979

		50.989

		50.999

		51.009

		51.019

		51.029

		51.039

		51.049

		51.059

		51.069

		51.079

		51.089

		51.099

		51.109

		51.119

		51.129

		51.139

		51.149

		51.159

		51.169

		51.179

		51.189

		51.199

		51.209

		51.219

		51.229

		51.239

		51.249

		51.259

		51.269

		51.279

		51.289

		51.299

		51.309

		51.319

		51.329

		51.339

		51.349

		51.359

		51.369

		51.379

		51.389

		51.399

		51.409

		51.419

		51.429

		51.439

		51.449

		51.459

		51.469

		51.479

		51.489

		51.499

		51.509

		51.519

		51.529

		51.539

		51.549

		51.56

		51.569

		51.579

		51.589

		51.599

		51.609

		51.619

		51.629

		51.639

		51.649

		51.659

		51.669

		51.679

		51.689

		51.699

		51.709

		51.719

		51.729

		51.739

		51.749

		51.759

		51.769

		51.779

		51.789

		51.799

		51.809

		51.819

		51.829

		51.839

		51.849

		51.859

		51.869

		51.879

		51.889

		51.899

		51.909

		51.919

		51.929

		51.939

		51.949

		51.959

		51.969

		51.979

		51.989

		51.999

		52.009

		52.019

		52.029

		52.039

		52.049

		52.059

		52.07

		52.079

		52.089

		52.099

		52.11

		52.119

		52.129

		52.139

		52.149

		52.159

		52.169

		52.182

		52.189

		52.199

		52.209

		52.219

		52.229

		52.239

		52.249

		52.259

		52.269

		52.279

		52.29

		52.299

		52.309

		52.319

		52.329

		52.339

		52.349

		52.359

		52.369

		52.379

		52.389

		52.401

		52.409

		52.419

		52.429

		52.439

		52.449

		52.459

		52.469

		52.479

		52.489

		52.499

		52.51

		52.519

		52.529

		52.539

		52.549

		52.559

		52.569

		52.579

		52.589

		52.599

		52.609

		52.62

		52.629

		52.639

		52.649

		52.659

		52.669

		52.679

		52.689

		52.699

		52.709

		52.719

		52.73

		52.739

		52.749

		52.759

		52.769

		52.779

		52.789

		52.799

		52.809

		52.819

		52.829

		52.84

		52.849

		52.859

		52.869

		52.879

		52.889

		52.899

		52.909

		52.919

		52.929

		52.939

		52.949

		52.959

		52.969

		52.979

		52.989

		52.999

		53.009

		53.019

		53.029

		53.039

		53.049

		53.059

		53.069

		53.079

		53.089

		53.099

		53.109

		53.119

		53.129

		53.139

		53.149

		53.159

		53.169

		53.179

		53.189

		53.199

		53.209

		53.219

		53.229

		53.239

		53.249

		53.259

		53.269

		53.279

		53.289

		53.299

		53.309

		53.319

		53.329

		53.339

		53.349

		53.359

		53.369

		53.379

		53.389

		53.399

		53.409

		53.419

		53.429

		53.439

		53.449

		53.459

		53.469

		53.479

		53.489

		53.499

		53.509

		53.519

		53.529

		53.539

		53.549

		53.559

		53.569

		53.579

		53.589

		53.599

		53.609

		53.619

		53.629

		53.639

		53.649

		53.659

		53.669

		53.679

		53.689

		53.699

		53.709

		53.719

		53.729

		53.739

		53.749

		53.759

		53.769

		53.779

		53.789

		53.799

		53.809

		53.819

		53.829

		53.839

		53.849

		53.859

		53.869

		53.879

		53.889

		53.899

		53.909

		53.919

		53.929

		53.939

		53.949

		53.959

		53.969

		53.979

		53.989

		53.999

		54.009

		54.019

		54.029

		54.039

		54.049

		54.059

		54.069

		54.079

		54.089

		54.099

		54.109

		54.119

		54.129

		54.139

		54.149

		54.159

		54.169

		54.179

		54.189

		54.199

		54.209

		54.219

		54.229

		54.239

		54.249

		54.259

		54.269

		54.279

		54.289

		54.299

		54.309

		54.319

		54.329

		54.339

		54.349

		54.359

		54.369

		54.379

		54.389

		54.399

		54.409

		54.419

		54.429

		54.439

		54.449

		54.459

		54.469

		54.479

		54.489

		54.499

		54.509

		54.519

		54.529

		54.539

		54.549

		54.559

		54.569

		54.579

		54.589

		54.599

		54.609

		54.619

		54.629

		54.639

		54.649

		54.659

		54.669

		54.679

		54.689

		54.699

		54.709

		54.719

		54.729

		54.739

		54.749

		54.759

		54.769

		54.779

		54.789

		54.799

		54.809

		54.819

		54.829

		54.839

		54.849

		54.859

		54.869

		54.879

		54.889

		54.899

		54.909

		54.919

		54.929

		54.939

		54.949

		54.959

		54.969

		54.979

		54.989

		54.999

		55.009

		55.019

		55.029

		55.039

		55.049

		55.059

		55.069

		55.079

		55.089

		55.099

		55.109

		55.119

		55.129

		55.139

		55.149

		55.159

		55.169

		55.179

		55.189

		55.199

		55.209

		55.219

		55.229

		55.239

		55.249

		55.259

		55.269

		55.279

		55.289

		55.299

		55.309

		55.319

		55.329

		55.339

		55.349

		55.359

		55.369

		55.379

		55.389

		55.399

		55.409

		55.419

		55.429

		55.439

		55.449

		55.459

		55.469

		55.479

		55.489

		55.499

		55.509

		55.519

		55.529

		55.539

		55.549

		55.559

		55.569

		55.579

		55.589

		55.599

		55.609

		55.619

		55.629

		55.639

		55.649

		55.659

		55.669

		55.679

		55.689

		55.699

		55.709

		55.719

		55.729

		55.739

		55.749

		55.759

		55.769

		55.779

		55.789

		55.799

		55.809

		55.819

		55.829

		55.839

		55.849

		55.859

		55.869

		55.879

		55.889

		55.899

		55.909

		55.919

		55.929

		55.939

		55.949

		55.959

		55.969

		55.979

		55.989

		55.999

		56.009

		56.019

		56.029

		56.039

		56.049

		56.059

		56.069

		56.079

		56.089

		56.099

		56.109

		56.119

		56.129

		56.139

		56.149

		56.159

		56.169

		56.179

		56.19

		56.199

		56.209

		56.219

		56.229

		56.239

		56.249

		56.259

		56.269

		56.279

		56.289

		56.3

		56.309

		56.319

		56.329

		56.34

		56.35

		56.36

		56.37

		56.38

		56.39

		56.4

		56.41

		56.42

		56.43

		56.44

		56.45

		56.46

		56.47

		56.48

		56.49

		56.5

		56.51

		56.52

		56.53

		56.54

		56.55

		56.561

		56.57

		56.58

		56.59

		56.6

		56.611

		56.62

		56.63

		56.64

		56.65

		56.66

		56.67

		56.68

		56.69

		56.7

		56.71

		56.72

		56.73

		56.74

		56.75

		56.761

		56.77

		56.78

		56.79

		56.8

		56.81

		56.82

		56.83

		56.84

		56.85

		56.86

		56.87

		56.88

		56.89

		56.9

		56.91

		56.92

		56.93

		56.94

		56.95

		56.96

		56.97

		56.98

		56.99

		57

		57.01

		57.02

		57.03

		57.041

		57.05

		57.06

		57.07

		57.08

		57.09

		57.1

		57.11

		57.12

		57.13

		57.14

		57.15

		57.16

		57.17

		57.18

		57.19

		57.2

		57.21

		57.22

		57.23

		57.24

		57.25

		57.26

		57.271

		57.28

		57.29

		57.3

		57.31

		57.32

		57.33

		57.34

		57.35

		57.36

		57.37

		57.38

		57.39

		57.4

		57.41

		57.42

		57.43

		57.441

		57.45

		57.46

		57.47

		57.48

		57.491

		57.5

		57.51

		57.52

		57.53

		57.54

		57.55

		57.56

		57.57

		57.58

		57.59

		57.6

		57.61

		57.62

		57.63

		57.64

		57.65

		57.661

		57.67

		57.68

		57.69

		57.7

		57.711

		57.72

		57.73

		57.74

		57.75

		57.76

		57.77

		57.78

		57.79

		57.8

		57.81

		57.82

		57.83

		57.84

		57.85

		57.86

		57.87

		57.88

		57.89

		57.9

		57.91

		57.92

		57.93

		57.94

		57.95

		57.96

		57.97

		57.98

		57.99

		58

		58.01

		58.02

		58.03

		58.044

		58.053

		58.063

		58.073

		58.083

		58.093

		58.103

		58.113

		58.123

		58.133

		58.143

		58.153

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617

		-1.0109289617



at 0.1Hz

at 1Hz

at 2Hz

at 3Hz

Time/sec.

Percentage Response

0.0032258065

-0.0021505376

0

0.0032258065

0

0

-0.0021505376

0.0032258065

0

0

-0.0021505376

0

0.0032258065

0

0.0086021505

0

-0.0021505376

0.0032258065

0.0032258065

0.0086021505

0

0

0.0032258065

0.0032258065

0.0032258065

0

0

0

0

-0.0021505376

0

0.0053763441

0

0

0.0032258065

0.0053763441

0

-0.0053763441

-0.0021505376

0.0053763441

0

0

0

0.0086021505

0.0053763441

0

0.0086021505

0

0.0032258065

0

0.0053763441

-0.0021505376

0

0

0.0032258065

0.0032258065

-0.0053763441

0.0086021505

0.0032258065

0

0.0032258065

0

0.0086021505

0

0

0

0

0

0

0.0086021505

0.0032258065

-0.0021505376

0

0

0.0032258065

0

0

0.0032258065

0

0

0.0032258065

0.0032258065

0

-0.0021505376

0

0

-0.0021505376

0.0032258065

0.0032258065

0

0.0086021505

0

0

0.0032258065

0

0

0

0

0.0032258065

0.0032258065

-0.0021505376

0.0032258065

0

0

0

0.0053763441

0

0.0053763441

0

0

0

0.0053763441

-0.0021505376

0

-0.0021505376

0

0.0086021505

0

0.0032258065

0.0032258065

0.0053763441

0

0

0

0

0

0.0032258065

0.0053763441

0

0

-0.0021505376

0

0

0.0032258065

0.0032258065

0.0032258065

0

-0.0021505376

0.0053763441

0

0

0.0053763441

0

-0.0021505376

0.0032258065

0.0032258065

0.0053763441

0.0032258065

0.0032258065

0.0053763441

0.0053763441

0.0032258065

0

0.0032258065

0

-0.0021505376

0.0032258065

0.0032258065

0

0.0053763441

0

0.0053763441

0.0032258065

0

0

-0.0021505376

0

0

0.0032258065

0.0053763441

0.0032258065

0.0032258065

0

0

0

0.0032258065

0.0086021505

0.0053763441

0.0086021505

0

0

0.0032258065

-0.0021505376

0.0053763441

0

0.0032258065

0.0032258065

0

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0

0

0

0.0032258065

0

0.0032258065

0.0032258065

0

0.0032258065

0

0

0

0.0032258065

0.0086021505

0

0

0.0032258065

0.0032258065

0

0.0053763441

0.0053763441

0.0032258065

0

0

0

0

0.0053763441

0

0.0053763441

0

0

0.0032258065

0

0.0032258065

0.0032258065

0

0.0053763441

0

0.0032258065

0

-0.0021505376

0.0053763441

0.0053763441

0

0.0053763441

0.0032258065

0.0053763441

0

0.0032258065

0

0

0

-0.0021505376

0.0053763441

0.0032258065

0.0032258065

0.0032258065

0.0086021505

0

0

0.0032258065

0.0032258065

0

-0.0021505376

0.0032258065

0

0

0.0086021505

0

0.0032258065

0

0

0

0

0.0032258065

0.0032258065

0

0

0.0032258065

0.0032258065

-0.0053763441

0

0.0086021505

0.0032258065

-0.0021505376

0.0032258065

0.0032258065

-0.0021505376

0

0

-0.0021505376

0.0086021505

-0.0021505376

0

-0.0021505376

0.0032258065

0.0032258065

0

0.0086021505

0

0.0032258065

0.0032258065

0.0086021505

0

0.0032258065

0.0086021505

-0.0021505376

0.0032258065

0.0032258065

0

0.0032258065

0

0

0

0

0.0053763441

0

-0.0021505376

0.0032258065

0

0.0032258065

0.0032258065

0.0032258065

0

0.0053763441

0

0

-0.0021505376

0.0053763441

0.0032258065

0.0086021505

-0.0021505376

0.0032258065

0

0.0032258065

0

0.0032258065

-0.0021505376

0.0032258065

0

0.0032258065

0.0032258065

0

0.0032258065

0.0032258065

0.0053763441

0

0.0032258065

0.0032258065

0

-0.0021505376

0

0.0086021505

0

0.0086021505

0

0.0032258065

-0.0021505376

0

0.0032258065

0

0

0

0

0

0

0.0032258065

0.0086021505

0

0.0032258065

0.0053763441

0

0.0032258065

0

0

0

0

0

0.0053763441

0.0053763441

0

0

0

-0.0021505376

0

-0.0021505376

0.0032258065

0

-0.0021505376

0

0

-0.0021505376

0

0

0

0.0053763441

0.0032258065

0

0

0.0032258065

0

-0.0021505376

-0.0021505376

0.0053763441

0

-0.0053763441

0.0053763441

0.0032258065

0

-0.0053763441

0.0032258065

0.0032258065

0.0032258065

0.0053763441

0

0

0.0053763441

0.0086021505

0

0

0.0032258065

0

0.0032258065

-0.0021505376

0.0032258065

-0.0021505376

0

0

0.0032258065

0.0032258065

0

0

0.0032258065

0.0032258065

0.0032258065

0

0.0032258065

0.0032258065

-0.0021505376

0

0.0032258065

0

-0.0021505376

0

0

0.0053763441

0.0032258065

0

0.0032258065

0

0.0053763441

0

0

0

0

0.0032258065

0.0032258065

0.0053763441

-0.0021505376

0

0

0.0032258065

0

0.0053763441

0.0032258065

-0.0021505376

0

-0.0021505376

0

0.0032258065

0.0053763441

-0.0021505376

0.0032258065

0

0

0.0032258065

0.0032258065

0

0.0053763441

0

0

0

0

0.0032258065

-0.0021505376

0

0.0032258065

0

0.0032258065

0

0

0.0032258065

0.0032258065

0.0032258065

-0.0021505376

0.0032258065

0

0.0032258065

-0.0021505376

0.0053763441

0.0086021505

0

-0.0021505376

0

0.0032258065

0

0

0.0053763441

-0.0021505376

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0

0.0053763441

0

-0.0021505376

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0

0

0

0.0053763441

0.0032258065

0.0053763441

0

0.0086021505

0.0032258065

0

0.0053763441

0

0.0032258065

0.0032258065

-0.0021505376

0

0.0032258065

0

0.0053763441

-0.0053763441

0.0032258065

0

0.0032258065

-0.0021505376

0.0032258065

0.0032258065

-0.0021505376

0

0

0

0.0032258065

0

0.0032258065

-0.0021505376

0.0032258065

0.0032258065

0

0.0086021505

0

0.0053763441

0

0

0

0

0

0

0.0086021505

0

0.0032258065

0

0

0.0032258065

0.0032258065

0.0032258065

0

-0.0021505376

0

0

0

-0.0021505376

0.0032258065

0.0032258065

0

0

0

0.0032258065

0.0053763441

0.0032258065

0.0032258065

0

0.0032258065

0.0053763441

0.0032258065

0.0032258065

0.0032258065

0

0

0.0032258065

0.0053763441

0.0053763441

0.0032258065

0

0

0

0.0032258065

0.0032258065

-0.0021505376

0

0

-0.0021505376

0.0032258065

0.0086021505

0.0032258065

-0.0021505376

0.0032258065

-0.0021505376

0.0032258065

0

0

0.0032258065

0.0086021505

0.0032258065

-0.0021505376

0.0107526882

0

-0.0021505376

0.0032258065

0.0032258065

0

0.0032258065

0.0053763441

0.0053763441

0

0.0032258065

0.0053763441

0

0.0032258065

0.0032258065

-0.0021505376

0

0

-0.0021505376

-0.0021505376

-0.0021505376

-0.0021505376

0

0

0

0

0

0.0086021505

0.0032258065

0.0053763441

0

0

0

0

0

-0.0021505376

0.0032258065

0

0

0

0

0

-0.0053763441

0.0032258065

0

0.0032258065

0.0053763441

0

0

-0.0021505376

0

0.0032258065

-0.0021505376

0.0053763441

0.0053763441

0

0.0053763441

0.0032258065

0

0.0032258065

-0.0021505376

0

-0.0053763441

0

0.0086021505

-0.0021505376

0

0

0

0

0

0

0.0032258065

0.0032258065

0.0032258065

0

0.0032258065

0.0032258065

0

0

0

0

0.0032258065

0.0032258065

0

0

0.0032258065

0

0

0

0.0053763441

0.0086021505

0

0.0032258065

0.0032258065

0.0032258065

0.0086021505

0.0086021505

0.0086021505

0.0129032258

0.0215053763

0.0290322581

0.0451612903

0.0688172043

0.0870967742

0.1107526882

0.1397849462

0.1709677419

0.2021505376

0.2311827957

0.2655913978

0.2989247312

0.3311827957

0.3655913978

0.4021505376

0.4279569892

0.4569892473

0.4806451613

0.5150537634

0.5387096774

0.5569892473

0.5827956989

0.6010752688

0.6172043011

0.635483871

0.6591397849

0.6720430108

0.6903225806

0.6989247312

0.7139784946

0.7268817204

0.735483871

0.7483870968

0.7591397849

0.7698924731

0.7774193548

0.788172043

0.7956989247

0.8086021505

0.8161290323

0.8193548387

0.8247311828

0.8322580645

0.8376344086

0.8505376344

0.8451612903

0.8591397849

0.8591397849

0.8666666667

0.8666666667

0.8720430108

0.8741935484

0.8795698925

0.8817204301

0.8903225806

0.8870967742

0.8924731183

0.8978494624

0.9107526882

0.9053763441

0.9086021505

0.9086021505

0.9107526882

0.9215053763

0.9161290323

0.9161290323

0.9215053763

0.9268817204

0.9268817204

0.9268817204

0.9344086022

0.9322580645

0.9290322581

0.9430107527

0.9397849462

0.9376344086

0.9397849462

0.9397849462

0.9430107527

0.9451612903

0.9451612903

0.9483870968

0.9483870968

0.952688172

0.9483870968

0.952688172

0.9580645161

0.9505376344

0.9580645161

0.9559139785

0.9559139785

0.9580645161

0.9580645161

0.9634408602

0.9612903226

0.9634408602

0.9634408602

0.9612903226

0.9634408602

0.9688172043

0.9666666667

0.9612903226

0.9688172043

0.9666666667

0.9688172043

0.9741935484

0.9741935484

0.9688172043

0.9688172043

0.9709677419

0.9709677419

0.976344086

0.9741935484

0.9741935484

0.9709677419

0.9741935484

0.976344086

0.9741935484

0.9795698925

0.976344086

0.9817204301

0.9795698925

0.976344086

0.976344086

0.976344086

0.9795698925

0.976344086

0.9741935484

0.9817204301

0.9795698925

0.9795698925

0.9795698925

0.9849462366

0.9795698925

0.9849462366

0.9795698925

0.9817204301

0.9817204301

0.9849462366

0.9849462366

0.9849462366

0.9795698925

0.9849462366

0.9795698925

0.9817204301

0.9817204301

0.9849462366

0.9817204301

0.9849462366

0.9903225806

0.9849462366

0.9849462366

0.9870967742

0.9849462366

0.9849462366

0.9817204301

0.9870967742

0.9849462366

0.9817204301

0.9849462366

0.9849462366

0.9870967742

0.9849462366

0.9870967742

0.9849462366

0.9903225806

0.9903225806

0.9946236559

0.9924731183

0.9924731183

0.9870967742

0.9870967742

0.9870967742

0.9870967742

0.9903225806

0.9924731183

0.9903225806

0.9903225806

0.9903225806

0.9870967742

0.9870967742

0.9870967742

0.9946236559

0.9903225806

0.9903225806

0.9903225806

0.9903225806

0.9946236559

0.9924731183

0.9903225806

0.9903225806

0.9924731183

0.9946236559

0.9903225806

0.9946236559

0.9946236559

0.9924731183

0.9903225806

0.9903225806

0.9946236559

0.9924731183

0.9924731183

0.9903225806

0.9903225806

0.9924731183

0.9946236559

1

0.9924731183

0.9924731183

0.9903225806

0.9924731183

0.9924731183

0.9903225806

0.9978494624

0.9924731183

0.9924731183

0.9978494624

0.9946236559

0.9924731183

0.9924731183

0.9903225806

0.9978494624

0.9903225806

0.9924731183

0.9946236559

0.9946236559

0.9924731183

0.9946236559

0.9946236559

0.9946236559

0.9924731183

0.9978494624

0.9924731183

0.9978494624

0.9946236559

0.9924731183

0.9924731183

0.9903225806

0.9978494624

0.9946236559

0.9946236559

0.9903225806

0.9924731183

0.9946236559

0.9946236559

0.9946236559

0.9946236559

0.9924731183

0.9946236559

0.9978494624

0.9946236559

0.9924731183

1

0.9946236559

0.9946236559

0.9946236559

0.9924731183

0.9946236559

0.9978494624

0.9978494624

0.9946236559

0.9978494624

0.9946236559

0.9924731183

0.9946236559

0.9978494624

0.9946236559

0.9978494624

0.9978494624

0.9946236559

0.9978494624

0.9946236559

0.9978494624

0.9978494624

0.9946236559

0.9946236559

0.9924731183

0.9946236559

0.9946236559

0.9946236559

0.9978494624

0.9946236559

0.9978494624

0.9978494624

0.9978494624

0.9978494624

0.9946236559

0.9946236559

0.9946236559

0.9946236559

0.9946236559

0.9978494624

0.9978494624

1

0.9978494624

0.9946236559

0.9946236559

1

1

0.9946236559

1

0.9978494624

0.9946236559

1

0.9924731183

1

0.9946236559

0.9978494624

1

0.9946236559

0.9924731183

0.9978494624

0.9924731183

0.9978494624

1

0.9924731183

0.9946236559

0.9978494624

0.9978494624

0.9978494624

0.9978494624

0.9946236559

1.0032258065

0.9946236559

1

1

0.9946236559

0.9946236559

1.0032258065

1

0.9946236559

0.9978494624

0.9978494624

0.9946236559

0.9978494624

0.9978494624

0.9946236559

1

1

0.9978494624

1.0032258065

1.0032258065

1

1

0.9978494624

0.9946236559

1

0.9946236559

0.9946236559

0.9978494624

0.9978494624

0.9978494624

0.9946236559

0.9946236559

0.9978494624

0.9946236559

1.0032258065

0.9978494624

0.9978494624

0.9946236559

0.9978494624

1.0032258065

1

0.9978494624

0.9978494624

0.9978494624

1

1

0.9946236559

0.9946236559

0.9978494624

0.9946236559

0.9978494624

0.9946236559

0.9978494624

1.0032258065

1

1

0.9978494624

0.9978494624

0.9978494624

0.9978494624

0.9946236559

1

1

0.9946236559

1.0053763441

0.9946236559

0.9978494624

0.9978494624

0.9946236559

0.9978494624

0.9978494624

0.9978494624

0.9978494624

0.9978494624

0.9946236559

0.9978494624

1.0032258065

0.9978494624

0.9978494624

1.0032258065

0.9978494624

0.9978494624

0.9946236559

0.9978494624

0.9978494624

0.9978494624

1.0032258065

1

1.0032258065

0.9978494624

1

0.9946236559

0.9978494624

0.9978494624

1.0032258065

0.9946236559

1

1.0053763441

1.0032258065

0.9978494624

1

1

0.9978494624

1

0.9978494624

0.9978494624

1

1

0.9978494624

0.9978494624

1

1.0053763441

0.9978494624

0.9978494624

0.9978494624

0.9978494624

1.0032258065

0.9946236559

1.0032258065

0.9978494624

0.9978494624

0.9978494624

0.9978494624

1

1

1

0.9978494624

0.9978494624

0.9978494624

0.9978494624

0.9946236559

1.0032258065

0.9978494624

0.9978494624

0.9978494624

1

0.9978494624

1

0.9978494624

0.9946236559

1

1

0.9978494624

1

1

1

1

1

1

1.0032258065

1.0053763441

0.9978494624

1

0.9946236559

1

1

0.9978494624

0.9978494624

1

1

1

1

1.0032258065

1

0.9978494624

0.9978494624

0.9978494624

1.0032258065

1.0032258065

0.9978494624

0.9978494624

0.9978494624

1

0.9978494624

0.9978494624

0.9978494624

0.9978494624

0.9978494624

1.0032258065

0.9978494624

1

1.0053763441

1

1

0.9978494624

0.9978494624

0.9978494624

1

1

0.9978494624

1

0.9978494624

0.9978494624

0.9978494624

1

1.0032258065

0.9978494624

1.0053763441

0.9978494624

0.9978494624

1

0.9978494624

1

1.0032258065

1.0053763441

1.0053763441

1

1.0032258065

1

0.9978494624

1

0.9978494624

1.0032258065

0.9978494624

1

0.9946236559

1

1

1.0032258065

1

0.9978494624

1.0032258065

0.9978494624

0.9978494624

0.9978494624

0.9978494624

1

1.0053763441

1.0032258065

1

1.0032258065

0.9978494624

0.9978494624

1

1

0.9978494624

1

1

1

1

1

1.0032258065

1

1.0053763441

1.0053763441

1

1.0032258065

1.0053763441

1.0032258065

1

1

0.9978494624

1

1.0032258065

1

1.0032258065

1.0053763441

1.0053763441

1.0032258065

0.9978494624

1

1.0032258065

1.0053763441

1

1

1

1.0053763441

1

1

1.0053763441

1.0032258065

1.0053763441

1

1

0.9978494624

1.0053763441

1

1

1.0053763441

1

1

1

1.0032258065

1.0032258065

1

0.9978494624

1

0.9978494624

1

0.9978494624

1

0.9946236559

0.9946236559

0.9870967742

0.9817204301

0.9688172043

0.9559139785

0.9344086022

0.9161290323

0.8817204301

0.8537634409

0.8268817204

0.7849462366

0.7591397849

0.7193548387

0.6903225806

0.6516129032

0.6225806452

0.588172043

0.5569892473

0.5301075269

0.5010752688

0.4752688172

0.4548387097

0.4387096774

0.4150537634

0.3860215054

0.3698924731

0.3516129032

0.3365591398

0.3225806452

0.3096774194

0.2967741935

0.2806451613

0.2677419355

0.2569892473

0.2494623656

0.2387096774

0.223655914

0.2204301075

0.2075268817

0.1967741935

0.1946236559

0.1870967742

0.1784946237

0.1731182796

0.1655913978

0.1602150538

0.1580645161

0.1494623656

0.1440860215

0.1397849462

0.1365591398

0.1311827957

0.1258064516

0.123655914

0.123655914

0.1161290323

0.1107526882

0.1107526882

0.1021505376

0.1021505376

0.1021505376

0.0978494624

0.0892473118

0.0892473118

0.0870967742

0.0838709677

0.0838709677

0.0817204301

0.0784946237

0.0817204301

0.076344086

0.0741935484

0.0688172043

0.0709677419

0.0709677419

0.0709677419

0.0655913978

0.0688172043

0.0634408602

0.0602150538

0.0548387097

0.0548387097

0.0602150538

0.0548387097

0.0548387097

0.052688172

0.0505376344

0.047311828

0.0548387097

0.0505376344

0.047311828

0.0419354839

0.0505376344

0.0451612903

0.0419354839

0.0419354839

0.0397849462

0.0419354839

0.0419354839

0.0397849462

0.0451612903

0.0397849462

0.0365591398

0.0365591398

0.0365591398

0.0311827957

0.0365591398

0.0344086022

0.0397849462

0.0365591398

0.0290322581

0.0344086022

0.0311827957

0.0344086022

0.0344086022

0.0290322581

0.023655914

0.0268817204

0.0268817204

0.0290322581

0.0344086022

0.0268817204

0.0268817204

0.023655914

0.0215053763

0.023655914

0.023655914

0.0268817204

0.0268817204

0.0215053763

0.0268817204

0.0268817204

0.023655914

0.023655914

0.0268817204

0.0215053763

0.023655914

0.0215053763

0.023655914

0.0215053763

0.0182795699

0.0182795699

0.023655914

0.0215053763

0.0161290323

0.0182795699

0.0161290323

0.0215053763

0.0182795699

0.0182795699

0.0182795699

0.0215053763

0.0182795699

0.0182795699

0.0161290323

0.0182795699

0.0129032258

0.0215053763

0.0182795699

0.0161290323

0.0182795699

0.0182795699

0.0129032258

0.0215053763

0.0182795699

0.0182795699

0.0161290323

0.0182795699

0.0182795699

0.0182795699

0.0129032258

0.0129032258

0.0107526882

0.0161290323

0.0182795699

0.0161290323

0.0129032258

0.0129032258

0.0161290323

0.0129032258

0.0129032258

0.0182795699

0.0129032258

0.0129032258

0.0161290323

0.0129032258

0.0129032258

0.0161290323

0.0107526882

0.0107526882

0.0129032258

0.0129032258

0.0107526882

0.0182795699

0.0129032258

0.0086021505

0.0129032258

0.0107526882

0.0107526882

0.0161290323

0.0129032258

0.0161290323

0.0129032258

0.0053763441

0.0107526882

0.0129032258

0.0161290323

0.0129032258

0.0107526882

0.0107526882

0.0107526882

0.0086021505

0.0086021505

0.0129032258

0.0086021505

0.0129032258

0.0053763441

0.0086021505

0.0129032258

0.0161290323

0.0129032258

0.0107526882

0.0053763441

0.0086021505

0.0086021505

0.0086021505

0.0107526882

0.0086021505

0.0161290323

0.0086021505

0.0107526882

0.0129032258

0.0129032258

0.0129032258

0.0107526882

0.0086021505

0.0086021505

0.0129032258

0.0053763441

0.0086021505

0.0107526882

0.0086021505

0.0086021505

0.0086021505

0.0129032258

0.0107526882

0.0107526882

0.0086021505

0.0086021505

0.0107526882

0.0129032258

0.0129032258

0.0053763441

0.0086021505

0.0053763441

0.0086021505

0.0086021505

0.0107526882

0.0032258065

0.0086021505

0.0053763441

0.0129032258

0.0086021505

0.0107526882

0.0086021505

0.0086021505

0.0053763441

0.0053763441

0.0053763441

0.0107526882

0.0107526882

0.0086021505

0.0053763441

0.0053763441

0.0053763441

0.0086021505

0.0086021505

0.0053763441

0.0086021505

0.0032258065

0.0107526882

0.0053763441

0.0053763441

0.0086021505

0

0.0053763441

0.0053763441

0.0053763441

0.0032258065

0.0053763441

0.0086021505

0

0.0053763441

0.0086021505

0.0086021505

0.0032258065

0.0053763441

0.0053763441

0.0053763441

0.0053763441

0.0086021505

0.0053763441

0.0086021505

0.0053763441

0.0032258065

0.0032258065

0.0086021505

0.0086021505

0.0053763441

0.0086021505

0.0086021505

0.0032258065

0.0032258065

0.0053763441

0.0032258065

0.0053763441

0.0053763441

0.0053763441

0.0053763441

0.0053763441

0.0053763441

0.0107526882

0.0086021505

0.0032258065

0.0107526882

0.0086021505

0.0086021505

0.0107526882

0

0.0053763441

0.0053763441

0.0032258065

0.0107526882

0.0053763441

0.0032258065

0.0053763441

0.0053763441

0.0032258065

0.0107526882

0.0032258065

0

0

0

0.0086021505

0.0053763441

0.0053763441

0.0032258065

0

0.0086021505

0.0032258065

0.0053763441

0.0032258065

0.0086021505

0.0053763441

0.0129032258

0.0032258065

0.0086021505

0.0053763441

0.0053763441

0.0053763441

0

0.0032258065

0.0032258065

0.0032258065

0.0053763441

0.0053763441

0.0053763441

0.0086021505

0.0086021505

0.0086021505

0.0032258065

0.0032258065

0.0032258065

0.0053763441

0.0107526882

0.0086021505

0.0053763441

0.0053763441

0.0032258065

0.0053763441

0.0107526882

0.0032258065

0.0365591398

0.0053763441

0.0032258065

0.0032258065

0.0053763441

0.0053763441

0.0032258065

0.0032258065

0.0053763441

0.0086021505

0.0086021505

0.0053763441

0.0032258065

0

0.0032258065

0.0032258065

0.0032258065

0.0107526882

0.0053763441

0.0107526882

0.0032258065

0

0.0032258065

0

0.0053763441

0.0032258065

0.0032258065

0.0086021505

0.0032258065

0.0053763441

0.0032258065

0.0053763441

0.0053763441

0.0107526882

0.0053763441

0.0086021505

0.0107526882

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0053763441

0.0032258065

0.0053763441

0.0032258065

0.0032258065

0.0086021505

0.0053763441

0.0053763441

0.0032258065

0.0032258065

0.0032258065

0.0086021505

0.0032258065

0.0032258065

0.0053763441

0.0032258065

0.0032258065

0.0086021505

0.0107526882

0.0053763441

0.0086021505

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0053763441

0.0032258065

0.0086021505

0.0032258065

0.0032258065

0.0032258065

0

0.0032258065

0.0032258065

0.0032258065

0

0.0032258065

0.0032258065

0.0032258065

0

0.0032258065

0.0107526882

0.0032258065

0.0086021505

0.0053763441

0.0053763441

0.0032258065

0.0032258065

0.0086021505

0.0032258065

0.0032258065

0

0.0032258065

0.0053763441

0.0086021505

0.0086021505

0.0053763441

0.0053763441

0.0032258065

0

0.0107526882

0.0032258065

0.0032258065

0.0086021505

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0053763441

0

0.0032258065

0.0032258065

0.0086021505

0.0053763441

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0086021505

0.0086021505

0.0032258065

0.0032258065

0.0032258065

0

0.0032258065

0.0053763441

0.0086021505

0.0032258065

0.0053763441

0.0032258065

0.0032258065

0.0032258065

0

-0.0021505376

0.0032258065

-0.0021505376

0.0086021505

0.0032258065

0.0032258065

0.0086021505

0

0.0053763441

0.0032258065

0.0086021505

0.0053763441

0.0086021505

0.0053763441

0.0032258065

0.0086021505

0

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0

0

0

0.0032258065

0.0032258065

0.0032258065

0

0.0032258065

0.0053763441

0.0032258065

0.0086021505

0

0

0.0032258065

0.0086021505

0.0032258065

0.0032258065

0.0053763441

0.0086021505

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0086021505

0.0086021505

0

0.0053763441

0.0032258065

0.0086021505

0.0053763441

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0032258065

0.0053763441

0.0032258065

0.0032258065

0.0032258065

0.0053763441

0.0032258065

0

0

0.0032258065

0.0086021505

0.0032258065

0.0053763441

0.0032258065

0.0032258065

0.0086021505

0

0.0086021505

0.0032258065

0.0053763441

0.0053763441

0.0086021505

0.0129032258

0.0215053763

0.0311827957

0.0419354839

0.0602150538

0.0784946237

0.1075268817

0.1311827957

0.1602150538

0.1967741935

0.2311827957

0.2655913978

0.2989247312

0.3258064516

0.3602150538

0.3913978495

0.4204301075

0.4494623656

0.4806451613

0.5064516129

0.5333333333

0.5537634409

0.5752688172

0.5989247312

0.6172043011

0.6333333333

0.6516129032

0.6698924731

0.6903225806

0.6956989247

0.711827957

0.7247311828

0.7408602151

0.7483870968

0.7612903226

0.7698924731

0.7774193548

0.788172043

0.7956989247

0.8032258065

0.8086021505

0.8193548387

0.8268817204

0.8301075269

0.8376344086

0.8451612903

0.8505376344

0.8505376344

0.8688172043

0.8666666667

0.8741935484

0.8774193548

0.8817204301

0.8795698925

0.8849462366

0.8903225806

0.8956989247

0.8956989247

0.8978494624

0.9032258065

0.9139784946

0.9086021505

0.9107526882

0.9161290323

0.9161290323

0.9215053763

0.9215053763

0.9247311828

0.9247311828

0.9247311828

0.9268817204

0.9322580645

0.9322580645

0.9344086022

0.9430107527

0.9344086022

0.9451612903

0.9430107527

0.9430107527

0.9483870968

0.9430107527

0.9451612903

0.9483870968

0.9483870968

0.9505376344

0.9559139785

0.952688172

0.952688172

0.952688172

0.9580645161

0.9580645161

0.9612903226

0.9559139785

0.9612903226

0.9580645161

0.9612903226

0.9580645161

0.9634408602

0.9634408602

0.9666666667

0.9634408602

0.9666666667

0.9634408602

0.9688172043

0.9666666667

0.9709677419

0.9688172043

0.9688172043

0.9709677419

0.9688172043

0.9709677419

0.9709677419

0.9741935484

0.9741935484

0.9709677419

0.976344086

0.9741935484

0.9741935484

0.9795698925

0.9795698925

0.976344086

0.976344086

0.9849462366

0.9795698925

0.976344086

0.976344086

0.976344086

0.9795698925

0.9849462366

0.9849462366

0.9795698925

0.9817204301

0.9795698925

0.9817204301

0.9795698925

0.9817204301

0.9817204301

0.9817204301

0.9817204301

0.9817204301

0.9795698925

0.9817204301

0.9903225806

0.9849462366

0.9903225806

0.9817204301

0.9849462366

0.9849462366

0.9849462366

0.9903225806

0.9817204301

0.9903225806

0.9849462366

0.9924731183

0.9903225806

0.9903225806

0.9870967742

0.9849462366

0.9903225806

0.9870967742

0.9903225806

0.9946236559

0.9924731183

0.9817204301

0.9924731183

0.9870967742

0.9870967742

0.9924731183

0.9849462366

0.9924731183

0.9849462366

0.9903225806

0.9870967742

0.9870967742

0.9903225806

0.9924731183

0.9903225806

0.9903225806

0.9978494624

0.9903225806

0.9924731183

0.9946236559

0.9903225806

0.9946236559

0.9946236559

0.9903225806

0.9903225806

0.9924731183

0.9924731183

0.9903225806

0.9903225806

0.9903225806

0.9903225806

0.9903225806

0.9903225806

0.9903225806

0.9924731183

0.9946236559

0.9924731183

0.9946236559

0.9924731183

0.9924731183

0.9924731183

0.9924731183

0.9946236559

0.9924731183

0.9924731183

0.9946236559

0.9924731183

0.9903225806

0.9946236559

0.9924731183

0.9946236559

0.9924731183

0.9924731183

0.9924731183

0.9946236559

0.9946236559

0.9946236559

0.9946236559

0.9924731183

0.9946236559

1

0.9946236559

0.9978494624

0.9946236559

0.9978494624

0.9946236559

0.9946236559

1

0.9978494624

0.9924731183

0.9978494624

0.9924731183

0.9946236559

0.9946236559

0.9946236559

0.9946236559

0.9924731183

0.9978494624

0.9946236559

0.9978494624

0.9924731183

0.9946236559

0.9924731183

0.9946236559

0.9924731183

0.9978494624

1.0032258065

1

0.9946236559

1

0.9946236559

1

1

0.9924731183

0.9978494624

0.9946236559

0.9946236559

0.9946236559

0.9946236559

0.9946236559

0.9978494624

0.9924731183

0.9978494624

1

0.9978494624

1

0.9946236559

0.9946236559

0.9946236559

0.9946236559

0.9946236559

1.0053763441

0.9924731183

0.9946236559

1

0.9978494624

0.9924731183

0.9946236559

0.9946236559

1

0.9978494624

0.9946236559

1.0032258065

0.9978494624

0.9946236559

0.9978494624

0.9946236559

1

0.9946236559

0.9924731183

1

0.9978494624

1.0032258065

1

0.9978494624

0.9946236559

0.9946236559

1

0.9978494624

0.9946236559

0.9946236559

0.9978494624

0.9978494624

0.9978494624

0.9978494624

0.9946236559

1

0.9946236559

0.9978494624

0.9978494624

0.9946236559

1

0.9978494624

1

0.9946236559

1

0.9978494624

0.9978494624

0.9946236559

1

1

1

0.9978494624

1.0053763441

0.9924731183

0.9946236559

1.0032258065

1

1.0032258065

0.9978494624

1.0032258065

0.9946236559

1

1.0032258065

0.9978494624

1

0.9978494624

0.9978494624

0.9978494624

0.9978494624

0.9978494624

0.9946236559

1

1

1.0032258065

1.0053763441

1

0.9978494624

0.9946236559

1

1.0032258065

0.9978494624

1

0.9978494624

1.0032258065

1

1

0.9946236559

0.9978494624

0.9978494624

1

0.9978494624

1

1

1

1

0.9946236559

1.0053763441

0.9978494624

1.0053763441

1.0053763441

1

0.9978494624

0.9946236559

0.9978494624

1.0053763441

1

1

1.0032258065

0.9978494624

0.9978494624

0.9978494624

0.9978494624

1.0032258065

0.9978494624

1.0032258065

0.9978494624

1

0.9978494624

0.9978494624

0.9978494624

0.9978494624

1.0053763441

1

1.0032258065

1

0.9946236559

1

1.0053763441

0.9978494624

1

0.9978494624

0.9978494624

0.9946236559

1

1.0032258065

0.9978494624

0.9946236559

1

1

0.9978494624

1

0.9978494624

1.0053763441

1

1

1

1

1.0053763441

0.9978494624

1

1.0053763441

0.9978494624

1.0032258065

1

0.9978494624

1

1

0.9978494624

1

0.9978494624

1.0053763441

0.9978494624

1

1

0.9978494624

0.9978494624

0.9978494624

1

0.9978494624

0.9978494624

0.9978494624

1

1

1

1

1

1

1

1

0.9978494624

1

1

1

1

1

1

1

1

1.0032258065

1.0053763441

0.9978494624

1

1.0032258065

0.9978494624

1

1

1

1

0.9978494624

1

1

1.0053763441

0.9978494624

1

1

1.0053763441

1

1.0053763441

1

1

0.9978494624

1.0032258065

1

1

0.9978494624

1.0053763441

1

1

0.9978494624

1

0.9978494624

1.0107526882

1.0053763441

1

1

0.9978494624

0.9978494624

1

1

1

1

1.0053763441

1.0053763441

1

1

0.9978494624

0.9978494624

1

1.0032258065

1

0.9978494624

0.9978494624

1

1

1.0053763441

0.9978494624

1

1

1

1.0032258065

0.9978494624

1

1.0053763441

1

1

1

0.9978494624

1

1.0053763441

1.0053763441

1

0.9978494624

1

0.9978494624

0.9978494624

1

1

1.0053763441

1.0053763441

1.0053763441

1.0032258065

1.0053763441

1.0032258065

1.0032258065

1

1

1

1

1

1

1

1

1

1.0032258065

1.0032258065

1.0032258065

1.0053763441

1

1.0032258065

1

1.0032258065

1

0.9978494624

1

1

1.0032258065

1

1

1.0032258065

1.0032258065

1.0053763441

1

1

1

1

1.0107526882

1.0032258065

1

1

1

0.9978494624

1

1

1.0086021505

1.0053763441

1.0053763441

1

1.0086021505

0.9978494624

1

1.0086021505

0.9978494624

1.0032258065

1

1.0086021505

1

1.0053763441

0.9978494624

1.0053763441

1.0086021505

0.9978494624

1.0032258065

1.0032258065

1.0032258065

1

1

0.9978494624

1.0053763441

0.9978494624

0.9924731183

0.9870967742

0.9741935484

0.9688172043

0.9483870968

0.9215053763

0.8870967742

0.8666666667

0.835483871

0.7956989247

0.7612903226

0.7322580645

0.6956989247

0.6612903226

0.6301075269

0.5956989247

0.564516129

0.535483871

0.5096774194

0.4860215054

0.4623655914

0.4387096774

0.4258064516

0.3989247312

0.3752688172

0.3602150538

0.3494623656

0.3258064516

0.3129032258

0.2967741935

0.2860215054

0.2731182796

0.2623655914

0.2494623656

0.2365591398

0.2290322581

0.2204301075

0.2150537634

0.1967741935

0.1946236559

0.1892473118

0.1817204301

0.176344086

0.1655913978

0.1602150538

0.1548387097

0.147311828

0.1494623656

0.1419354839

0.1419354839

0.1311827957

0.1258064516

0.123655914

0.1182795699

0.1182795699

0.1161290323

0.1107526882

0.1053763441

0.1053763441

0.1075268817

0.0978494624

0.0946236559

0.0870967742

0.0870967742

0.0838709677

0.0838709677

0.0784946237

0.0817204301

0.076344086

0.0741935484

0.0741935484

0.0741935484

0.0741935484

0.0709677419

0.0688172043

0.0709677419

0.0634408602

0.0634408602

0.0634408602

0.0580645161

0.0580645161

0.0548387097

0.0580645161

0.052688172

0.0580645161

0.0548387097

0.052688172

0.0505376344

0.047311828

0.047311828

0.047311828

0.047311828

0.047311828

0.0451612903

0.0451612903

0.047311828

0.0419354839

0.0419354839

0.0451612903

0.0505376344

0.0602150538

0.0655913978

0.0784946237

0.0924731183

0.1129032258

0.1397849462

0.1634408602

0.1946236559

0.223655914

0.2623655914

0.2946236559

0.3225806452

0.3569892473

0.3913978495

0.4172043011

0.4494623656

0.4784946237

0.5043010753

0.5301075269

0.5537634409

0.5806451613

0.5989247312

0.6225806452

0.6408602151

0.6537634409

0.6806451613

0.6903225806

0.7064516129

0.7225806452

0.7322580645

0.7483870968

0.7559139785

0.7666666667

0.7795698925

0.7903225806

0.7978494624

0.8086021505

0.8161290323

0.8193548387

0.8301075269

0.8376344086

0.8430107527

0.8483870968

0.8537634409

0.8666666667

0.8688172043

0.8688172043

0.8741935484

0.8795698925

0.8849462366

0.8849462366

0.8903225806

0.8903225806

0.8795698925

0.8720430108

0.8634408602

0.8505376344

0.8268817204

0.8086021505

0.7774193548

0.7462365591

0.7193548387

0.688172043

0.6591397849

0.6225806452

0.5989247312

0.564516129

0.5387096774

0.5096774194

0.4784946237

0.4516129032

0.435483871

0.4043010753

0.3838709677

0.3655913978

0.3462365591

0.3279569892

0.3129032258

0.2989247312

0.2784946237

0.2677419355

0.2569892473

0.247311828

0.2311827957

0.223655914

0.2129032258

0.2021505376

0.1946236559

0.1817204301

0.1784946237

0.1731182796

0.1602150538

0.1548387097

0.147311828

0.1419354839

0.1397849462

0.1344086022

0.1344086022

0.1290322581

0.1290322581

0.1311827957

0.1397849462

0.1494623656

0.1602150538

0.1784946237

0.2075268817

0.223655914

0.2569892473

0.2806451613

0.3129032258

0.3440860215

0.3731182796

0.4096774194

0.4311827957

0.4591397849

0.4935483871

0.5247311828

0.5440860215

0.5677419355

0.5903225806

0.611827957

0.6333333333

0.6569892473

0.6698924731

0.6849462366

0.7032258065

0.7247311828

0.7322580645

0.7430107527

0.7591397849

0.7698924731

0.7795698925

0.788172043

0.8010752688

0.8086021505

0.8161290323

0.8301075269

0.8376344086

0.8397849462

0.8483870968

0.8537634409

0.8591397849

0.8688172043

0.8720430108

0.8741935484

0.8817204301

0.8849462366

0.8903225806

0.8903225806

0.8956989247

0.8903225806

0.9010752688

0.8924731183

0.8849462366

0.8741935484

0.8559139785

0.8376344086

0.8139784946

0.788172043

0.7612903226

0.7268817204

0.6956989247

0.6698924731

0.6376344086

0.6043010753

0.5720430108

0.5408602151

0.511827957

0.488172043

0.464516129

0.435483871

0.411827957

0.3892473118

0.3698924731

0.3516129032

0.3333333333

0.3225806452

0.3021505376

0.2913978495

0.2731182796

0.2602150538

0.252688172

0.2311827957

0.2258064516

0.2150537634

0.2053763441

0.1967741935

0.1892473118

0.176344086

0.1731182796

0.1677419355

0.1602150538

0.1548387097

0.1440860215

0.1419354839

0.1365591398

0.1365591398

0.1311827957

0.1311827957

0.1311827957

0.1397849462

0.147311828

0.1634408602

0.176344086

0.1967741935

0.223655914

0.252688172

0.2784946237

0.3075268817

0.3419354839

0.3698924731

0.4021505376

0.4333333333

0.464516129

0.4860215054

0.5150537634

0.5387096774

0.5698924731

0.5935483871

0.611827957

0.6301075269

0.6537634409

0.6698924731

0.6827956989

0.7010752688

0.7172043011

0.7322580645

0.7483870968

0.7537634409

0.7666666667

0.7774193548

0.7903225806

0.7978494624

0.8064516129

0.8161290323

0.8247311828

0.8322580645

0.8430107527

0.8430107527

0.8537634409

0.8559139785

0.8634408602

0.8741935484

0.8774193548

0.8795698925

0.8870967742

0.8924731183

0.8903225806

0.8903225806

0.8903225806

0.8924731183

0.8924731183

0.8849462366

0.8795698925

0.8591397849

0.8451612903

0.8161290323

0.7956989247

0.7612903226

0.735483871

0.7010752688

0.6698924731

0.6408602151

0.6043010753

0.5752688172

0.5462365591

0.5172043011

0.488172043

0.464516129

0.4408602151

0.4172043011

0.3935483871

0.3784946237

0.3548387097

0.3387096774

0.3225806452

0.3075268817

0.2913978495

0.2784946237

0.2623655914

0.2494623656

0.2387096774

0.2258064516

0.2182795699

0.2096774194

0.2

0.1913978495

0.1817204301

0.1731182796

0.1677419355

0.1655913978

0.152688172

0.147311828

0.1440860215

0.1365591398

0.1397849462

0.1311827957

0.1344086022

0.1311827957

0.1365591398

0.1419354839

0.152688172

0.1709677419

0.1913978495

0.2075268817

0.2333333333

0.2623655914

0.2967741935

0.3279569892

0.3569892473

0.3838709677

0.4172043011

0.4462365591

0.4731182796

0.5010752688

0.535483871

0.5569892473

0.5774193548

0.5989247312

0.6225806452

0.6430107527

0.6569892473

0.6774193548

0.6956989247

0.7064516129

0.7268817204

0.7430107527

0.7537634409

0.7698924731

0.7741935484

0.7849462366

0.7924731183

0.811827957

0.811827957

0.8215053763

0.8268817204

0.835483871

0.8451612903

0.8483870968

0.8591397849

0.8634408602

0.8688172043

0.8720430108

0.8795698925

0.8817204301

0.8870967742

0.8924731183

0.8924731183

0.8924731183

0.8978494624

0.9032258065

0.8956989247

0.8817204301

0.8688172043

0.8537634409

0.8322580645

0.8086021505

0.7795698925

0.7505376344

0.7193548387

0.688172043

0.6612903226

0.6247311828

0.5903225806

0.5623655914

0.5333333333

0.5043010753

0.4784946237

0.4569892473

0.4279569892

0.4096774194

0.3860215054

0.3677419355

0.3462365591

0.3311827957

0.3096774194

0.2989247312

0.2806451613

0.2709677419

0.2548387097

0.2440860215

0.2311827957

0.223655914

0.2129032258

0.2021505376

0.1946236559

0.1892473118

0.1838709677

0.176344086

0.1602150538

0.1548387097

0.152688172

0.1440860215

0.1440860215

0.1365591398

0.1365591398

0.1311827957

0.1290322581

0.1365591398

0.1397849462

0.1494623656

0.1677419355

0.1817204301

0.2021505376

0.2333333333

0.2569892473

0.2892473118

0.3182795699

0.3494623656

0.3860215054

0.411827957

0.4440860215

0.4698924731

0.5043010753

0.5225806452

0.5462365591

0.5720430108

0.5935483871

0.6172043011

0.6376344086

0.6537634409

0.6720430108

0.6956989247

0.7032258065

0.7193548387

0.7376344086

0.7483870968

0.764516129

0.7720430108

0.7795698925

0.7903225806

0.8032258065

0.811827957

0.8215053763

0.8322580645

0.835483871

0.8397849462

0.8483870968

0.8537634409

0.8612903226

0.8666666667

0.8720430108

0.8774193548

0.8795698925

0.8817204301

0.8903225806

0.8903225806

0.8956989247

0.9010752688

0.8956989247

0.8924731183

0.8870967742

0.8741935484

0.8612903226

0.8376344086

0.8215053763

0.7924731183

0.7666666667

0.735483871

0.7032258065

0.6720430108

0.6376344086

0.6064516129

0.5774193548

0.5462365591

0.5172043011

0.4913978495

0.464516129

0.4387096774

0.4172043011

0.3967741935

0.3752688172

0.3548387097

0.3333333333

0.3204301075

0.3043010753

0.2913978495

0.2806451613

0.2623655914

0.2494623656

0.2333333333

0.2290322581

0.2182795699

0.2075268817

0.2

0.1892473118

0.1817204301

0.1731182796

0.1677419355

0.1602150538

0.1548387097

0.147311828

0.1440860215

0.1365591398

0.1311827957

0.1290322581

0.1258064516

0.1290322581

0.1344086022

0.1344086022

0.152688172

0.1602150538

0.1731182796

0.2

0.2150537634

0.2440860215

0.2731182796

0.3043010753

0.3333333333

0.3655913978

0.3935483871

0.4258064516

0.4548387097

0.4827956989

0.5150537634

0.535483871

0.5591397849

0.5860215054

0.6043010753

0.6279569892

0.6462365591

0.6666666667

0.6827956989

0.6989247312

0.711827957

0.7301075269

0.7462365591

0.7559139785

0.764516129

0.7774193548

0.788172043

0.7956989247

0.811827957

0.811827957

0.8268817204

0.8301075269

0.8376344086

0.8430107527

0.8537634409

0.8591397849

0.8612903226

0.8666666667

0.8795698925

0.8795698925

0.8795698925

0.8849462366

0.8924731183

0.8924731183

0.8978494624

0.8956989247

0.8903225806

0.8849462366

0.8741935484

0.8612903226

0.8430107527

0.8215053763

0.7956989247

0.764516129

0.735483871

0.7032258065

0.6698924731

0.6408602151

0.6064516129

0.5806451613

0.5483870968

0.5204301075

0.4967741935

0.464516129

0.4408602151

0.4150537634

0.3967741935

0.3698924731

0.3569892473

0.3387096774

0.3225806452

0.3096774194

0.2913978495

0.276344086

0.2655913978

0.2494623656

0.2365591398

0.2311827957

0.2150537634

0.2053763441

0.1967741935

0.1946236559

0.1838709677

0.1731182796

0.1709677419

0.1634408602

0.1602150538

0.147311828

0.1419354839

0.1344086022

0.1365591398

0.1258064516

0.1290322581

0.1258064516

0.1311827957

0.1397849462

0.1494623656

0.1602150538

0.1784946237

0.2

0.2258064516

0.247311828

0.276344086

0.3096774194

0.3387096774

0.3698924731

0.4021505376

0.4333333333

0.464516129

0.488172043

0.5150537634

0.5387096774

0.564516129

0.5903225806

0.6139784946

0.6301075269

0.6537634409

0.6698924731

0.6849462366

0.7010752688

0.7193548387

0.7322580645

0.7430107527

0.7559139785

0.7698924731

0.7795698925

0.7903225806

0.7978494624

0.811827957

0.8161290323

0.8268817204

0.8322580645

0.8397849462

0.8505376344

0.8537634409

0.8559139785

0.8634408602

0.8688172043

0.8774193548

0.8795698925

0.8849462366

0.8870967742

0.8903225806

0.8956989247

0.8924731183

0.9010752688

0.8924731183

0.8870967742

0.8774193548

0.8666666667

0.8376344086

0.8215053763

0.7956989247

0.7612903226

0.735483871

0.7010752688

0.6720430108

0.6376344086

0.6064516129

0.5806451613

0.5440860215

0.5150537634

0.488172043

0.464516129

0.4387096774

0.4204301075

0.3935483871

0.3731182796

0.3516129032

0.3419354839

0.3258064516

0.3096774194

0.2892473118

0.2731182796

0.2602150538

0.247311828

0.2419354839

0.223655914

0.2150537634

0.2075268817

0.1967741935

0.1870967742

0.1870967742

0.1784946237

0.1655913978

0.1580645161

0.152688172

0.147311828

0.1494623656

0.1365591398

0.1311827957

0.1290322581

0.1258064516

0.1258064516

0.1290322581

0.1344086022

0.1440860215

0.152688172

0.1709677419

0.1946236559

0.2129032258

0.2387096774

0.2709677419

0.3043010753

0.3311827957

0.3623655914

0.3892473118

0.4258064516

0.4516129032

0.4806451613

0.5064516129

0.5387096774

0.5591397849

0.5860215054

0.6064516129

0.6279569892

0.6462365591

0.6698924731

0.6903225806

0.6935483871

0.7193548387

0.7301075269

0.7408602151

0.7505376344

0.7612903226

0.7827956989

0.7849462366

0.7956989247

0.8064516129

0.8161290323

0.8247311828

0.8322580645

0.8430107527

0.8483870968

0.8537634409

0.8559139785

0.8591397849

0.8688172043

0.8720430108

0.8849462366

0.8795698925

0.8849462366

0.8903225806

0.8956989247

0.9032258065

0.8956989247

0.8956989247

0.8849462366

0.8795698925

0.8634408602

0.8483870968

0.8301075269

0.8010752688

0.7741935484

0.7430107527

0.711827957

0.6806451613

0.6516129032

0.6225806452

0.5860215054

0.5569892473

0.5247311828

0.4989247312

0.4698924731

0.4548387097

0.4279569892

0.4043010753

0.3806451613

0.3623655914

0.3419354839

0.3225806452

0.3150537634

0.2967741935

0.2784946237

0.2655913978

0.252688172

0.2440860215

0.2333333333

0.2204301075

0.2096774194

0.2

0.1967741935

0.1838709677

0.176344086

0.1677419355

0.1655913978

0.1548387097

0.152688172

0.1419354839

0.1397849462

0.1311827957

0.1258064516

0.1290322581

0.123655914

0.1290322581

0.1344086022

0.1419354839

0.1548387097

0.1677419355

0.1892473118

0.2075268817

0.2333333333

0.2602150538

0.2913978495

0.3225806452

0.3516129032

0.3838709677

0.4204301075

0.4440860215

0.4731182796

0.5064516129

0.5301075269

0.5537634409

0.5752688172

0.5989247312

0.6225806452

0.6430107527

0.6666666667

0.6774193548

0.6956989247

0.7086021505

0.7268817204

0.7301075269

0.7483870968

0.7612903226

0.7720430108

0.788172043

0.7956989247

0.8032258065

0.8139784946

0.8247311828

0.8268817204

0.8376344086

0.8397849462

0.8451612903

0.8591397849

0.8634408602

0.8666666667

0.8741935484

0.8817204301

0.8795698925

0.8849462366

0.8924731183

0.8956989247

0.8924731183

0.8924731183

0.8903225806

0.8870967742

0.8795698925

0.8666666667

0.8483870968

0.8322580645

0.8010752688

0.7741935484

0.7408602151

0.711827957

0.6774193548

0.6483870968

0.6172043011

0.5860215054

0.5537634409

0.5247311828

0.5043010753

0.4698924731

0.4462365591

0.4225806452

0.4043010753

0.3784946237

0.3602150538

0.3419354839

0.3258064516

0.3129032258

0.2913978495

0.2838709677

0.2655913978

0.252688172

0.2387096774

0.2333333333

0.2182795699

0.2150537634

0.2

0.1892473118

0.1784946237

0.1731182796

0.1655913978

0.1655913978

0.1548387097

0.1494623656

0.1440860215

0.1397849462

0.1344086022

0.1365591398

0.1311827957

0.1344086022

0.1397849462

0.1419354839

0.152688172

0.1709677419

0.1913978495

0.2096774194

0.2387096774

0.2623655914

0.2946236559

0.3279569892

0.3548387097

0.3860215054

0.4172043011

0.4462365591

0.4752688172

0.5010752688

0.5301075269

0.5591397849

0.5806451613

0.6043010753

0.6193548387

0.6483870968

0.6591397849

0.6774193548

0.6956989247

0.7139784946

0.7268817204

0.7376344086

0.7505376344

0.764516129

0.7741935484

0.7827956989

0.7924731183

0.8032258065

0.8161290323

0.8215053763

0.8301075269

0.835483871

0.8430107527

0.8505376344

0.8559139785

0.8612903226

0.8666666667

0.8688172043

0.8774193548

0.8817204301

0.8849462366

0.8924731183

0.8956989247

0.9010752688

0.8924731183

0.8956989247

0.8978494624

0.8849462366

0.8774193548

0.8612903226

0.8430107527

0.8161290323

0.7903225806

0.7666666667

0.735483871

0.7032258065

0.6720430108

0.6376344086

0.6096774194

0.5752688172

0.5516129032

0.5225806452

0.4913978495

0.464516129

0.4408602151

0.4225806452

0.3989247312

0.3731182796

0.3516129032

0.3365591398

0.3182795699

0.3075268817

0.2892473118

0.2731182796

0.2602150538

0.2494623656

0.2365591398

0.2311827957

0.2150537634

0.2075268817

0.1946236559

0.1870967742

0.1817204301

0.1731182796

0.1709677419

0.1655913978

0.1580645161

0.1440860215

0.1397849462

0.1344086022

0.1311827957

0.1311827957

0.1290322581

0.1311827957

0.1397849462

0.1494623656

0.1634408602

0.1817204301

0.2053763441

0.2290322581

0.2569892473

0.2860215054

0.3150537634

0.3462365591

0.3784946237

0.4075268817

0.4387096774

0.464516129

0.4989247312

0.5204301075

0.5462365591

0.5720430108

0.5935483871

0.6172043011

0.6408602151

0.6591397849

0.6720430108

0.6903225806

0.7032258065

0.7193548387

0.735483871

0.7483870968

0.7591397849

0.7720430108

0.7795698925

0.788172043

0.8010752688

0.811827957

0.8215053763

0.8247311828

0.8301075269

0.8376344086

0.8451612903

0.8537634409

0.8612903226

0.8666666667

0.8688172043

0.8720430108

0.8795698925

0.8817204301

0.8870967742

0.8924731183

0.9010752688

0.8978494624

0.8956989247

0.9010752688

0.8956989247

0.8870967742

0.8741935484

0.8612903226

0.8483870968

0.8161290323

0.7956989247

0.764516129

0.7322580645

0.6989247312

0.6698924731

0.635483871

0.6064516129

0.5752688172

0.5440860215

0.5172043011

0.488172043

0.4591397849

0.4387096774

0.4172043011

0.3913978495

0.3731182796

0.3548387097

0.3365591398

0.3204301075

0.3043010753

0.2892473118

0.2731182796

0.2602150538

0.2494623656

0.2387096774

0.2290322581

0.2150537634

0.2053763441

0.1967741935

0.1892473118

0.1817204301

0.176344086

0.1634408602

0.1602150538

0.152688172

0.1494623656

0.1419354839

0.1365591398

0.1344086022

0.1397849462

0.1397849462

0.147311828

0.1580645161

0.176344086

0.1913978495

0.2150537634

0.2387096774

0.2677419355

0.3021505376

0.3258064516

0.3602150538

0.3913978495

0.4172043011

0.4516129032

0.4827956989

0.5096774194

0.5333333333

0.5537634409

0.5827956989

0.6010752688

0.6279569892

0.6430107527

0.664516129

0.6774193548

0.6989247312

0.7064516129

0.7139784946

0.7225806452

0.7247311828

0.7225806452

0.7193548387

0.711827957

0.6989247312

0.6774193548

0.6537634409

0.6279569892

0.6010752688

0.5752688172

0.5537634409

0.5279569892

0.5010752688

0.4731182796

0.4462365591

0.4279569892

0.3989247312

0.3784946237

0.3602150538

0.3419354839

0.3258064516

0.3129032258

0.3096774194

0.3021505376

0.3021505376

0.3075268817

0.3182795699

0.3279569892

0.3462365591

0.3623655914

0.3860215054

0.4096774194

0.4387096774

0.4591397849

0.4860215054

0.511827957

0.5387096774

0.5623655914

0.588172043

0.6064516129

0.6279569892

0.6483870968

0.6752688172

0.6849462366

0.7010752688

0.711827957

0.7268817204

0.735483871

0.7430107527

0.7483870968

0.7537634409

0.7483870968

0.7408602151

0.735483871

0.7172043011

0.6989247312

0.6774193548

0.6516129032

0.6247311828

0.5956989247

0.5720430108

0.5462365591

0.5172043011

0.488172043

0.4623655914

0.4387096774

0.4172043011

0.3967741935

0.3752688172

0.3548387097

0.3387096774

0.3225806452

0.3150537634

0.3129032258

0.3096774194

0.3075268817

0.3182795699

0.3311827957

0.3419354839

0.3677419355

0.3860215054

0.4150537634

0.4333333333

0.4623655914

0.488172043

0.5150537634

0.5408602151

0.5623655914

0.5827956989

0.6064516129

0.6279569892

0.6483870968

0.6698924731

0.6849462366

0.7010752688

0.7172043011

0.7268817204

0.7430107527

0.7430107527

0.7537634409

0.7537634409

0.7505376344

0.7462365591

0.7322580645

0.711827957

0.6989247312

0.6806451613

0.6462365591

0.6193548387

0.5935483871

0.5698924731

0.5387096774

0.5096774194

0.488172043

0.4591397849

0.435483871

0.4172043011

0.3913978495

0.3655913978

0.3569892473

0.3387096774

0.3225806452

0.3182795699

0.3129032258

0.3129032258

0.3150537634

0.3225806452

0.3365591398

0.3602150538

0.3806451613

0.3967741935

0.4172043011

0.4440860215

0.4698924731

0.4913978495

0.5225806452

0.5440860215

0.5677419355

0.5903225806

0.611827957

0.6333333333

0.6591397849

0.6698924731

0.6903225806

0.7032258065

0.7172043011

0.7301075269

0.7430107527

0.7505376344

0.7505376344

0.7483870968

0.7430107527

0.735483871

0.7193548387

0.7010752688

0.6806451613

0.6612903226

0.6333333333

0.6043010753

0.5860215054

0.5516129032

0.5225806452

0.4935483871

0.4731182796

0.4462365591

0.4225806452

0.4021505376

0.3806451613

0.3602150538

0.3440860215

0.3333333333

0.3225806452

0.3204301075

0.3182795699

0.3182795699

0.3225806452

0.3365591398

0.3569892473

0.3731182796

0.3989247312

0.4150537634

0.4408602151

0.4677419355

0.4935483871

0.5172043011

0.5440860215

0.564516129

0.5903225806

0.611827957

0.6301075269

0.6569892473

0.6698924731

0.688172043

0.7032258065

0.7225806452

0.735483871

0.7430107527

0.7537634409

0.7591397849

0.7612903226

0.7591397849

0.7505376344

0.7462365591

0.7301075269

0.7139784946

0.688172043

0.6612903226

0.6333333333

0.6043010753

0.5806451613

0.5537634409

0.5225806452

0.5043010753

0.4731182796

0.4494623656

0.4279569892

0.4021505376

0.3806451613

0.3623655914

0.3494623656

0.3333333333

0.3258064516

0.3204301075

0.3204301075

0.3258064516

0.3333333333

0.3494623656

0.3623655914

0.3784946237

0.4043010753

0.4279569892

0.4569892473

0.4827956989

0.5010752688

0.5301075269

0.5537634409

0.5752688172

0.6010752688

0.6193548387

0.6430107527

0.664516129

0.6827956989

0.6956989247

0.7086021505

0.7225806452

0.7376344086

0.7462365591

0.7505376344

0.7537634409

0.7462365591

0.7462365591

0.7376344086

0.7268817204

0.7010752688

0.6806451613

0.6516129032

0.6279569892

0.5956989247

0.5752688172

0.5440860215

0.5172043011

0.4935483871

0.4623655914

0.4408602151

0.4204301075

0.3989247312

0.3752688172

0.3548387097

0.3419354839

0.3258064516

0.3225806452

0.3204301075

0.3096774194

0.3129032258

0.3204301075

0.3365591398

0.3462365591

0.3698924731

0.3935483871

0.4204301075

0.4387096774

0.4752688172

0.4913978495

0.5172043011

0.5408602151

0.5677419355

0.5860215054

0.6096774194

0.6301075269

0.6516129032

0.6698924731

0.6849462366

0.7010752688

0.7172043011

0.7301075269

0.7462365591

0.7462365591

0.7462365591

0.7505376344

0.7462365591

0.735483871

0.7193548387

0.7064516129

0.6827956989

0.6569892473

0.635483871

0.6096774194

0.5774193548

0.5483870968

0.5172043011

0.4967741935

0.4698924731

0.4440860215

0.4204301075

0.4021505376

0.3752688172

0.3602150538

0.3440860215

0.3333333333

0.3225806452

0.3150537634

0.3150537634

0.3204301075

0.3258064516

0.3365591398

0.3655913978

0.3784946237

0.4021505376

0.4225806452

0.4462365591

0.4698924731

0.4989247312

0.5225806452

0.5440860215

0.5720430108

0.5903225806

0.611827957

0.635483871

0.6569892473

0.6720430108

0.6903225806

0.7032258065

0.7247311828

0.7322580645

0.7408602151

0.7537634409

0.7591397849

0.7612903226

0.7559139785

0.7462365591

0.7376344086

0.7193548387

0.6989247312

0.6752688172

0.6483870968

0.6225806452

0.5935483871

0.564516129

0.5387096774

0.511827957

0.4860215054

0.4591397849

0.435483871

0.4204301075

0.3913978495

0.3731182796

0.3548387097

0.3419354839

0.3279569892

0.3204301075

0.3204301075

0.3182795699

0.3225806452

0.3365591398

0.3494623656

0.3698924731

0.3892473118

0.411827957

0.435483871

0.4591397849

0.4806451613

0.5096774194

0.535483871

0.5569892473

0.5860215054

0.6064516129

0.6301075269

0.6462365591

0.664516129

0.6806451613

0.6956989247

0.7086021505

0.7322580645

0.7376344086

0.7537634409

0.7591397849

0.7591397849

0.7537634409

0.7483870968

0.735483871

0.7268817204

0.7032258065

0.6827956989

0.6537634409

0.6279569892

0.6064516129

0.5720430108

0.5462365591

0.5204301075

0.488172043

0.4677419355

0.4440860215

0.4258064516

0.3967741935

0.3752688172

0.3655913978

0.3419354839

0.3279569892

0.3279569892

0.3204301075

0.3150537634

0.3150537634

0.3311827957

0.3387096774

0.3548387097

0.3752688172

0.3989247312

0.4225806452

0.4462365591

0.4731182796

0.4989247312

0.5225806452

0.5462365591

0.5698924731

0.5956989247

0.6139784946

0.635483871

0.6569892473

0.6752688172

0.6903225806

0.7086021505

0.7225806452

0.7408602151

0.7483870968

0.7505376344

0.7559139785

0.7537634409

0.7505376344

0.7408602151

0.7268817204

0.7086021505

0.6903225806

0.6666666667

0.6408602151

0.6139784946

0.588172043

0.5569892473

0.5387096774

0.5064516129

0.4752688172

0.4516129032

0.4333333333

0.4075268817

0.3838709677

0.3655913978

0.3494623656

0.3365591398

0.3279569892

0.3204301075

0.3150537634

0.3150537634

0.3258064516

0.3387096774

0.3494623656

0.3655913978

0.3860215054

0.4096774194

0.4333333333

0.4591397849

0.4860215054

0.5172043011

0.5333333333

0.5623655914

0.5806451613

0.6064516129

0.6279569892

0.6430107527

0.6698924731

0.6849462366

0.6989247312

0.7139784946

0.7301075269

0.7430107527

0.7537634409

0.7483870968

0.7537634409

0.7483870968

0.7462365591

0.7268817204

0.7086021505

0.6903225806

0.6698924731

0.6430107527

0.6139784946

0.5860215054

0.5591397849

0.5333333333

0.511827957

0.4784946237

0.4516129032

0.4279569892

0.4096774194

0.3838709677

0.3655913978

0.3548387097

0.3365591398

0.3225806452

0.3204301075

0.3150537634

0.3150537634

0.3204301075

0.3333333333

0.3494623656

0.3655913978

0.3860215054

0.4096774194

0.4387096774

0.4548387097

0.4860215054

0.5096774194

0.5301075269

0.5569892473

0.588172043

0.6043010753

0.6247311828

0.6483870968

0.6666666667

0.6849462366

0.7032258065

0.7139784946

0.7268817204

0.7408602151

0.7483870968

0.7559139785

0.7559139785

0.7591397849

0.7462365591

0.7376344086

0.7172043011

0.6989247312

0.6806451613

0.6516129032

0.6279569892

0.5989247312

0.5720430108

0.5516129032

0.5172043011

0.488172043

0.4623655914

0.4408602151

0.4225806452

0.3967741935

0.3752688172

0.3569892473

0.3419354839

0.3279569892

0.3204301075

0.3182795699

0.3204301075

0.3204301075

0.3365591398

0.3462365591

0.3677419355

0.3838709677

0.4096774194

0.4333333333

0.4548387097

0.4806451613

0.5096774194

0.535483871

0.5537634409

0.5774193548

0.6043010753

0.6279569892

0.6408602151

0.664516129

0.6774193548

0.6956989247

0.7139784946

0.7247311828

0.7430107527

0.7483870968

0.7505376344

0.7591397849

0.7559139785

0.7537634409

0.7462365591

0.7247311828

0.7086021505

0.688172043

0.6591397849

0.635483871

0.611827957

0.5827956989

0.5537634409

0.5247311828

0.4989247312

0.4784946237

0.4462365591

0.4225806452

0.4043010753

0.3892473118

0.3655913978

0.3462365591

0.3333333333

0.3225806452

0.3150537634

0.3150537634

0.3204301075

0.3279569892

0.3311827957

0.3494623656

0.3623655914

0.3838709677

0.4096774194

0.4333333333

0.4548387097

0.4860215054

0.5064516129

0.5301075269

0.5516129032

0.5774193548

0.6043010753

0.6247311828

0.6430107527

0.6612903226

0.6827956989

0.6956989247

0.711827957

0.7247311828

0.7376344086

0.7505376344

0.7537634409

0.7559139785

0.7505376344

0.7537634409

0.7376344086

0.7193548387

0.7064516129

0.6827956989

0.6537634409

0.6279569892

0.6064516129

0.5774193548

0.5483870968

0.5150537634

0.4913978495

0.4698924731

0.4408602151

0.4172043011

0.3989247312

0.3806451613

0.3623655914

0.3440860215

0.3333333333

0.3204301075

0.3182795699

0.3150537634

0.3182795699

0.3225806452

0.3387096774

0.3569892473

0.3806451613

0.3967741935

0.4204301075

0.4462365591

0.4698924731

0.4935483871

0.5204301075

0.5440860215

0.5720430108

0.5903225806

0.6139784946

0.6301075269

0.6537634409

0.6720430108

0.688172043

0.7064516129

0.7193548387

0.7322580645

0.7430107527

0.7505376344

0.7591397849

0.7591397849

0.7537634409

0.7462365591

0.7322580645

0.7172043011

0.6935483871

0.6698924731

0.6516129032

0.6193548387

0.5903225806

0.5698924731

0.535483871

0.5096774194

0.4806451613

0.4591397849

0.435483871

0.4096774194

0.3892473118

0.3698924731

0.3516129032

0.3419354839

0.3311827957

0.3225806452

0.3204301075

0.3225806452

0.3279569892

0.3387096774

0.3569892473

0.3752688172

0.3967741935

0.4172043011

0.4387096774

0.464516129

0.4913978495

0.5204301075

0.5387096774

0.564516129

0.5860215054

0.6096774194

0.635483871

0.6483870968

0.6698924731

0.688172043

0.7032258065

0.7172043011

0.735483871

0.7408602151

0.7505376344

0.7537634409

0.7559139785

0.7505376344

0.7462365591

0.7322580645

0.7139784946

0.6989247312

0.6666666667

0.6462365591

0.6172043011

0.5935483871

0.5623655914

0.5440860215

0.5096774194

0.4827956989

0.4569892473

0.435483871

0.4096774194

0.3892473118

0.3731182796

0.3569892473

0.3365591398

0.3258064516

0.3150537634

0.3129032258

0.3096774194

0.3182795699

0.3279569892

0.3440860215

0.3602150538

0.3892473118

0.4075268817

0.4279569892

0.4623655914

0.4806451613

0.511827957

0.5408602151

0.5537634409

0.5774193548

0.5989247312

0.6225806452

0.6483870968

0.6612903226

0.6806451613

0.6956989247

0.7086021505

0.7268817204

0.7408602151

0.7430107527

0.7483870968

0.7505376344

0.7537634409

0.7430107527

0.7301075269

0.711827957

0.6935483871

0.6666666667

0.6462365591

0.6193548387

0.5935483871

0.564516129

0.5387096774

0.5096774194

0.4827956989

0.4591397849

0.4333333333

0.411827957

0.3892473118

0.3698924731

0.3494623656

0.3365591398

0.3258064516

0.3225806452

0.3129032258

0.3129032258

0.3258064516

0.3258064516

0.3387096774

0.3569892473

0.3752688172

0.3967741935

0.4204301075

0.4408602151

0.4784946237

0.4989247312

0.5204301075

0.5537634409

0.5698924731

0.5903225806

0.6225806452

0.6333333333

0.6591397849

0.6806451613

0.6935483871

0.7064516129

0.7247311828

0.7322580645

0.7430107527

0.7483870968

0.7537634409

0.7559139785

0.7483870968

0.735483871

0.7247311828

0.6989247312

0.6806451613

0.6516129032

0.6279569892

0.6010752688

0.5720430108

0.5440860215

0.5204301075

0.4967741935

0.4677419355

0.4408602151

0.4204301075

0.3967741935

0.3752688172

0.3569892473

0.3419354839

0.3365591398

0.3182795699

0.3129032258

0.3129032258

0.3129032258

0.3258064516

0.3333333333

0.3494623656

0.3655913978

0.3860215054

0.411827957

0.435483871

0.4623655914

0.4913978495

0.5150537634

0.5462365591

0.5623655914

0.5903225806

0.6096774194

0.6279569892

0.6483870968

0.664516129

0.6849462366

0.6989247312

0.7139784946

0.7247311828

0.7408602151

0.7430107527

0.7462365591

0.7462365591

0.7430107527

0.7322580645

0.7225806452

0.7032258065

0.6806451613

0.6591397849

0.635483871

0.6010752688

0.5806451613

0.5516129032

0.5204301075

0.4967741935

0.4784946237

0.4494623656

0.4225806452

0.3967741935

0.3806451613

0.3602150538

0.3419354839

0.3311827957

0.3150537634

0.3129032258

0.3096774194

0.3096774194

0.3182795699

0.3204301075

0.3365591398

0.3548387097

0.3752688172

0.4021505376

0.4225806452

0.4462365591

0.4752688172

0.4989247312

0.5247311828

0.5516129032

0.5752688172

0.6010752688

0.6172043011

0.6430107527

0.664516129

0.6774193548

0.6903225806

0.7064516129

0.7225806452

0.7322580645

0.7376344086

0.7462365591

0.7505376344

0.7462365591

0.7430107527

0.7268817204

0.7086021505

0.6935483871

0.6698924731

0.6462365591

0.6139784946

0.5956989247

0.5623655914

0.5408602151

0.5064516129

0.4913978495

0.4569892473

0.4387096774

0.4204301075

0.3989247312

0.3698924731

0.3602150538

0.3494623656

0.3333333333

0.3279569892

0.3311827957

0.3333333333

0.3462365591

0.3602150538

0.3698924731

0.3913978495

0.411827957

0.4333333333

0.464516129

0.488172043

0.5150537634

0.5387096774

0.5591397849

0.5774193548

0.6010752688

0.6172043011

0.635483871

0.6483870968

0.6483870968

0.6569892473

0.6483870968

0.6376344086

0.6225806452

0.6096774194

0.5935483871

0.5698924731

0.5462365591

0.5225806452

0.5043010753

0.4784946237

0.4591397849

0.4408602151

0.4279569892

0.4150537634

0.4150537634

0.4096774194

0.4172043011

0.4225806452

0.435483871

0.4548387097

0.4752688172

0.4967741935

0.511827957

0.5333333333

0.5537634409

0.5827956989

0.5956989247

0.611827957

0.6301075269

0.6483870968

0.6569892473

0.664516129

0.6720430108

0.6666666667

0.664516129

0.6516129032

0.6376344086

0.6172043011

0.6010752688

0.5752688172

0.5483870968

0.5247311828

0.5043010753

0.4827956989

0.4623655914

0.4462365591

0.435483871

0.4279569892

0.4258064516

0.4258064516

0.4333333333

0.4440860215

0.464516129

0.4806451613

0.4989247312

0.5204301075

0.5440860215

0.5591397849

0.5806451613

0.6064516129

0.6193548387

0.635483871

0.6483870968

0.6612903226

0.664516129

0.6698924731

0.6666666667

0.6591397849

0.6462365591

0.6279569892

0.6139784946

0.5935483871

0.564516129

0.5440860215

0.5247311828

0.4935483871

0.4752688172

0.4591397849

0.4462365591

0.4333333333

0.4279569892

0.4311827957

0.4311827957

0.4408602151

0.4548387097

0.4677419355

0.4860215054

0.5010752688

0.5279569892

0.5483870968

0.5720430108

0.588172043

0.6064516129

0.6247311828

0.6483870968

0.6612903226

0.664516129

0.6752688172

0.6752688172

0.6752688172

0.664516129

0.6483870968

0.6333333333

0.6172043011

0.5903225806

0.5677419355

0.5440860215

0.5204301075

0.4967741935

0.4784946237

0.464516129

0.4462365591

0.4408602151

0.435483871

0.4408602151

0.4408602151

0.4494623656

0.4569892473

0.4731182796

0.4935483871

0.5096774194

0.5301075269

0.5483870968

0.5752688172

0.5903225806

0.6172043011

0.6333333333

0.6483870968

0.6569892473

0.6666666667

0.6752688172

0.6752688172

0.6752688172

0.6666666667

0.6537634409

0.6301075269

0.611827957

0.5956989247

0.5698924731

0.5483870968

0.5204301075

0.4967741935

0.4806451613

0.4591397849

0.4462365591

0.4387096774

0.4333333333

0.4279569892

0.4333333333

0.4494623656

0.4569892473

0.4698924731

0.4913978495

0.5096774194

0.5301075269

0.5516129032

0.5720430108

0.5935483871

0.611827957

0.6333333333

0.6462365591

0.6569892473

0.6666666667

0.6806451613

0.6752688172

0.6666666667

0.6666666667

0.6483870968

0.6408602151

0.6139784946

0.5935483871

0.5698924731

0.5483870968

0.5247311828

0.5043010753

0.4827956989

0.464516129

0.4516129032

0.4408602151

0.4333333333

0.435483871

0.4387096774

0.4494623656

0.4591397849

0.4806451613

0.4967741935

0.511827957

0.5333333333

0.5516129032

0.5720430108

0.5989247312

0.6139784946

0.6333333333

0.6462365591

0.6591397849

0.6752688172

0.6774193548

0.6774193548

0.6752688172

0.664516129

0.6516129032

0.6333333333

0.6172043011

0.5935483871

0.5698924731

0.5462365591

0.5247311828

0.5010752688

0.4806451613

0.464516129

0.4516129032

0.4408602151

0.4387096774

0.4440860215

0.4440860215

0.4516129032

0.4591397849

0.4806451613

0.4989247312

0.5172043011

0.5387096774

0.5569892473

0.5774193548

0.6043010753

0.6279569892

0.635483871

0.6537634409

0.664516129

0.6827956989

0.6849462366

0.6849462366

0.688172043

0.6752688172

0.6612903226

0.6462365591

0.6333333333

0.6064516129

0.5827956989

0.5591397849

0.5333333333

0.5096774194

0.488172043

0.4698924731

0.4623655914

0.4516129032

0.4516129032

0.4548387097

0.4569892473

0.4698924731

0.4806451613

0.5010752688

0.511827957

0.535483871

0.5537634409

0.5935483871

0.5956989247

0.6172043011

0.6376344086

0.6537634409

0.664516129

0.688172043

0.6935483871

0.6989247312

0.7010752688

0.6956989247

0.6849462366

0.6752688172

0.6591397849

0.635483871

0.611827957

0.5935483871

0.5623655914

0.5387096774

0.5247311828

0.4967741935

0.4806451613

0.464516129

0.4623655914

0.4516129032

0.4462365591

0.4569892473

0.4752688172

0.4784946237

0.4935483871

0.5096774194

0.5301075269

0.5537634409

0.5720430108

0.588172043

0.6096774194

0.6301075269

0.6516129032

0.6612903226

0.6720430108

0.6774193548

0.6849462366

0.688172043

0.6806451613

0.6720430108

0.6537634409

0.6279569892

0.6172043011

0.5956989247

0.5720430108

0.5462365591

0.5279569892

0.5064516129

0.4806451613

0.4731182796

0.4548387097

0.4494623656

0.4494623656

0.464516129

0.4623655914

0.4752688172

0.4967741935

0.5064516129

0.5301075269

0.5483870968

0.5698924731

0.588172043

0.6064516129

0.6301075269

0.6483870968

0.6591397849

0.6774193548

0.6849462366

0.6903225806

0.6956989247

0.6935483871

0.6849462366

0.6720430108

0.6591397849

0.635483871

0.6172043011

0.5956989247

0.5677419355

0.5516129032

0.5204301075

0.5010752688

0.4827956989

0.4677419355

0.4548387097

0.4462365591

0.4494623656

0.4494623656

0.4548387097

0.464516129

0.4806451613

0.4967741935

0.5172043011

0.535483871

0.5537634409

0.5774193548

0.5956989247

0.6172043011

0.6333333333

0.6516129032

0.6698924731

0.6774193548

0.6827956989

0.6903225806

0.6827956989

0.6827956989

0.6612903226

0.6537634409

0.635483871

0.6096774194

0.5860215054

0.5591397849

0.5333333333

0.5204301075

0.4913978495

0.4731182796

0.4591397849

0.4494623656

0.4440860215

0.4408602151

0.4462365591

0.4548387097

0.464516129

0.4784946237

0.5010752688

0.5172043011

0.5440860215

0.5591397849

0.5827956989

0.6010752688

0.6225806452

0.635483871

0.6483870968

0.6612903226

0.6774193548

0.6774193548

0.6806451613

0.6806451613

0.6698924731

0.6569892473

0.6430107527

0.6139784946

0.5956989247

0.5720430108

0.5483870968

0.5172043011

0.5010752688

0.4806451613

0.4623655914

0.4462365591

0.435483871

0.4333333333

0.4333333333

0.4387096774

0.4462365591

0.4591397849

0.4731182796

0.4935483871

0.5172043011

0.5301075269

0.5537634409

0.5720430108

0.5935483871

0.6172043011

0.6333333333

0.6483870968

0.6537634409

0.6612903226

0.6698924731

0.6806451613

0.6666666667

0.6569892473

0.6462365591

0.6301075269

0.611827957

0.5903225806

0.5591397849

0.535483871

0.5150537634

0.4967741935

0.4731182796

0.4548387097

0.4440860215

0.4387096774

0.4279569892

0.4333333333

0.4387096774

0.4494623656

0.464516129

0.4784946237

0.4967741935

0.5150537634

0.5333333333

0.5537634409

0.5774193548

0.5989247312

0.6172043011

0.6333333333

0.6483870968

0.6591397849

0.6698924731

0.6698924731

0.6720430108

0.664516129

0.6537634409

0.6376344086

0.6193548387

0.5935483871

0.5806451613

0.5537634409

0.5301075269

0.5064516129

0.4860215054

0.4731182796

0.4494623656

0.4387096774

0.435483871

0.4258064516

0.435483871

0.4387096774

0.4440860215

0.4591397849

0.4731182796

0.4935483871

0.5150537634

0.535483871

0.5537634409

0.5774193548

0.5956989247

0.6193548387

0.6279569892

0.6537634409

0.664516129

0.6698924731

0.6666666667

0.6698924731

0.6698924731

0.6537634409

0.6430107527

0.6193548387

0.5989247312

0.5752688172

0.5537634409

0.5279569892

0.5064516129

0.488172043

0.4677419355

0.4494623656

0.4440860215

0.4387096774

0.435483871

0.4387096774

0.4440860215

0.4548387097

0.4677419355

0.4913978495

0.5064516129

0.5247311828

0.5483870968

0.5720430108

0.5860215054

0.6064516129

0.6279569892

0.6430107527

0.6569892473

0.6698924731

0.6752688172

0.6806451613

0.6752688172

0.664516129

0.6569892473

0.6376344086

0.6193548387

0.5956989247

0.5752688172

0.5537634409

0.5279569892

0.5064516129

0.4827956989

0.4677419355

0.4548387097

0.4440860215

0.4333333333

0.4387096774

0.4408602151

0.4408602151

0.4516129032

0.4698924731

0.4860215054

0.5043010753

0.5247311828

0.5440860215

0.564516129

0.5860215054

0.6064516129

0.6193548387

0.6376344086

0.6569892473

0.664516129

0.6720430108

0.6720430108

0.6698924731

0.6591397849

0.6516129032

0.635483871

0.6172043011

0.5935483871

0.5698924731

0.5537634409

0.5225806452

0.4967741935

0.4784946237

0.4591397849

0.4494623656

0.4387096774

0.4387096774

0.4311827957

0.4387096774

0.4440860215

0.4591397849

0.4731182796

0.488172043

0.511827957

0.5333333333

0.5483870968

0.5720430108

0.5903225806

0.611827957

0.6301075269

0.6483870968

0.6591397849

0.6666666667

0.6752688172

0.6698924731

0.6774193548

0.6612903226

0.6516129032

0.635483871

0.611827957

0.5989247312

0.5677419355

0.5440860215

0.5247311828

0.5010752688

0.4784946237

0.4591397849

0.4494623656

0.4408602151

0.4387096774

0.435483871

0.4387096774

0.4494623656

0.4591397849

0.4784946237

0.4967741935

0.5150537634

0.5387096774

0.5537634409

0.5752688172

0.6010752688

0.6139784946

0.6301075269

0.6430107527

0.6591397849

0.664516129

0.6752688172

0.6698924731

0.6666666667

0.6591397849

0.6462365591

0.6279569892

0.611827957

0.588172043

0.564516129

0.5408602151

0.5204301075

0.4967741935

0.4752688172

0.4591397849

0.4494623656

0.4311827957

0.4279569892

0.4311827957

0.4387096774

0.4440860215

0.4591397849

0.4752688172

0.488172043

0.511827957

0.5301075269

0.5569892473

0.5720430108

0.5903225806

0.6139784946

0.6301075269

0.6430107527

0.6591397849

0.6698924731

0.6752688172

0.6720430108

0.6698924731

0.6612903226

0.6462365591

0.6301075269

0.611827957

0.5860215054

0.5677419355

0.5408602151

0.5225806452

0.5010752688

0.4784946237

0.4591397849

0.4462365591

0.435483871

0.4311827957

0.4408602151

0.4440860215

0.4494623656

0.4677419355

0.4784946237

0.4967741935

0.511827957

0.5333333333

0.5569892473

0.5752688172

0.5989247312

0.6172043011

0.6333333333

0.6483870968

0.6591397849

0.6698924731

0.6774193548

0.6720430108

0.6666666667

0.6612903226

0.6537634409

0.6376344086

0.6096774194

0.5903225806

0.564516129

0.5440860215

0.5172043011

0.4967741935

0.4698924731

0.464516129

0.4462365591

0.4333333333

0.4311827957

0.4311827957

0.4333333333

0.4440860215

0.4591397849

0.4731182796

0.4935483871

0.5172043011

0.5333333333

0.5537634409

0.5774193548

0.5935483871

0.6225806452

0.6279569892

0.6516129032

0.6591397849

0.6666666667

0.6666666667

0.6698924731

0.664516129

0.6537634409

0.6430107527

0.6193548387

0.6043010753

0.5827956989

0.5569892473

0.5387096774

0.5096774194

0.488172043

0.4677419355

0.4516129032

0.4387096774

0.435483871

0.4225806452

0.4204301075

0.4311827957

0.4440860215

0.4516129032

0.4677419355

0.4860215054

0.5064516129

0.5301075269

0.5462365591

0.5677419355

0.5860215054

0.6043010753

0.6279569892

0.6376344086

0.6462365591

0.6612903226

0.664516129

0.6569892473

0.6516129032

0.6462365591

0.6333333333

0.6172043011

0.5935483871

0.5752688172

0.5483870968

0.5279569892

0.5010752688

0.4806451613

0.464516129

0.4408602151

0.4279569892

0.4204301075

0.4204301075

0.4172043011

0.4225806452

0.4311827957

0.4408602151

0.4591397849

0.4784946237

0.4935483871

0.5204301075

0.5408602151

0.5623655914

0.5774193548

0.5989247312

0.6172043011

0.6376344086

0.6408602151

0.6516129032

0.6569892473

0.6537634409

0.6569892473

0.6462365591

0.6279569892

0.6139784946

0.588172043

0.564516129

0.5462365591

0.5279569892

0.5010752688

0.4752688172

0.4569892473

0.4408602151

0.4311827957

0.4204301075

0.4150537634

0.4204301075

0.4204301075

0.4311827957

0.4462365591

0.4623655914

0.4806451613

0.4989247312

0.5225806452

0.5408602151

0.5623655914

0.588172043

0.6064516129

0.6172043011

0.6301075269

0.6430107527

0.6537634409

0.6516129032

0.6516129032

0.6462365591

0.6333333333

0.6225806452

0.6010752688

0.5860215054

0.5569892473

0.535483871

0.5096774194

0.488172043

0.4698924731

0.4462365591

0.4408602151

0.4258064516

0.4150537634

0.4075268817

0.4172043011

0.4172043011

0.4279569892

0.4440860215

0.4591397849

0.4784946237

0.5010752688

0.5225806452

0.5440860215

0.564516129

0.5827956989

0.6010752688

0.6172043011

0.635483871

0.6462365591

0.6516129032

0.6516129032

0.6537634409

0.6516129032

0.635483871

0.6225806452

0.6010752688

0.5752688172

0.5516129032

0.5279569892

0.511827957

0.488172043

0.4677419355

0.4462365591

0.435483871

0.4258064516

0.4204301075

0.4204301075

0.4258064516

0.435483871

0.4440860215

0.4623655914

0.4784946237

0.4967741935

0.5247311828

0.5387096774

0.5569892473

0.5806451613

0.5989247312

0.6139784946

0.635483871

0.6430107527

0.6569892473

0.6612903226

0.664516129

0.6612903226

0.6516129032

0.6376344086

0.6193548387

0.5989247312

0.5774193548

0.5537634409

0.5333333333

0.511827957

0.488172043

0.4677419355

0.4516129032

0.4408602151

0.435483871

0.4258064516

0.4311827957

0.4311827957

0.4408602151

0.4548387097

0.4677419355

0.4967741935

0.5064516129

0.5247311828

0.5440860215

0.5677419355

0.588172043

0.611827957

0.6225806452

0.6462365591

0.6516129032

0.6612903226

0.664516129

0.6612903226

0.6591397849

0.6462365591

0.6333333333

0.6172043011

0.5956989247

0.5698924731

0.5591397849

0.5279569892

0.5010752688

0.488172043

0.4623655914

0.4494623656

0.4333333333

0.4258064516

0.4225806452

0.4225806452

0.4333333333

0.4440860215

0.4569892473

0.464516129

0.4860215054

0.5043010753

0.5333333333

0.5462365591

0.5677419355

0.5903225806

0.6064516129

0.6247311828

0.6408602151

0.6483870968

0.664516129

0.664516129

0.6698924731

0.6537634409

0.6516129032

0.6301075269

0.611827957

0.5935483871

0.5752688172

0.5516129032

0.5247311828

0.5064516129

0.4860215054

0.4623655914

0.4494623656

0.4333333333

0.4279569892

0.4258064516

0.4279569892

0.435483871

0.4548387097

0.464516129

0.4752688172

0.4967741935

0.5150537634

0.5333333333

0.5591397849

0.5774193548

0.5935483871

0.6139784946

0.635483871

0.6462365591

0.6591397849

0.664516129

0.6752688172

0.6752688172

0.6612903226

0.6537634409

0.6376344086

0.6172043011

0.5935483871

0.5720430108

0.5516129032

0.535483871

0.5096774194

0.4806451613

0.464516129

0.4494623656

0.4387096774

0.435483871

0.4333333333

0.4311827957

0.4408602151

0.4494623656

0.4698924731

0.4860215054

0.5010752688

0.5172043011

0.5387096774

0.5623655914

0.5806451613

0.5989247312

0.6193548387

0.6408602151

0.6516129032

0.6591397849

0.6698924731

0.6698924731

0.6752688172

0.6666666667

0.6516129032

0.635483871

0.6225806452

0.5935483871

0.5774193548

0.5537634409

0.5247311828

0.5043010753

0.4827956989

0.464516129

0.4548387097

0.4462365591

0.4408602151

0.4279569892

0.4408602151

0.4387096774

0.4494623656

0.4677419355

0.488172043

0.5043010753

0.5225806452

0.5462365591

0.564516129

0.588172043

0.6043010753

0.6225806452

0.6430107527

0.6537634409

0.6612903226

0.6698924731

0.6698924731

0.6666666667

0.664516129

0.6462365591

0.6279569892

0.611827957

0.5860215054

0.5623655914

0.5440860215

0.5204301075

0.4967741935

0.4752688172



FO134

		0		0.00933		0.0032258065

		0.01		0.00928		-0.0021505376

		0.02		0.0093		0

		0.03		0.00933		0.0032258065

		0.04		0.0093		0

		0.05		0.0093		0

		0.06		0.00928		-0.0021505376

		0.07		0.00933		0.0032258065

		0.08		0.0093		0

		0.095		0.0093		0

		0.103		0.00928		-0.0021505376

		0.112		0.0093		0

		0.122		0.00933		0.0032258065

		0.131		0.0093		0

		0.142		0.00938		0.0086021505

		0.152		0.0093		0

		0.162		0.00928		-0.0021505376

		0.172		0.00933		0.0032258065

		0.182		0.00933		0.0032258065

		0.192		0.00938		0.0086021505

		0.202		0.0093		0

		0.212		0.0093		0

		0.222		0.00933		0.0032258065

		0.232		0.00933		0.0032258065

		0.242		0.00933		0.0032258065

		0.252		0.0093		0

		0.262		0.0093		0

		0.272		0.0093		0

		0.282		0.0093		0

		0.292		0.00928		-0.0021505376

		0.302		0.0093		0

		0.312		0.00935		0.0053763441

		0.322		0.0093		0

		0.332		0.0093		0

		0.342		0.00933		0.0032258065

		0.352		0.00935		0.0053763441

		0.362		0.0093		0

		0.372		0.00925		-0.0053763441

		0.382		0.00928		-0.0021505376

		0.392		0.00935		0.0053763441

		0.402		0.0093		0

		0.412		0.0093		0

		0.422		0.0093		0

		0.432		0.00938		0.0086021505

		0.442		0.00935		0.0053763441

		0.452		0.0093		0

		0.462		0.00938		0.0086021505

		0.472		0.0093		0

		0.482		0.00933		0.0032258065

		0.492		0.0093		0

		0.502		0.00935		0.0053763441

		0.512		0.00928		-0.0021505376

		0.522		0.0093		0

		0.532		0.0093		0

		0.542		0.00933		0.0032258065

		0.552		0.00933		0.0032258065

		0.562		0.00925		-0.0053763441

		0.572		0.00938		0.0086021505

		0.582		0.00933		0.0032258065

		0.592		0.0093		0

		0.602		0.00933		0.0032258065

		0.612		0.0093		0

		0.622		0.00938		0.0086021505

		0.632		0.0093		0

		0.642		0.0093		0

		0.652		0.0093		0

		0.662		0.0093		0

		0.672		0.0093		0

		0.682		0.0093		0

		0.692		0.00938		0.0086021505

		0.702		0.00933		0.0032258065

		0.712		0.00928		-0.0021505376

		0.722		0.0093		0

		0.732		0.0093		0

		0.742		0.00933		0.0032258065

		0.752		0.0093		0

		0.762		0.0093		0

		0.772		0.00933		0.0032258065

		0.782		0.0093		0

		0.792		0.0093		0

		0.802		0.00933		0.0032258065

		0.812		0.00933		0.0032258065

		0.822		0.0093		0

		0.832		0.00928		-0.0021505376

		0.842		0.0093		0

		0.854		0.0093		0

		0.864		0.00928		-0.0021505376

		0.874		0.00933		0.0032258065

		0.884		0.00933		0.0032258065

		0.895		0.0093		0

		0.904		0.00938		0.0086021505

		0.914		0.0093		0

		0.925		0.0093		0

		0.934		0.00933		0.0032258065

		0.944		0.0093		0

		0.954		0.0093		0

		0.964		0.0093		0

		0.974		0.0093		0

		0.984		0.00933		0.0032258065

		0.994		0.00933		0.0032258065

		1.004		0.00928		-0.0021505376

		1.014		0.00933		0.0032258065

		1.024		0.0093		0

		1.035		0.0093		0

		1.044		0.0093		0

		1.054		0.00935		0.0053763441

		1.064		0.0093		0

		1.074		0.00935		0.0053763441

		1.084		0.0093		0

		1.094		0.0093		0

		1.106		0.0093		0

		1.114		0.00935		0.0053763441

		1.124		0.00928		-0.0021505376

		1.134		0.0093		0

		1.145		0.00928		-0.0021505376

		1.154		0.0093		0

		1.164		0.00938		0.0086021505

		1.174		0.0093		0

		1.184		0.00933		0.0032258065

		1.194		0.00933		0.0032258065

		1.204		0.00935		0.0053763441

		1.214		0.0093		0

		1.224		0.0093		0

		1.234		0.0093		0

		1.244		0.0093		0

		1.255		0.0093		0

		1.264		0.00933		0.0032258065

		1.274		0.00935		0.0053763441

		1.284		0.0093		0

		1.294		0.0093		0

		1.304		0.00928		-0.0021505376

		1.314		0.0093		0

		1.324		0.0093		0

		1.334		0.00933		0.0032258065

		1.344		0.00933		0.0032258065

		1.354		0.00933		0.0032258065

		1.365		0.0093		0

		1.374		0.00928		-0.0021505376

		1.384		0.00935		0.0053763441

		1.394		0.0093		0

		1.404		0.0093		0

		1.414		0.00935		0.0053763441

		1.424		0.0093		0

		1.436		0.00928		-0.0021505376

		1.446		0.00933		0.0032258065

		1.456		0.00933		0.0032258065

		1.466		0.00935		0.0053763441

		1.478		0.00933		0.0032258065

		1.487		0.00933		0.0032258065

		1.497		0.00935		0.0053763441

		1.507		0.00935		0.0053763441

		1.517		0.00933		0.0032258065

		1.527		0.0093		0

		1.537		0.00933		0.0032258065

		1.547		0.0093		0

		1.557		0.00928		-0.0021505376

		1.567		0.00933		0.0032258065

		1.577		0.00933		0.0032258065

		1.587		0.0093		0

		1.597		0.00935		0.0053763441

		1.607		0.0093		0

		1.618		0.00935		0.0053763441

		1.627		0.00933		0.0032258065

		1.637		0.0093		0

		1.647		0.0093		0

		1.657		0.00928		-0.0021505376

		1.667		0.0093		0

		1.677		0.0093		0

		1.687		0.00933		0.0032258065

		1.697		0.00935		0.0053763441

		1.707		0.00933		0.0032258065

		1.717		0.00933		0.0032258065

		1.727		0.0093		0

		1.739		0.0093		0

		1.749		0.0093		0

		1.758		0.00933		0.0032258065

		1.768		0.00938		0.0086021505

		1.778		0.00935		0.0053763441

		1.79		0.00938		0.0086021505

		1.799		0.0093		0

		1.809		0.0093		0

		1.819		0.00933		0.0032258065

		1.829		0.00928		-0.0021505376

		1.839		0.00935		0.0053763441

		1.849		0.0093		0

		1.859		0.00933		0.0032258065

		1.869		0.00933		0.0032258065

		1.88		0.0093		0

		1.889		0.00933		0.0032258065

		1.899		0.00933		0.0032258065

		1.91		0.00933		0.0032258065

		1.919		0.00933		0.0032258065

		1.929		0.0093		0

		1.939		0.0093		0

		1.949		0.0093		0

		1.959		0.00933		0.0032258065

		1.971		0.0093		0

		1.98		0.00933		0.0032258065

		1.99		0.00933		0.0032258065

		2		0.0093		0

		2.009		0.00933		0.0032258065

		2.019		0.0093		0

		2.03		0.0093		0

		2.04		0.0093		0

		2.05		0.00933		0.0032258065

		2.06		0.00938		0.0086021505

		2.07		0.0093		0

		2.08		0.0093		0

		2.09		0.00933		0.0032258065

		2.1		0.00933		0.0032258065

		2.11		0.0093		0

		2.12		0.00935		0.0053763441

		2.13		0.00935		0.0053763441

		2.14		0.00933		0.0032258065

		2.15		0.0093		0

		2.16		0.0093		0

		2.17		0.0093		0

		2.18		0.0093		0

		2.19		0.00935		0.0053763441

		2.2		0.0093		0

		2.21		0.00935		0.0053763441

		2.22		0.0093		0

		2.23		0.0093		0

		2.24		0.00933		0.0032258065

		2.25		0.0093		0

		2.259		0.00933		0.0032258065

		2.27		0.00933		0.0032258065

		2.28		0.0093		0

		2.29		0.00935		0.0053763441

		2.3		0.0093		0

		2.31		0.00933		0.0032258065

		2.32		0.0093		0

		2.33		0.00928		-0.0021505376

		2.34		0.00935		0.0053763441

		2.35		0.00935		0.0053763441

		2.36		0.0093		0

		2.373		0.00935		0.0053763441

		2.383		0.00933		0.0032258065

		2.393		0.00935		0.0053763441

		2.403		0.0093		0

		2.413		0.00933		0.0032258065

		2.423		0.0093		0

		2.433		0.0093		0

		2.443		0.0093		0

		2.453		0.00928		-0.0021505376

		2.463		0.00935		0.0053763441

		2.473		0.00933		0.0032258065

		2.483		0.00933		0.0032258065

		2.493		0.00933		0.0032258065

		2.503		0.00938		0.0086021505

		2.513		0.0093		0

		2.523		0.0093		0

		2.533		0.00933		0.0032258065

		2.543		0.00933		0.0032258065

		2.553		0.0093		0

		2.563		0.00928		-0.0021505376

		2.573		0.00933		0.0032258065

		2.583		0.0093		0

		2.593		0.0093		0

		2.603		0.00938		0.0086021505

		2.613		0.0093		0

		2.627		0.00933		0.0032258065

		2.633		0.0093		0

		2.643		0.0093		0

		2.653		0.0093		0

		2.663		0.0093		0

		2.673		0.00933		0.0032258065

		2.683		0.00933		0.0032258065

		2.693		0.0093		0

		2.703		0.0093		0

		2.713		0.00933		0.0032258065

		2.723		0.00933		0.0032258065

		2.733		0.00925		-0.0053763441

		2.743		0.0093		0

		2.753		0.00938		0.0086021505

		2.763		0.00933		0.0032258065

		2.773		0.00928		-0.0021505376

		2.783		0.00933		0.0032258065

		2.793		0.00933		0.0032258065

		2.803		0.00928		-0.0021505376

		2.813		0.0093		0

		2.823		0.0093		0

		2.833		0.00928		-0.0021505376

		2.843		0.00938		0.0086021505

		2.853		0.00928		-0.0021505376

		2.863		0.0093		0

		2.873		0.00928		-0.0021505376

		2.883		0.00933		0.0032258065

		2.893		0.00933		0.0032258065

		2.903		0.0093		0

		2.913		0.00938		0.0086021505

		2.923		0.0093		0

		2.933		0.00933		0.0032258065

		2.943		0.00933		0.0032258065

		2.953		0.00938		0.0086021505

		2.963		0.0093		0

		2.973		0.00933		0.0032258065

		2.983		0.00938		0.0086021505

		2.993		0.00928		-0.0021505376

		3.003		0.00933		0.0032258065

		3.013		0.00933		0.0032258065

		3.023		0.0093		0

		3.033		0.00933		0.0032258065

		3.043		0.0093		0

		3.053		0.0093		0

		3.063		0.0093		0

		3.073		0.0093		0

		3.083		0.00935		0.0053763441

		3.094		0.0093		0

		3.103		0.00928		-0.0021505376

		3.113		0.00933		0.0032258065

		3.123		0.0093		0

		3.133		0.00933		0.0032258065

		3.143		0.00933		0.0032258065

		3.153		0.00933		0.0032258065

		3.163		0.0093		0

		3.173		0.00935		0.0053763441

		3.183		0.0093		0

		3.193		0.0093		0

		3.203		0.00928		-0.0021505376

		3.213		0.00935		0.0053763441

		3.223		0.00933		0.0032258065

		3.233		0.00938		0.0086021505

		3.243		0.00928		-0.0021505376

		3.253		0.00933		0.0032258065

		3.263		0.0093		0

		3.273		0.00933		0.0032258065

		3.283		0.0093		0

		3.293		0.00933		0.0032258065

		3.303		0.00928		-0.0021505376

		3.313		0.00933		0.0032258065

		3.323		0.0093		0

		3.333		0.00933		0.0032258065

		3.343		0.00933		0.0032258065

		3.353		0.0093		0

		3.363		0.00933		0.0032258065

		3.373		0.00933		0.0032258065

		3.386		0.00935		0.0053763441

		3.393		0.0093		0

		3.403		0.00933		0.0032258065

		3.413		0.00933		0.0032258065

		3.423		0.0093		0

		3.433		0.00928		-0.0021505376

		3.443		0.0093		0

		3.453		0.00938		0.0086021505

		3.463		0.0093		0

		3.473		0.00938		0.0086021505

		3.483		0.0093		0

		3.493		0.00933		0.0032258065

		3.503		0.00928		-0.0021505376

		3.513		0.0093		0

		3.523		0.00933		0.0032258065

		3.533		0.0093		0

		3.543		0.0093		0

		3.553		0.0093		0

		3.563		0.0093		0

		3.573		0.0093		0

		3.583		0.0093		0

		3.593		0.00933		0.0032258065

		3.603		0.00938		0.0086021505

		3.613		0.0093		0

		3.623		0.00933		0.0032258065

		3.633		0.00935		0.0053763441

		3.643		0.0093		0

		3.653		0.00933		0.0032258065

		3.663		0.0093		0

		3.673		0.0093		0

		3.683		0.0093		0

		3.693		0.0093		0

		3.703		0.0093		0

		3.713		0.00935		0.0053763441

		3.723		0.00935		0.0053763441

		3.733		0.0093		0

		3.743		0.0093		0

		3.753		0.0093		0

		3.763		0.00928		-0.0021505376

		3.773		0.0093		0

		3.783		0.00928		-0.0021505376

		3.793		0.00933		0.0032258065

		3.803		0.0093		0

		3.813		0.00928		-0.0021505376

		3.823		0.0093		0

		3.833		0.0093		0

		3.843		0.00928		-0.0021505376

		3.853		0.0093		0

		3.863		0.0093		0

		3.873		0.0093		0

		3.883		0.00935		0.0053763441

		3.893		0.00933		0.0032258065

		3.903		0.0093		0

		3.913		0.0093		0

		3.923		0.00933		0.0032258065

		3.933		0.0093		0

		3.943		0.00928		-0.0021505376

		3.953		0.00928		-0.0021505376

		3.963		0.00935		0.0053763441

		3.973		0.0093		0

		3.983		0.00925		-0.0053763441

		3.993		0.00935		0.0053763441

		4.003		0.00933		0.0032258065

		4.013		0.0093		0

		4.023		0.00925		-0.0053763441

		4.032		0.00933		0.0032258065

		4.043		0.00933		0.0032258065

		4.052		0.00933		0.0032258065

		4.063		0.00935		0.0053763441

		4.073		0.0093		0

		4.083		0.0093		0

		4.093		0.00935		0.0053763441

		4.103		0.00938		0.0086021505

		4.113		0.0093		0

		4.123		0.0093		0

		4.133		0.00933		0.0032258065

		4.144		0.0093		0

		4.154		0.00933		0.0032258065

		4.164		0.00928		-0.0021505376

		4.174		0.00933		0.0032258065

		4.184		0.00928		-0.0021505376

		4.194		0.0093		0

		4.203		0.0093		0

		4.214		0.00933		0.0032258065

		4.223		0.00933		0.0032258065

		4.234		0.0093		0

		4.244		0.0093		0

		4.254		0.00933		0.0032258065

		4.264		0.00933		0.0032258065

		4.274		0.00933		0.0032258065

		4.284		0.0093		0

		4.294		0.00933		0.0032258065

		4.304		0.00933		0.0032258065

		4.314		0.00928		-0.0021505376

		4.324		0.0093		0

		4.334		0.00933		0.0032258065

		4.344		0.0093		0

		4.354		0.00928		-0.0021505376

		4.364		0.0093		0

		4.374		0.0093		0

		4.384		0.00935		0.0053763441

		4.397		0.00933		0.0032258065

		4.404		0.0093		0

		4.414		0.00933		0.0032258065

		4.424		0.0093		0

		4.434		0.00935		0.0053763441

		4.444		0.0093		0

		4.453		0.0093		0

		4.464		0.0093		0

		4.473		0.0093		0

		4.484		0.00933		0.0032258065

		4.494		0.00933		0.0032258065

		4.504		0.00935		0.0053763441

		4.514		0.00928		-0.0021505376

		4.524		0.0093		0

		4.534		0.0093		0

		4.544		0.00933		0.0032258065

		4.554		0.0093		0

		4.564		0.00935		0.0053763441

		4.574		0.00933		0.0032258065

		4.584		0.00928		-0.0021505376

		4.594		0.0093		0

		4.604		0.00928		-0.0021505376

		4.614		0.0093		0

		4.624		0.00933		0.0032258065

		4.634		0.00935		0.0053763441

		4.644		0.00928		-0.0021505376

		4.654		0.00933		0.0032258065

		4.664		0.0093		0

		4.674		0.0093		0

		4.684		0.00933		0.0032258065

		4.694		0.00933		0.0032258065

		4.703		0.0093		0

		4.714		0.00935		0.0053763441

		4.724		0.0093		0

		4.734		0.0093		0

		4.744		0.0093		0

		4.754		0.0093		0

		4.764		0.00933		0.0032258065

		4.774		0.00928		-0.0021505376

		4.784		0.0093		0

		4.794		0.00933		0.0032258065

		4.804		0.0093		0

		4.814		0.00933		0.0032258065

		4.825		0.0093		0

		4.834		0.0093		0

		4.844		0.00933		0.0032258065

		4.854		0.00933		0.0032258065

		4.864		0.00933		0.0032258065

		4.874		0.00928		-0.0021505376

		4.884		0.00933		0.0032258065

		4.894		0.0093		0

		4.904		0.00933		0.0032258065

		4.914		0.00928		-0.0021505376

		4.924		0.00935		0.0053763441

		4.935		0.00938		0.0086021505

		4.944		0.0093		0

		4.954		0.00928		-0.0021505376

		4.964		0.0093		0

		4.974		0.00933		0.0032258065

		4.984		0.0093		0

		4.994		0.0093		0

		5.004		0.00935		0.0053763441

		5.014		0.00928		-0.0021505376

		5.024		0.00933		0.0032258065

		5.034		0.00933		0.0032258065

		5.045		0.00933		0.0032258065

		5.054		0.00933		0.0032258065

		5.064		0.0093		0

		5.074		0.00935		0.0053763441

		5.084		0.0093		0

		5.094		0.00928		-0.0021505376

		5.104		0.00933		0.0032258065

		5.114		0.00933		0.0032258065

		5.124		0.00933		0.0032258065

		5.134		0.00933		0.0032258065

		5.144		0.0093		0

		5.156		0.0093		0

		5.164		0.0093		0

		5.174		0.00935		0.0053763441

		5.184		0.00933		0.0032258065

		5.194		0.00935		0.0053763441

		5.204		0.0093		0

		5.214		0.00938		0.0086021505

		5.224		0.00933		0.0032258065

		5.234		0.0093		0

		5.244		0.00935		0.0053763441

		5.254		0.0093		0

		5.264		0.00933		0.0032258065

		5.274		0.00933		0.0032258065

		5.284		0.00928		-0.0021505376

		5.294		0.0093		0

		5.304		0.00933		0.0032258065

		5.314		0.0093		0

		5.324		0.00935		0.0053763441

		5.334		0.00925		-0.0053763441

		5.344		0.00933		0.0032258065

		5.354		0.0093		0

		5.364		0.00933		0.0032258065

		5.374		0.00928		-0.0021505376

		5.384		0.00933		0.0032258065

		5.394		0.00933		0.0032258065

		5.407		0.00928		-0.0021505376

		5.414		0.0093		0

		5.424		0.0093		0

		5.434		0.0093		0

		5.444		0.00933		0.0032258065

		5.454		0.0093		0

		5.464		0.00933		0.0032258065

		5.474		0.00928		-0.0021505376

		5.484		0.00933		0.0032258065

		5.494		0.00933		0.0032258065

		5.504		0.0093		0

		5.514		0.00938		0.0086021505

		5.524		0.0093		0

		5.534		0.00935		0.0053763441

		5.544		0.0093		0

		5.554		0.0093		0

		5.564		0.0093		0

		5.574		0.0093		0

		5.584		0.0093		0

		5.595		0.0093		0

		5.604		0.00938		0.0086021505

		5.614		0.0093		0

		5.624		0.00933		0.0032258065

		5.634		0.0093		0

		5.644		0.0093		0

		5.654		0.00933		0.0032258065

		5.664		0.00933		0.0032258065

		5.674		0.00933		0.0032258065

		5.684		0.0093		0

		5.694		0.00928		-0.0021505376

		5.705		0.0093		0

		5.714		0.0093		0

		5.724		0.0093		0

		5.734		0.00928		-0.0021505376

		5.744		0.00933		0.0032258065

		5.754		0.00933		0.0032258065

		5.764		0.0093		0

		5.774		0.0093		0

		5.784		0.0093		0

		5.794		0.00933		0.0032258065

		5.804		0.00935		0.0053763441

		5.815		0.00933		0.0032258065

		5.824		0.00933		0.0032258065

		5.834		0.0093		0

		5.844		0.00933		0.0032258065

		5.854		0.00935		0.0053763441

		5.864		0.00933		0.0032258065

		5.874		0.00933		0.0032258065

		5.884		0.00933		0.0032258065

		5.894		0.0093		0

		5.904		0.0093		0

		5.914		0.00933		0.0032258065

		5.925		0.00935		0.0053763441

		5.934		0.00935		0.0053763441

		5.944		0.00933		0.0032258065

		5.954		0.0093		0

		5.964		0.0093		0

		5.974		0.0093		0

		5.984		0.00933		0.0032258065

		5.994		0.00933		0.0032258065

		6.004		0.00928		-0.0021505376

		6.014		0.0093		0

		6.024		0.0093		0

		6.035		0.00928		-0.0021505376

		6.044		0.00933		0.0032258065

		6.054		0.00938		0.0086021505

		6.064		0.00933		0.0032258065

		6.074		0.00928		-0.0021505376

		6.084		0.00933		0.0032258065

		6.094		0.00928		-0.0021505376

		6.104		0.00933		0.0032258065

		6.114		0.0093		0

		6.124		0.0093		0

		6.134		0.00933		0.0032258065

		6.145		0.00938		0.0086021505

		6.154		0.00933		0.0032258065

		6.165		0.00928		-0.0021505376

		6.174		0.0094		0.0107526882

		6.184		0.0093		0

		6.194		0.00928		-0.0021505376

		6.204		0.00933		0.0032258065

		6.214		0.00933		0.0032258065

		6.224		0.0093		0

		6.234		0.00933		0.0032258065

		6.244		0.00935		0.0053763441

		6.255		0.00935		0.0053763441

		6.264		0.0093		0

		6.274		0.00933		0.0032258065

		6.284		0.00935		0.0053763441

		6.294		0.0093		0

		6.304		0.00933		0.0032258065

		6.314		0.00933		0.0032258065

		6.324		0.00928		-0.0021505376

		6.334		0.0093		0

		6.344		0.0093		0

		6.354		0.00928		-0.0021505376

		6.365		0.00928		-0.0021505376

		6.374		0.00928		-0.0021505376

		6.384		0.00928		-0.0021505376

		6.394		0.0093		0

		6.404		0.0093		0

		6.417		0.0093		0

		6.424		0.0093		0

		6.434		0.0093		0

		6.444		0.00938		0.0086021505

		6.454		0.00933		0.0032258065

		6.464		0.00935		0.0053763441

		6.474		0.0093		0

		6.484		0.0093		0

		6.494		0.0093		0

		6.504		0.0093		0

		6.514		0.0093		0

		6.524		0.00928		-0.0021505376

		6.534		0.00933		0.0032258065

		6.544		0.0093		0

		6.554		0.0093		0

		6.564		0.0093		0

		6.574		0.0093		0

		6.584		0.0093		0

		6.594		0.00925		-0.0053763441

		6.604		0.00933		0.0032258065

		6.614		0.0093		0

		6.624		0.00933		0.0032258065

		6.634		0.00935		0.0053763441

		6.644		0.0093		0

		6.654		0.0093		0

		6.664		0.00928		-0.0021505376

		6.674		0.0093		0

		6.684		0.00933		0.0032258065

		6.694		0.00928		-0.0021505376

		6.704		0.00935		0.0053763441

		6.714		0.00935		0.0053763441

		6.724		0.0093		0

		6.734		0.00935		0.0053763441

		6.744		0.00933		0.0032258065

		6.754		0.0093		0

		6.764		0.00933		0.0032258065

		6.774		0.00928		-0.0021505376

		6.784		0.0093		0

		6.794		0.00925		-0.0053763441

		6.804		0.0093		0

		6.814		0.00938		0.0086021505

		6.824		0.00928		-0.0021505376

		6.834		0.0093		0

		6.844		0.0093		0

		6.854		0.0093		0

		6.864		0.0093		0

		6.874		0.0093		0

		6.884		0.0093		0

		6.894		0.00933		0.0032258065

		6.904		0.00933		0.0032258065

		6.914		0.00933		0.0032258065

		6.924		0.0093		0

		6.934		0.00933		0.0032258065

		6.944		0.00933		0.0032258065

		6.954		0.0093		0

		6.964		0.0093		0

		6.974		0.0093		0

		6.984		0.0093		0

		6.994		0.00933		0.0032258065

		7.004		0.00933		0.0032258065

		7.014		0.0093		0

		7.024		0.0093		0

		7.034		0.00933		0.0032258065

		7.044		0.0093		0

		7.054		0.0093		0

		7.064		0.0093		0

		7.074		0.00935		0.0053763441

		7.084		0.00938		0.0086021505

		7.094		0.0093		0

		7.104		0.00933		0.0032258065

		7.114		0.00933		0.0032258065

		7.124		0.00933		0.0032258065

		7.134		0.00938		0.0086021505

		7.144		0.00938		0.0086021505

		7.154		0.00938		0.0086021505

		7.164		0.00942		0.0129032258

		7.176		0.0095		0.0215053763

		7.185		0.00957		0.0290322581

		7.195		0.00972		0.0451612903

		7.205		0.00994		0.0688172043

		7.215		0.01011		0.0870967742

		7.225		0.01033		0.1107526882

		7.235		0.0106		0.1397849462

		7.245		0.01089		0.1709677419

		7.255		0.01118		0.2021505376

		7.265		0.01145		0.2311827957

		7.275		0.01177		0.2655913978

		7.285		0.01208		0.2989247312

		7.295		0.01238		0.3311827957

		7.305		0.0127		0.3655913978

		7.315		0.01304		0.4021505376

		7.325		0.01328		0.4279569892

		7.335		0.01355		0.4569892473

		7.345		0.01377		0.4806451613

		7.355		0.01409		0.5150537634

		7.365		0.01431		0.5387096774

		7.375		0.01448		0.5569892473

		7.385		0.01472		0.5827956989

		7.395		0.01489		0.6010752688

		7.405		0.01504		0.6172043011

		7.415		0.01521		0.635483871

		7.428		0.01543		0.6591397849

		7.435		0.01555		0.6720430108

		7.445		0.01572		0.6903225806

		7.455		0.0158		0.6989247312

		7.465		0.01594		0.7139784946

		7.475		0.01606		0.7268817204

		7.485		0.01614		0.735483871

		7.495		0.01626		0.7483870968

		7.505		0.01636		0.7591397849

		7.515		0.01646		0.7698924731

		7.525		0.01653		0.7774193548

		7.535		0.01663		0.788172043

		7.545		0.0167		0.7956989247

		7.555		0.01682		0.8086021505

		7.565		0.01689		0.8161290323

		7.575		0.01692		0.8193548387

		7.585		0.01697		0.8247311828

		7.595		0.01704		0.8322580645

		7.605		0.01709		0.8376344086

		7.615		0.01721		0.8505376344

		7.625		0.01716		0.8451612903

		7.635		0.01729		0.8591397849

		7.645		0.01729		0.8591397849

		7.655		0.01736		0.8666666667

		7.665		0.01736		0.8666666667

		7.675		0.01741		0.8720430108

		7.685		0.01743		0.8741935484

		7.695		0.01748		0.8795698925

		7.705		0.0175		0.8817204301

		7.715		0.01758		0.8903225806

		7.725		0.01755		0.8870967742

		7.735		0.0176		0.8924731183

		7.745		0.01765		0.8978494624

		7.755		0.01777		0.9107526882

		7.765		0.01772		0.9053763441

		7.775		0.01775		0.9086021505

		7.785		0.01775		0.9086021505

		7.795		0.01777		0.9107526882

		7.805		0.01787		0.9215053763

		7.815		0.01782		0.9161290323

		7.825		0.01782		0.9161290323

		7.835		0.01787		0.9215053763

		7.845		0.01792		0.9268817204

		7.855		0.01792		0.9268817204

		7.865		0.01792		0.9268817204

		7.875		0.01799		0.9344086022

		7.885		0.01797		0.9322580645

		7.895		0.01794		0.9290322581

		7.905		0.01807		0.9430107527

		7.915		0.01804		0.9397849462

		7.925		0.01802		0.9376344086

		7.935		0.01804		0.9397849462

		7.945		0.01804		0.9397849462

		7.955		0.01807		0.9430107527

		7.965		0.01809		0.9451612903

		7.975		0.01809		0.9451612903

		7.985		0.01812		0.9483870968

		7.995		0.01812		0.9483870968

		8.005		0.01816		0.952688172

		8.015		0.01812		0.9483870968

		8.025		0.01816		0.952688172

		8.035		0.01821		0.9580645161

		8.045		0.01814		0.9505376344

		8.055		0.01821		0.9580645161

		8.065		0.01819		0.9559139785

		8.075		0.01819		0.9559139785

		8.085		0.01821		0.9580645161

		8.095		0.01821		0.9580645161

		8.105		0.01826		0.9634408602

		8.115		0.01824		0.9612903226

		8.125		0.01826		0.9634408602

		8.135		0.01826		0.9634408602

		8.145		0.01824		0.9612903226

		8.155		0.01826		0.9634408602

		8.164		0.01831		0.9688172043

		8.175		0.01829		0.9666666667

		8.185		0.01824		0.9612903226

		8.195		0.01831		0.9688172043

		8.204		0.01829		0.9666666667

		8.215		0.01831		0.9688172043

		8.225		0.01836		0.9741935484

		8.235		0.01836		0.9741935484

		8.245		0.01831		0.9688172043

		8.255		0.01831		0.9688172043

		8.265		0.01833		0.9709677419

		8.275		0.01833		0.9709677419

		8.285		0.01838		0.976344086

		8.295		0.01836		0.9741935484

		8.305		0.01836		0.9741935484

		8.315		0.01833		0.9709677419

		8.325		0.01836		0.9741935484

		8.335		0.01838		0.976344086

		8.345		0.01836		0.9741935484

		8.355		0.01841		0.9795698925

		8.365		0.01838		0.976344086

		8.375		0.01843		0.9817204301

		8.385		0.01841		0.9795698925

		8.395		0.01838		0.976344086

		8.405		0.01838		0.976344086

		8.414		0.01838		0.976344086

		8.425		0.01841		0.9795698925

		8.437		0.01838		0.976344086

		8.447		0.01836		0.9741935484

		8.457		0.01843		0.9817204301

		8.467		0.01841		0.9795698925

		8.477		0.01841		0.9795698925

		8.487		0.01841		0.9795698925

		8.497		0.01846		0.9849462366

		8.506		0.01841		0.9795698925

		8.517		0.01846		0.9849462366

		8.527		0.01841		0.9795698925

		8.537		0.01843		0.9817204301

		8.546		0.01843		0.9817204301

		8.557		0.01846		0.9849462366

		8.567		0.01846		0.9849462366

		8.577		0.01846		0.9849462366

		8.587		0.01841		0.9795698925

		8.597		0.01846		0.9849462366

		8.607		0.01841		0.9795698925

		8.617		0.01843		0.9817204301

		8.627		0.01843		0.9817204301

		8.637		0.01846		0.9849462366

		8.647		0.01843		0.9817204301

		8.657		0.01846		0.9849462366

		8.667		0.01851		0.9903225806

		8.677		0.01846		0.9849462366

		8.69		0.01846		0.9849462366

		8.697		0.01848		0.9870967742

		8.707		0.01846		0.9849462366

		8.717		0.01846		0.9849462366

		8.727		0.01843		0.9817204301

		8.737		0.01848		0.9870967742

		8.747		0.01846		0.9849462366

		8.756		0.01843		0.9817204301

		8.767		0.01846		0.9849462366

		8.777		0.01846		0.9849462366

		8.787		0.01848		0.9870967742

		8.797		0.01846		0.9849462366

		8.807		0.01848		0.9870967742

		8.817		0.01846		0.9849462366

		8.827		0.01851		0.9903225806

		8.837		0.01851		0.9903225806

		8.847		0.01855		0.9946236559

		8.857		0.01853		0.9924731183

		8.867		0.01853		0.9924731183

		8.877		0.01848		0.9870967742

		8.887		0.01848		0.9870967742

		8.897		0.01848		0.9870967742

		8.907		0.01848		0.9870967742

		8.917		0.01851		0.9903225806

		8.927		0.01853		0.9924731183

		8.937		0.01851		0.9903225806

		8.947		0.01851		0.9903225806

		8.957		0.01851		0.9903225806

		8.967		0.01848		0.9870967742

		8.977		0.01848		0.9870967742

		8.987		0.01848		0.9870967742

		8.997		0.01855		0.9946236559

		9.006		0.01851		0.9903225806

		9.017		0.01851		0.9903225806

		9.027		0.01851		0.9903225806

		9.037		0.01851		0.9903225806

		9.047		0.01855		0.9946236559

		9.057		0.01853		0.9924731183

		9.067		0.01851		0.9903225806

		9.077		0.01851		0.9903225806

		9.087		0.01853		0.9924731183

		9.097		0.01855		0.9946236559

		9.107		0.01851		0.9903225806

		9.117		0.01855		0.9946236559

		9.127		0.01855		0.9946236559

		9.137		0.01853		0.9924731183

		9.147		0.01851		0.9903225806

		9.157		0.01851		0.9903225806

		9.167		0.01855		0.9946236559

		9.177		0.01853		0.9924731183

		9.187		0.01853		0.9924731183

		9.197		0.01851		0.9903225806

		9.207		0.01851		0.9903225806

		9.217		0.01853		0.9924731183

		9.227		0.01855		0.9946236559

		9.237		0.0186		1

		9.247		0.01853		0.9924731183

		9.256		0.01853		0.9924731183

		9.267		0.01851		0.9903225806

		9.277		0.01853		0.9924731183

		9.287		0.01853		0.9924731183

		9.297		0.01851		0.9903225806

		9.307		0.01858		0.9978494624

		9.317		0.01853		0.9924731183

		9.327		0.01853		0.9924731183

		9.337		0.01858		0.9978494624

		9.347		0.01855		0.9946236559

		9.357		0.01853		0.9924731183

		9.367		0.01853		0.9924731183

		9.377		0.01851		0.9903225806

		9.387		0.01858		0.9978494624

		9.397		0.01851		0.9903225806

		9.407		0.01853		0.9924731183

		9.417		0.01855		0.9946236559

		9.427		0.01855		0.9946236559

		9.437		0.01853		0.9924731183

		9.447		0.01855		0.9946236559

		9.457		0.01855		0.9946236559

		9.467		0.01855		0.9946236559

		9.477		0.01853		0.9924731183

		9.487		0.01858		0.9978494624

		9.497		0.01853		0.9924731183

		9.507		0.01858		0.9978494624

		9.517		0.01855		0.9946236559

		9.527		0.01853		0.9924731183

		9.537		0.01853		0.9924731183

		9.547		0.01851		0.9903225806

		9.557		0.01858		0.9978494624

		9.567		0.01855		0.9946236559

		9.577		0.01855		0.9946236559

		9.587		0.01851		0.9903225806

		9.597		0.01853		0.9924731183

		9.607		0.01855		0.9946236559

		9.617		0.01855		0.9946236559

		9.627		0.01855		0.9946236559

		9.637		0.01855		0.9946236559

		9.647		0.01853		0.9924731183

		9.657		0.01855		0.9946236559

		9.667		0.01858		0.9978494624

		9.677		0.01855		0.9946236559

		9.687		0.01853		0.9924731183

		9.7		0.0186		1

		9.707		0.01855		0.9946236559

		9.717		0.01855		0.9946236559

		9.727		0.01855		0.9946236559

		9.737		0.01853		0.9924731183

		9.747		0.01855		0.9946236559

		9.757		0.01858		0.9978494624

		9.767		0.01858		0.9978494624

		9.777		0.01855		0.9946236559

		9.787		0.01858		0.9978494624

		9.797		0.01855		0.9946236559

		9.807		0.01853		0.9924731183

		9.817		0.01855		0.9946236559

		9.827		0.01858		0.9978494624

		9.837		0.01855		0.9946236559

		9.847		0.01858		0.9978494624

		9.857		0.01858		0.9978494624

		9.867		0.01855		0.9946236559

		9.877		0.01858		0.9978494624

		9.887		0.01855		0.9946236559

		9.897		0.01858		0.9978494624

		9.907		0.01858		0.9978494624

		9.917		0.01855		0.9946236559

		9.927		0.01855		0.9946236559

		9.937		0.01853		0.9924731183

		9.947		0.01855		0.9946236559

		9.957		0.01855		0.9946236559

		9.967		0.01855		0.9946236559

		9.977		0.01858		0.9978494624

		9.987		0.01855		0.9946236559

		9.997		0.01858		0.9978494624

		10.007		0.01858		0.9978494624

		10.017		0.01858		0.9978494624

		10.027		0.01858		0.9978494624

		10.037		0.01855		0.9946236559

		10.047		0.01855		0.9946236559

		10.057		0.01855		0.9946236559

		10.067		0.01855		0.9946236559

		10.077		0.01855		0.9946236559

		10.087		0.01858		0.9978494624

		10.097		0.01858		0.9978494624

		10.107		0.0186		1

		10.117		0.01858		0.9978494624

		10.127		0.01855		0.9946236559

		10.137		0.01855		0.9946236559

		10.147		0.0186		1

		10.157		0.0186		1

		10.167		0.01855		0.9946236559

		10.177		0.0186		1

		10.187		0.01858		0.9978494624

		10.197		0.01855		0.9946236559

		10.207		0.0186		1

		10.217		0.01853		0.9924731183

		10.227		0.0186		1

		10.237		0.01855		0.9946236559

		10.247		0.01858		0.9978494624

		10.257		0.0186		1

		10.267		0.01855		0.9946236559

		10.277		0.01853		0.9924731183

		10.287		0.01858		0.9978494624

		10.297		0.01853		0.9924731183

		10.307		0.01858		0.9978494624

		10.317		0.0186		1

		10.327		0.01853		0.9924731183

		10.337		0.01855		0.9946236559

		10.347		0.01858		0.9978494624

		10.357		0.01858		0.9978494624

		10.367		0.01858		0.9978494624

		10.377		0.01858		0.9978494624

		10.387		0.01855		0.9946236559

		10.397		0.01863		1.0032258065

		10.407		0.01855		0.9946236559

		10.417		0.0186		1

		10.427		0.0186		1

		10.437		0.01855		0.9946236559

		10.447		0.01855		0.9946236559

		10.459		0.01863		1.0032258065

		10.468		0.0186		1

		10.478		0.01855		0.9946236559

		10.488		0.01858		0.9978494624

		10.498		0.01858		0.9978494624

		10.508		0.01855		0.9946236559

		10.518		0.01858		0.9978494624

		10.528		0.01858		0.9978494624

		10.538		0.01855		0.9946236559

		10.548		0.0186		1

		10.558		0.0186		1

		10.568		0.01858		0.9978494624

		10.578		0.01863		1.0032258065

		10.588		0.01863		1.0032258065

		10.598		0.0186		1

		10.608		0.0186		1

		10.618		0.01858		0.9978494624

		10.628		0.01855		0.9946236559

		10.638		0.0186		1

		10.649		0.01855		0.9946236559

		10.658		0.01855		0.9946236559

		10.668		0.01858		0.9978494624

		10.678		0.01858		0.9978494624

		10.688		0.01858		0.9978494624

		10.698		0.01855		0.9946236559

		10.711		0.01855		0.9946236559

		10.718		0.01858		0.9978494624

		10.728		0.01855		0.9946236559

		10.738		0.01863		1.0032258065

		10.748		0.01858		0.9978494624

		10.759		0.01858		0.9978494624

		10.768		0.01855		0.9946236559

		10.778		0.01858		0.9978494624

		10.788		0.01863		1.0032258065

		10.798		0.0186		1

		10.808		0.01858		0.9978494624

		10.818		0.01858		0.9978494624

		10.828		0.01858		0.9978494624

		10.838		0.0186		1

		10.848		0.0186		1

		10.858		0.01855		0.9946236559

		10.869		0.01855		0.9946236559

		10.878		0.01858		0.9978494624

		10.888		0.01855		0.9946236559

		10.898		0.01858		0.9978494624

		10.908		0.01855		0.9946236559

		10.918		0.01858		0.9978494624

		10.928		0.01863		1.0032258065

		10.938		0.0186		1

		10.948		0.0186		1

		10.958		0.01858		0.9978494624

		10.968		0.01858		0.9978494624

		10.979		0.01858		0.9978494624

		10.988		0.01858		0.9978494624

		10.998		0.01855		0.9946236559

		11.008		0.0186		1

		11.018		0.0186		1

		11.028		0.01855		0.9946236559

		11.038		0.01865		1.0053763441

		11.048		0.01855		0.9946236559

		11.058		0.01858		0.9978494624

		11.068		0.01858		0.9978494624

		11.078		0.01855		0.9946236559

		11.089		0.01858		0.9978494624

		11.098		0.01858		0.9978494624

		11.108		0.01858		0.9978494624

		11.118		0.01858		0.9978494624

		11.128		0.01858		0.9978494624

		11.138		0.01855		0.9946236559

		11.148		0.01858		0.9978494624

		11.158		0.01863		1.0032258065

		11.168		0.01858		0.9978494624

		11.178		0.01858		0.9978494624

		11.188		0.01863		1.0032258065

		11.199		0.01858		0.9978494624

		11.208		0.01858		0.9978494624

		11.218		0.01855		0.9946236559

		11.228		0.01858		0.9978494624

		11.238		0.01858		0.9978494624

		11.248		0.01858		0.9978494624

		11.258		0.01863		1.0032258065

		11.268		0.0186		1

		11.278		0.01863		1.0032258065

		11.288		0.01858		0.9978494624

		11.298		0.0186		1

		11.309		0.01855		0.9946236559

		11.318		0.01858		0.9978494624

		11.328		0.01858		0.9978494624

		11.338		0.01863		1.0032258065

		11.348		0.01855		0.9946236559

		11.358		0.0186		1

		11.368		0.01865		1.0053763441

		11.378		0.01863		1.0032258065

		11.388		0.01858		0.9978494624

		11.398		0.0186		1

		11.408		0.0186		1

		11.418		0.01858		0.9978494624

		11.428		0.0186		1

		11.438		0.01858		0.9978494624

		11.448		0.01858		0.9978494624

		11.458		0.0186		1

		11.469		0.0186		1

		11.478		0.01858		0.9978494624

		11.488		0.01858		0.9978494624

		11.498		0.0186		1

		11.508		0.01865		1.0053763441

		11.518		0.01858		0.9978494624

		11.528		0.01858		0.9978494624

		11.538		0.01858		0.9978494624

		11.548		0.01858		0.9978494624

		11.558		0.01863		1.0032258065

		11.568		0.01855		0.9946236559

		11.578		0.01863		1.0032258065

		11.588		0.01858		0.9978494624

		11.598		0.01858		0.9978494624

		11.608		0.01858		0.9978494624

		11.618		0.01858		0.9978494624

		11.628		0.0186		1

		11.638		0.0186		1

		11.648		0.0186		1

		11.658		0.01858		0.9978494624

		11.668		0.01858		0.9978494624

		11.678		0.01858		0.9978494624

		11.688		0.01858		0.9978494624

		11.698		0.01855		0.9946236559

		11.708		0.01863		1.0032258065

		11.721		0.01858		0.9978494624

		11.728		0.01858		0.9978494624

		11.738		0.01858		0.9978494624

		11.748		0.0186		1

		11.758		0.01858		0.9978494624

		11.768		0.0186		1

		11.778		0.01858		0.9978494624

		11.788		0.01855		0.9946236559

		11.798		0.0186		1

		11.808		0.0186		1

		11.818		0.01858		0.9978494624

		11.828		0.0186		1

		11.838		0.0186		1

		11.848		0.0186		1

		11.858		0.0186		1

		11.868		0.0186		1

		11.878		0.0186		1

		11.888		0.01863		1.0032258065

		11.898		0.01865		1.0053763441

		11.908		0.01858		0.9978494624

		11.918		0.0186		1

		11.928		0.01855		0.9946236559

		11.938		0.0186		1

		11.948		0.0186		1

		11.958		0.01858		0.9978494624

		11.968		0.01858		0.9978494624

		11.978		0.0186		1

		11.988		0.0186		1

		11.998		0.0186		1

		12.008		0.0186		1

		12.018		0.01863		1.0032258065

		12.028		0.0186		1

		12.038		0.01858		0.9978494624

		12.048		0.01858		0.9978494624

		12.058		0.01858		0.9978494624

		12.068		0.01863		1.0032258065

		12.078		0.01863		1.0032258065

		12.088		0.01858		0.9978494624

		12.098		0.01858		0.9978494624

		12.108		0.01858		0.9978494624

		12.118		0.0186		1

		12.128		0.01858		0.9978494624

		12.138		0.01858		0.9978494624

		12.148		0.01858		0.9978494624

		12.158		0.01858		0.9978494624

		12.168		0.01858		0.9978494624

		12.178		0.01863		1.0032258065

		12.188		0.01858		0.9978494624

		12.198		0.0186		1

		12.208		0.01865		1.0053763441

		12.218		0.0186		1

		12.228		0.0186		1

		12.238		0.01858		0.9978494624

		12.248		0.01858		0.9978494624

		12.258		0.01858		0.9978494624

		12.268		0.0186		1

		12.278		0.0186		1

		12.288		0.01858		0.9978494624

		12.298		0.0186		1

		12.308		0.01858		0.9978494624

		12.318		0.01858		0.9978494624

		12.328		0.01858		0.9978494624

		12.338		0.0186		1

		12.348		0.01863		1.0032258065

		12.358		0.01858		0.9978494624

		12.368		0.01865		1.0053763441

		12.378		0.01858		0.9978494624

		12.388		0.01858		0.9978494624

		12.398		0.0186		1

		12.408		0.01858		0.9978494624

		12.418		0.0186		1

		12.428		0.01863		1.0032258065

		12.438		0.01865		1.0053763441

		12.448		0.01865		1.0053763441

		12.458		0.0186		1

		12.468		0.01863		1.0032258065

		12.478		0.0186		1

		12.488		0.01858		0.9978494624

		12.498		0.0186		1

		12.508		0.01858		0.9978494624

		12.518		0.01863		1.0032258065

		12.528		0.01858		0.9978494624

		12.538		0.0186		1

		12.548		0.01855		0.9946236559

		12.558		0.0186		1

		12.568		0.0186		1

		12.578		0.01863		1.0032258065

		12.588		0.0186		1

		12.598		0.01858		0.9978494624

		12.608		0.01863		1.0032258065

		12.618		0.01858		0.9978494624

		12.628		0.01858		0.9978494624

		12.638		0.01858		0.9978494624

		12.648		0.01858		0.9978494624

		12.658		0.0186		1

		12.668		0.01865		1.0053763441

		12.678		0.01863		1.0032258065

		12.688		0.0186		1

		12.698		0.01863		1.0032258065

		12.708		0.01858		0.9978494624

		12.718		0.01858		0.9978494624

		12.731		0.0186		1

		12.74		0.0186		1

		12.75		0.01858		0.9978494624

		12.76		0.0186		1

		12.77		0.0186		1

		12.78		0.0186		1

		12.79		0.0186		1

		12.8		0.0186		1

		12.81		0.01863		1.0032258065

		12.82		0.0186		1

		12.83		0.01865		1.0053763441

		12.84		0.01865		1.0053763441

		12.85		0.0186		1

		12.86		0.01863		1.0032258065

		12.87		0.01865		1.0053763441

		12.88		0.01863		1.0032258065

		12.89		0.0186		1

		12.9		0.0186		1

		12.91		0.01858		0.9978494624

		12.92		0.0186		1

		12.93		0.01863		1.0032258065

		12.94		0.0186		1

		12.95		0.01863		1.0032258065

		12.96		0.01865		1.0053763441

		12.97		0.01865		1.0053763441

		12.983		0.01863		1.0032258065

		12.99		0.01858		0.9978494624

		13		0.0186		1

		13.01		0.01863		1.0032258065

		13.02		0.01865		1.0053763441

		13.03		0.0186		1

		13.04		0.0186		1

		13.05		0.0186		1

		13.06		0.01865		1.0053763441

		13.07		0.0186		1

		13.08		0.0186		1

		13.09		0.01865		1.0053763441

		13.1		0.01863		1.0032258065

		13.11		0.01865		1.0053763441

		13.122		0.0186		1

		13.131		0.0186		1

		13.141		0.01858		0.9978494624

		13.151		0.01865		1.0053763441

		13.161		0.0186		1

		13.171		0.0186		1

		13.181		0.01865		1.0053763441

		13.191		0.0186		1

		13.201		0.0186		1

		13.211		0.0186		1

		13.221		0.01863		1.0032258065

		13.231		0.01863		1.0032258065

		13.241		0.0186		1

		13.251		0.01858		0.9978494624

		13.261		0.0186		1

		13.271		0.01858		0.9978494624

		13.281		0.0186		1

		13.291		0.01858		0.9978494624

		13.301		0.0186		1

		13.311		0.01855		0.9946236559

		13.321		0.01855		0.9946236559

		13.331		0.01848		0.9870967742

		13.341		0.01843		0.9817204301

		13.351		0.01831		0.9688172043

		13.361		0.01819		0.9559139785

		13.371		0.01799		0.9344086022

		13.381		0.01782		0.9161290323

		13.391		0.0175		0.8817204301

		13.401		0.01724		0.8537634409

		13.411		0.01699		0.8268817204

		13.421		0.0166		0.7849462366

		13.431		0.01636		0.7591397849

		13.441		0.01599		0.7193548387

		13.451		0.01572		0.6903225806

		13.461		0.01536		0.6516129032

		13.471		0.01509		0.6225806452

		13.481		0.01477		0.588172043

		13.491		0.01448		0.5569892473

		13.501		0.01423		0.5301075269

		13.511		0.01396		0.5010752688

		13.521		0.01372		0.4752688172

		13.531		0.01353		0.4548387097

		13.541		0.01338		0.4387096774

		13.551		0.01316		0.4150537634

		13.561		0.01289		0.3860215054

		13.571		0.01274		0.3698924731

		13.581		0.01257		0.3516129032

		13.591		0.01243		0.3365591398

		13.601		0.0123		0.3225806452

		13.611		0.01218		0.3096774194

		13.621		0.01206		0.2967741935

		13.631		0.01191		0.2806451613

		13.641		0.01179		0.2677419355

		13.651		0.01169		0.2569892473

		13.661		0.01162		0.2494623656

		13.671		0.01152		0.2387096774

		13.681		0.01138		0.223655914

		13.691		0.01135		0.2204301075

		13.701		0.01123		0.2075268817

		13.711		0.01113		0.1967741935

		13.721		0.01111		0.1946236559

		13.731		0.01104		0.1870967742

		13.741		0.01096		0.1784946237

		13.751		0.01091		0.1731182796

		13.761		0.01084		0.1655913978

		13.771		0.01079		0.1602150538

		13.781		0.01077		0.1580645161

		13.791		0.01069		0.1494623656

		13.801		0.01064		0.1440860215

		13.811		0.0106		0.1397849462

		13.821		0.01057		0.1365591398

		13.831		0.01052		0.1311827957

		13.841		0.01047		0.1258064516

		13.851		0.01045		0.123655914

		13.861		0.01045		0.123655914

		13.871		0.01038		0.1161290323

		13.881		0.01033		0.1107526882

		13.891		0.01033		0.1107526882

		13.901		0.01025		0.1021505376

		13.911		0.01025		0.1021505376

		13.921		0.01025		0.1021505376

		13.931		0.01021		0.0978494624

		13.941		0.01013		0.0892473118

		13.951		0.01013		0.0892473118

		13.961		0.01011		0.0870967742

		13.971		0.01008		0.0838709677

		13.981		0.01008		0.0838709677

		13.991		0.01006		0.0817204301

		14.001		0.01003		0.0784946237

		14.011		0.01006		0.0817204301

		14.021		0.01001		0.076344086

		14.031		0.00999		0.0741935484

		14.041		0.00994		0.0688172043

		14.051		0.00996		0.0709677419

		14.061		0.00996		0.0709677419

		14.071		0.00996		0.0709677419

		14.081		0.00991		0.0655913978

		14.091		0.00994		0.0688172043

		14.101		0.00989		0.0634408602

		14.111		0.00986		0.0602150538

		14.121		0.00981		0.0548387097

		14.131		0.00981		0.0548387097

		14.141		0.00986		0.0602150538

		14.151		0.00981		0.0548387097

		14.161		0.00981		0.0548387097

		14.171		0.00979		0.052688172

		14.181		0.00977		0.0505376344

		14.191		0.00974		0.047311828

		14.201		0.00981		0.0548387097

		14.211		0.00977		0.0505376344

		14.221		0.00974		0.047311828

		14.231		0.00969		0.0419354839

		14.241		0.00977		0.0505376344

		14.251		0.00972		0.0451612903

		14.261		0.00969		0.0419354839

		14.271		0.00969		0.0419354839

		14.281		0.00967		0.0397849462

		14.291		0.00969		0.0419354839

		14.301		0.00969		0.0419354839

		14.311		0.00967		0.0397849462

		14.321		0.00972		0.0451612903

		14.331		0.00967		0.0397849462

		14.341		0.00964		0.0365591398

		14.351		0.00964		0.0365591398

		14.361		0.00964		0.0365591398

		14.371		0.00959		0.0311827957

		14.381		0.00964		0.0365591398

		14.391		0.00962		0.0344086022

		14.401		0.00967		0.0397849462

		14.411		0.00964		0.0365591398

		14.421		0.00957		0.0290322581

		14.431		0.00962		0.0344086022

		14.441		0.00959		0.0311827957

		14.451		0.00962		0.0344086022

		14.461		0.00962		0.0344086022

		14.471		0.00957		0.0290322581

		14.481		0.00952		0.023655914

		14.491		0.00955		0.0268817204

		14.501		0.00955		0.0268817204

		14.511		0.00957		0.0290322581

		14.521		0.00962		0.0344086022

		14.531		0.00955		0.0268817204

		14.541		0.00955		0.0268817204

		14.551		0.00952		0.023655914

		14.561		0.0095		0.0215053763

		14.571		0.00952		0.023655914

		14.581		0.00952		0.023655914

		14.591		0.00955		0.0268817204

		14.601		0.00955		0.0268817204

		14.611		0.0095		0.0215053763

		14.621		0.00955		0.0268817204

		14.631		0.00955		0.0268817204

		14.641		0.00952		0.023655914

		14.651		0.00952		0.023655914

		14.661		0.00955		0.0268817204

		14.671		0.0095		0.0215053763

		14.681		0.00952		0.023655914

		14.691		0.0095		0.0215053763

		14.701		0.00952		0.023655914

		14.711		0.0095		0.0215053763

		14.721		0.00947		0.0182795699

		14.731		0.00947		0.0182795699

		14.741		0.00952		0.023655914

		14.751		0.0095		0.0215053763

		14.761		0.00945		0.0161290323

		14.771		0.00947		0.0182795699

		14.781		0.00945		0.0161290323

		14.791		0.0095		0.0215053763

		14.801		0.00947		0.0182795699

		14.811		0.00947		0.0182795699

		14.821		0.00947		0.0182795699

		14.831		0.0095		0.0215053763

		14.841		0.00947		0.0182795699

		14.851		0.00947		0.0182795699

		14.861		0.00945		0.0161290323

		14.871		0.00947		0.0182795699

		14.881		0.00942		0.0129032258

		14.891		0.0095		0.0215053763

		14.901		0.00947		0.0182795699

		14.911		0.00945		0.0161290323

		14.921		0.00947		0.0182795699

		14.931		0.00947		0.0182795699

		14.941		0.00942		0.0129032258

		14.951		0.0095		0.0215053763

		14.961		0.00947		0.0182795699

		14.971		0.00947		0.0182795699

		14.981		0.00945		0.0161290323

		14.991		0.00947		0.0182795699

		15.001		0.00947		0.0182795699

		15.011		0.00947		0.0182795699

		15.021		0.00942		0.0129032258

		15.031		0.00942		0.0129032258

		15.041		0.0094		0.0107526882

		15.051		0.00945		0.0161290323

		15.061		0.00947		0.0182795699

		15.071		0.00945		0.0161290323

		15.081		0.00942		0.0129032258

		15.091		0.00942		0.0129032258

		15.101		0.00945		0.0161290323

		15.111		0.00942		0.0129032258

		15.121		0.00942		0.0129032258

		15.131		0.00947		0.0182795699

		15.141		0.00942		0.0129032258

		15.151		0.00942		0.0129032258

		15.161		0.00945		0.0161290323

		15.171		0.00942		0.0129032258

		15.181		0.00942		0.0129032258

		15.191		0.00945		0.0161290323

		15.201		0.0094		0.0107526882

		15.211		0.0094		0.0107526882

		15.221		0.00942		0.0129032258

		15.231		0.00942		0.0129032258

		15.241		0.0094		0.0107526882

		15.251		0.00947		0.0182795699

		15.261		0.00942		0.0129032258

		15.271		0.00938		0.0086021505

		15.281		0.00942		0.0129032258

		15.291		0.0094		0.0107526882

		15.301		0.0094		0.0107526882

		15.311		0.00945		0.0161290323

		15.321		0.00942		0.0129032258

		15.331		0.00945		0.0161290323

		15.341		0.00942		0.0129032258

		15.351		0.00935		0.0053763441

		15.361		0.0094		0.0107526882

		15.371		0.00942		0.0129032258

		15.381		0.00945		0.0161290323

		15.391		0.00942		0.0129032258

		15.401		0.0094		0.0107526882

		15.411		0.0094		0.0107526882

		15.421		0.0094		0.0107526882

		15.431		0.00938		0.0086021505

		15.441		0.00938		0.0086021505

		15.451		0.00942		0.0129032258

		15.461		0.00938		0.0086021505

		15.471		0.00942		0.0129032258

		15.481		0.00935		0.0053763441

		15.491		0.00938		0.0086021505

		15.501		0.00942		0.0129032258

		15.511		0.00945		0.0161290323

		15.521		0.00942		0.0129032258

		15.531		0.0094		0.0107526882

		15.541		0.00935		0.0053763441

		15.551		0.00938		0.0086021505

		15.561		0.00938		0.0086021505

		15.571		0.00938		0.0086021505

		15.581		0.0094		0.0107526882

		15.591		0.00938		0.0086021505

		15.601		0.00945		0.0161290323

		15.611		0.00938		0.0086021505

		15.621		0.0094		0.0107526882

		15.631		0.00942		0.0129032258

		15.641		0.00942		0.0129032258

		15.651		0.00942		0.0129032258

		15.661		0.0094		0.0107526882

		15.671		0.00938		0.0086021505

		15.681		0.00938		0.0086021505

		15.691		0.00942		0.0129032258

		15.701		0.00935		0.0053763441

		15.711		0.00938		0.0086021505

		15.721		0.0094		0.0107526882

		15.731		0.00938		0.0086021505

		15.741		0.00938		0.0086021505

		15.751		0.00938		0.0086021505

		15.761		0.00942		0.0129032258

		15.771		0.0094		0.0107526882

		15.781		0.0094		0.0107526882

		15.791		0.00938		0.0086021505

		15.801		0.00938		0.0086021505

		15.811		0.0094		0.0107526882

		15.821		0.00942		0.0129032258

		15.831		0.00942		0.0129032258

		15.841		0.00935		0.0053763441

		15.851		0.00938		0.0086021505

		15.861		0.00935		0.0053763441

		15.871		0.00938		0.0086021505

		15.881		0.00938		0.0086021505

		15.891		0.0094		0.0107526882

		15.901		0.00933		0.0032258065

		15.911		0.00938		0.0086021505

		15.921		0.00935		0.0053763441

		15.931		0.00942		0.0129032258

		15.941		0.00938		0.0086021505

		15.951		0.0094		0.0107526882

		15.961		0.00938		0.0086021505

		15.971		0.00938		0.0086021505

		15.981		0.00935		0.0053763441

		15.991		0.00935		0.0053763441

		16.001		0.00935		0.0053763441

		16.011		0.0094		0.0107526882

		16.02		0.0094		0.0107526882

		16.031		0.00938		0.0086021505

		16.041		0.00935		0.0053763441

		16.051		0.00935		0.0053763441

		16.061		0.00935		0.0053763441

		16.071		0.00938		0.0086021505

		16.081		0.00938		0.0086021505

		16.091		0.00935		0.0053763441

		16.1		0.00938		0.0086021505

		16.111		0.00933		0.0032258065

		16.121		0.0094		0.0107526882

		16.131		0.00935		0.0053763441

		16.141		0.00935		0.0053763441

		16.151		0.00938		0.0086021505

		16.161		0.0093		0

		16.171		0.00935		0.0053763441

		16.181		0.00935		0.0053763441

		16.19		0.00935		0.0053763441

		16.201		0.00933		0.0032258065

		16.211		0.00935		0.0053763441

		16.221		0.00938		0.0086021505

		16.231		0.0093		0

		16.241		0.00935		0.0053763441

		16.251		0.00938		0.0086021505

		16.261		0.00938		0.0086021505

		16.27		0.00933		0.0032258065

		16.281		0.00935		0.0053763441

		16.291		0.00935		0.0053763441

		16.301		0.00935		0.0053763441

		16.311		0.00935		0.0053763441

		16.321		0.00938		0.0086021505

		16.331		0.00935		0.0053763441

		16.341		0.00938		0.0086021505

		16.35		0.00935		0.0053763441

		16.361		0.00933		0.0032258065

		16.371		0.00933		0.0032258065

		16.381		0.00938		0.0086021505

		16.391		0.00938		0.0086021505

		16.401		0.00935		0.0053763441

		16.411		0.00938		0.0086021505

		16.421		0.00938		0.0086021505

		16.431		0.00933		0.0032258065

		16.44		0.00933		0.0032258065

		16.451		0.00935		0.0053763441

		16.461		0.00933		0.0032258065

		16.472		0.00935		0.0053763441

		16.481		0.00935		0.0053763441

		16.491		0.00935		0.0053763441

		16.501		0.00935		0.0053763441

		16.511		0.00935		0.0053763441

		16.52		0.00935		0.0053763441

		16.531		0.0094		0.0107526882

		16.541		0.00938		0.0086021505

		16.551		0.00933		0.0032258065

		16.561		0.0094		0.0107526882

		16.571		0.00938		0.0086021505

		16.582		0.00938		0.0086021505

		16.591		0.0094		0.0107526882

		16.601		0.0093		0

		16.611		0.00935		0.0053763441

		16.621		0.00935		0.0053763441

		16.631		0.00933		0.0032258065

		16.641		0.0094		0.0107526882

		16.651		0.00935		0.0053763441

		16.661		0.00933		0.0032258065

		16.671		0.00935		0.0053763441

		16.681		0.00935		0.0053763441

		16.692		0.00933		0.0032258065

		16.701		0.0094		0.0107526882

		16.711		0.00933		0.0032258065

		16.721		0.0093		0

		16.731		0.0093		0

		16.741		0.0093		0

		16.751		0.00938		0.0086021505

		16.761		0.00935		0.0053763441

		16.77		0.00935		0.0053763441

		16.781		0.00933		0.0032258065

		16.791		0.0093		0

		16.802		0.00938		0.0086021505

		16.811		0.00933		0.0032258065

		16.821		0.00935		0.0053763441

		16.831		0.00933		0.0032258065

		16.841		0.00938		0.0086021505

		16.851		0.00935		0.0053763441

		16.861		0.00942		0.0129032258

		16.871		0.00933		0.0032258065

		16.881		0.00938		0.0086021505

		16.891		0.00935		0.0053763441

		16.901		0.00935		0.0053763441

		16.912		0.00935		0.0053763441

		16.921		0.0093		0

		16.931		0.00933		0.0032258065

		16.94		0.00933		0.0032258065

		16.951		0.00933		0.0032258065

		16.961		0.00935		0.0053763441

		16.971		0.00935		0.0053763441

		16.981		0.00935		0.0053763441

		16.991		0.00938		0.0086021505

		17.001		0.00938		0.0086021505

		17.011		0.00938		0.0086021505

		17.022		0.00933		0.0032258065

		17.031		0.00933		0.0032258065

		17.041		0.00933		0.0032258065

		17.051		0.00935		0.0053763441

		17.061		0.0094		0.0107526882

		17.071		0.00938		0.0086021505

		17.081		0.00935		0.0053763441

		17.091		0.00935		0.0053763441

		17.101		0.00933		0.0032258065

		17.111		0.00935		0.0053763441

		17.121		0.0094		0.0107526882

		17.131		0.00933		0.0032258065

		17.141		0.00964		0.0365591398

		17.151		0.00935		0.0053763441

		17.161		0.00933		0.0032258065

		17.171		0.00933		0.0032258065

		17.181		0.00935		0.0053763441

		17.19		0.00935		0.0053763441

		17.201		0.00933		0.0032258065

		17.211		0.00933		0.0032258065

		17.221		0.00935		0.0053763441

		17.231		0.00938		0.0086021505

		17.241		0.00938		0.0086021505

		17.251		0.00935		0.0053763441

		17.261		0.00933		0.0032258065

		17.27		0.0093		0

		17.281		0.00933		0.0032258065

		17.291		0.00933		0.0032258065

		17.301		0.00933		0.0032258065

		17.311		0.0094		0.0107526882

		17.321		0.00935		0.0053763441

		17.331		0.0094		0.0107526882

		17.341		0.00933		0.0032258065

		17.351		0.0093		0

		17.361		0.00933		0.0032258065

		17.371		0.0093		0

		17.381		0.00935		0.0053763441

		17.391		0.00933		0.0032258065

		17.401		0.00933		0.0032258065

		17.411		0.00938		0.0086021505

		17.421		0.00933		0.0032258065

		17.431		0.00935		0.0053763441

		17.44		0.00933		0.0032258065

		17.451		0.00935		0.0053763441

		17.461		0.00935		0.0053763441

		17.471		0.0094		0.0107526882

		17.481		0.00935		0.0053763441

		17.491		0.00938		0.0086021505

		17.501		0.0094		0.0107526882

		17.511		0.00933		0.0032258065

		17.52		0.00933		0.0032258065

		17.531		0.00933		0.0032258065

		17.541		0.00933		0.0032258065

		17.551		0.00933		0.0032258065

		17.561		0.00935		0.0053763441

		17.571		0.00933		0.0032258065

		17.581		0.00935		0.0053763441

		17.591		0.00933		0.0032258065

		17.601		0.00933		0.0032258065

		17.611		0.00938		0.0086021505

		17.621		0.00935		0.0053763441

		17.631		0.00935		0.0053763441

		17.641		0.00933		0.0032258065

		17.651		0.00933		0.0032258065

		17.661		0.00933		0.0032258065

		17.671		0.00938		0.0086021505

		17.681		0.00933		0.0032258065

		17.69		0.00933		0.0032258065

		17.701		0.00935		0.0053763441

		17.711		0.00933		0.0032258065

		17.721		0.00933		0.0032258065

		17.731		0.00938		0.0086021505

		17.741		0.0094		0.0107526882

		17.751		0.00935		0.0053763441

		17.761		0.00938		0.0086021505

		17.77		0.00933		0.0032258065

		17.781		0.00933		0.0032258065

		17.791		0.00933		0.0032258065

		17.801		0.00933		0.0032258065

		17.811		0.00933		0.0032258065

		17.821		0.00933		0.0032258065

		17.831		0.00933		0.0032258065

		17.841		0.00933		0.0032258065

		17.851		0.00933		0.0032258065

		17.861		0.00933		0.0032258065

		17.871		0.00935		0.0053763441

		17.881		0.00933		0.0032258065

		17.891		0.00938		0.0086021505

		17.901		0.00933		0.0032258065

		17.911		0.00933		0.0032258065

		17.921		0.00933		0.0032258065

		17.931		0.0093		0

		17.94		0.00933		0.0032258065

		17.951		0.00933		0.0032258065

		17.961		0.00933		0.0032258065

		17.971		0.0093		0

		17.981		0.00933		0.0032258065

		17.991		0.00933		0.0032258065

		18.001		0.00933		0.0032258065

		18.011		0.0093		0

		18.02		0.00933		0.0032258065

		18.031		0.0094		0.0107526882

		18.041		0.00933		0.0032258065

		18.051		0.00938		0.0086021505

		18.061		0.00935		0.0053763441

		18.071		0.00935		0.0053763441

		18.081		0.00933		0.0032258065

		18.091		0.00933		0.0032258065

		18.101		0.00938		0.0086021505

		18.111		0.00933		0.0032258065

		18.121		0.00933		0.0032258065

		18.131		0.0093		0

		18.141		0.00933		0.0032258065

		18.151		0.00935		0.0053763441

		18.161		0.00938		0.0086021505

		18.171		0.00938		0.0086021505

		18.181		0.00935		0.0053763441

		18.19		0.00935		0.0053763441

		18.201		0.00933		0.0032258065

		18.211		0.0093		0

		18.221		0.0094		0.0107526882

		18.231		0.00933		0.0032258065

		18.241		0.00933		0.0032258065

		18.251		0.00938		0.0086021505

		18.261		0.00933		0.0032258065

		18.271		0.00933		0.0032258065

		18.281		0.00933		0.0032258065

		18.291		0.00933		0.0032258065

		18.301		0.00935		0.0053763441

		18.311		0.0093		0

		18.321		0.00933		0.0032258065

		18.331		0.00933		0.0032258065

		18.341		0.00938		0.0086021505

		18.351		0.00935		0.0053763441

		18.361		0.00933		0.0032258065

		18.371		0.00933		0.0032258065

		18.381		0.00933		0.0032258065

		18.391		0.00933		0.0032258065

		18.401		0.00938		0.0086021505

		18.411		0.00938		0.0086021505

		18.421		0.00933		0.0032258065

		18.431		0.00933		0.0032258065

		18.44		0.00933		0.0032258065

		18.451		0.0093		0

		18.461		0.00933		0.0032258065

		18.471		0.00935		0.0053763441

		18.481		0.00938		0.0086021505

		18.491		0.00933		0.0032258065

		18.501		0.00935		0.0053763441

		18.511		0.00933		0.0032258065

		18.521		0.00933		0.0032258065

		18.531		0.00933		0.0032258065

		18.541		0.0093		0

		18.551		0.00928		-0.0021505376

		18.561		0.00933		0.0032258065

		18.571		0.00928		-0.0021505376

		18.581		0.00938		0.0086021505

		18.591		0.00933		0.0032258065

		18.601		0.00933		0.0032258065

		18.611		0.00938		0.0086021505

		18.621		0.0093		0

		18.631		0.00935		0.0053763441

		18.641		0.00933		0.0032258065

		18.651		0.00938		0.0086021505

		18.661		0.00935		0.0053763441

		18.671		0.00938		0.0086021505

		18.681		0.00935		0.0053763441

		18.69		0.00933		0.0032258065

		18.701		0.00938		0.0086021505

		18.711		0.0093		0

		18.721		0.00933		0.0032258065

		18.731		0.00933		0.0032258065

		18.741		0.00933		0.0032258065

		18.751		0.00933		0.0032258065

		18.761		0.0093		0

		18.771		0.0093		0

		18.781		0.0093		0

		18.791		0.00933		0.0032258065

		18.801		0.00933		0.0032258065

		18.811		0.00933		0.0032258065

		18.821		0.0093		0

		18.831		0.00933		0.0032258065

		18.841		0.00935		0.0053763441

		18.851		0.00933		0.0032258065

		18.861		0.00938		0.0086021505

		18.871		0.0093		0

		18.881		0.0093		0

		18.891		0.00933		0.0032258065

		18.901		0.00938		0.0086021505

		18.911		0.00933		0.0032258065

		18.921		0.00933		0.0032258065

		18.931		0.00935		0.0053763441

		18.94		0.00938		0.0086021505

		18.951		0.00933		0.0032258065

		18.961		0.00933		0.0032258065

		18.971		0.00933		0.0032258065

		18.981		0.00933		0.0032258065

		18.991		0.00933		0.0032258065

		19.001		0.00938		0.0086021505

		19.011		0.00938		0.0086021505

		19.021		0.0093		0

		19.031		0.00935		0.0053763441

		19.041		0.00933		0.0032258065

		19.051		0.00938		0.0086021505

		19.061		0.00935		0.0053763441

		19.071		0.00933		0.0032258065

		19.081		0.00933		0.0032258065

		19.091		0.00933		0.0032258065

		19.101		0.00933		0.0032258065

		19.111		0.00933		0.0032258065

		19.121		0.00933		0.0032258065

		19.131		0.00935		0.0053763441

		19.141		0.00933		0.0032258065

		19.151		0.00933		0.0032258065

		19.161		0.00933		0.0032258065

		19.171		0.00935		0.0053763441

		19.181		0.00933		0.0032258065

		19.19		0.0093		0

		19.201		0.0093		0

		19.211		0.00933		0.0032258065

		19.221		0.00938		0.0086021505

		19.231		0.00933		0.0032258065

		19.241		0.00935		0.0053763441

		19.251		0.00933		0.0032258065

		19.261		0.00933		0.0032258065

		19.271		0.00938		0.0086021505

		19.281		0.0093		0

		19.291		0.00938		0.0086021505

		19.301		0.00933		0.0032258065

		19.311		0.00935		0.0053763441

		19.321		0.00935		0.0053763441

		19.331		0.00938		0.0086021505

		19.341		0.00942		0.0129032258

		19.351		0.0095		0.0215053763

		19.361		0.00959		0.0311827957

		19.371		0.00969		0.0419354839

		19.381		0.00986		0.0602150538

		19.391		0.01003		0.0784946237

		19.401		0.0103		0.1075268817

		19.411		0.01052		0.1311827957

		19.421		0.01079		0.1602150538

		19.431		0.01113		0.1967741935

		19.44		0.01145		0.2311827957

		19.451		0.01177		0.2655913978

		19.461		0.01208		0.2989247312

		19.471		0.01233		0.3258064516

		19.481		0.01265		0.3602150538

		19.491		0.01294		0.3913978495

		19.501		0.01321		0.4204301075

		19.511		0.01348		0.4494623656

		19.521		0.01377		0.4806451613

		19.531		0.01401		0.5064516129

		19.541		0.01426		0.5333333333

		19.551		0.01445		0.5537634409

		19.561		0.01465		0.5752688172

		19.571		0.01487		0.5989247312

		19.581		0.01504		0.6172043011

		19.591		0.01519		0.6333333333

		19.601		0.01536		0.6516129032

		19.611		0.01553		0.6698924731

		19.621		0.01572		0.6903225806

		19.631		0.01577		0.6956989247

		19.641		0.01592		0.711827957

		19.651		0.01604		0.7247311828

		19.661		0.01619		0.7408602151

		19.671		0.01626		0.7483870968

		19.681		0.01638		0.7612903226

		19.69		0.01646		0.7698924731

		19.701		0.01653		0.7774193548

		19.711		0.01663		0.788172043

		19.721		0.0167		0.7956989247

		19.731		0.01677		0.8032258065

		19.741		0.01682		0.8086021505

		19.751		0.01692		0.8193548387

		19.761		0.01699		0.8268817204

		19.771		0.01702		0.8301075269

		19.781		0.01709		0.8376344086

		19.791		0.01716		0.8451612903

		19.801		0.01721		0.8505376344

		19.811		0.01721		0.8505376344

		19.821		0.01738		0.8688172043

		19.831		0.01736		0.8666666667

		19.841		0.01743		0.8741935484

		19.851		0.01746		0.8774193548

		19.861		0.0175		0.8817204301

		19.871		0.01748		0.8795698925

		19.881		0.01753		0.8849462366

		19.891		0.01758		0.8903225806

		19.901		0.01763		0.8956989247

		19.911		0.01763		0.8956989247

		19.921		0.01765		0.8978494624

		19.931		0.0177		0.9032258065

		19.941		0.0178		0.9139784946

		19.951		0.01775		0.9086021505

		19.961		0.01777		0.9107526882

		19.971		0.01782		0.9161290323

		19.981		0.01782		0.9161290323

		19.991		0.01787		0.9215053763

		20.001		0.01787		0.9215053763

		20.011		0.0179		0.9247311828

		20.021		0.0179		0.9247311828

		20.031		0.0179		0.9247311828

		20.041		0.01792		0.9268817204

		20.051		0.01797		0.9322580645

		20.061		0.01797		0.9322580645

		20.071		0.01799		0.9344086022

		20.081		0.01807		0.9430107527

		20.091		0.01799		0.9344086022

		20.101		0.01809		0.9451612903

		20.111		0.01807		0.9430107527

		20.121		0.01807		0.9430107527

		20.131		0.01812		0.9483870968

		20.141		0.01807		0.9430107527

		20.151		0.01809		0.9451612903

		20.161		0.01812		0.9483870968

		20.171		0.01812		0.9483870968

		20.181		0.01814		0.9505376344

		20.191		0.01819		0.9559139785

		20.201		0.01816		0.952688172

		20.211		0.01816		0.952688172

		20.221		0.01816		0.952688172

		20.231		0.01821		0.9580645161

		20.241		0.01821		0.9580645161

		20.251		0.01824		0.9612903226

		20.261		0.01819		0.9559139785

		20.271		0.01824		0.9612903226

		20.281		0.01821		0.9580645161

		20.291		0.01824		0.9612903226

		20.301		0.01821		0.9580645161

		20.311		0.01826		0.9634408602

		20.321		0.01826		0.9634408602

		20.331		0.01829		0.9666666667

		20.341		0.01826		0.9634408602

		20.351		0.01829		0.9666666667

		20.361		0.01826		0.9634408602

		20.371		0.01831		0.9688172043

		20.381		0.01829		0.9666666667

		20.391		0.01833		0.9709677419

		20.401		0.01831		0.9688172043

		20.411		0.01831		0.9688172043

		20.421		0.01833		0.9709677419

		20.431		0.01831		0.9688172043

		20.441		0.01833		0.9709677419

		20.451		0.01833		0.9709677419

		20.461		0.01836		0.9741935484

		20.471		0.01836		0.9741935484

		20.481		0.01833		0.9709677419

		20.493		0.01838		0.976344086

		20.503		0.01836		0.9741935484

		20.513		0.01836		0.9741935484

		20.523		0.01841		0.9795698925

		20.533		0.01841		0.9795698925

		20.543		0.01838		0.976344086

		20.553		0.01838		0.976344086

		20.563		0.01846		0.9849462366

		20.573		0.01841		0.9795698925

		20.583		0.01838		0.976344086

		20.593		0.01838		0.976344086

		20.603		0.01838		0.976344086

		20.613		0.01841		0.9795698925

		20.623		0.01846		0.9849462366

		20.633		0.01846		0.9849462366

		20.643		0.01841		0.9795698925

		20.653		0.01843		0.9817204301

		20.663		0.01841		0.9795698925

		20.673		0.01843		0.9817204301

		20.683		0.01841		0.9795698925

		20.693		0.01843		0.9817204301

		20.703		0.01843		0.9817204301

		20.713		0.01843		0.9817204301

		20.723		0.01843		0.9817204301

		20.733		0.01843		0.9817204301

		20.743		0.01841		0.9795698925

		20.753		0.01843		0.9817204301

		20.763		0.01851		0.9903225806

		20.773		0.01846		0.9849462366

		20.783		0.01851		0.9903225806

		20.793		0.01843		0.9817204301

		20.803		0.01846		0.9849462366

		20.813		0.01846		0.9849462366

		20.823		0.01846		0.9849462366

		20.833		0.01851		0.9903225806

		20.843		0.01843		0.9817204301

		20.853		0.01851		0.9903225806

		20.863		0.01846		0.9849462366

		20.873		0.01853		0.9924731183

		20.883		0.01851		0.9903225806

		20.893		0.01851		0.9903225806

		20.903		0.01848		0.9870967742

		20.913		0.01846		0.9849462366

		20.923		0.01851		0.9903225806

		20.933		0.01848		0.9870967742

		20.943		0.01851		0.9903225806

		20.953		0.01855		0.9946236559

		20.963		0.01853		0.9924731183

		20.973		0.01843		0.9817204301

		20.983		0.01853		0.9924731183

		20.993		0.01848		0.9870967742

		21.003		0.01848		0.9870967742

		21.013		0.01853		0.9924731183

		21.023		0.01846		0.9849462366

		21.033		0.01853		0.9924731183

		21.043		0.01846		0.9849462366

		21.053		0.01851		0.9903225806

		21.063		0.01848		0.9870967742

		21.073		0.01848		0.9870967742

		21.083		0.01851		0.9903225806

		21.093		0.01853		0.9924731183

		21.103		0.01851		0.9903225806

		21.113		0.01851		0.9903225806

		21.123		0.01858		0.9978494624

		21.133		0.01851		0.9903225806

		21.143		0.01853		0.9924731183

		21.153		0.01855		0.9946236559

		21.163		0.01851		0.9903225806

		21.173		0.01855		0.9946236559

		21.183		0.01855		0.9946236559

		21.193		0.01851		0.9903225806

		21.203		0.01851		0.9903225806

		21.213		0.01853		0.9924731183

		21.223		0.01853		0.9924731183

		21.233		0.01851		0.9903225806

		21.243		0.01851		0.9903225806

		21.253		0.01851		0.9903225806

		21.263		0.01851		0.9903225806

		21.273		0.01851		0.9903225806

		21.283		0.01851		0.9903225806

		21.293		0.01851		0.9903225806

		21.303		0.01853		0.9924731183

		21.313		0.01855		0.9946236559

		21.323		0.01853		0.9924731183

		21.333		0.01855		0.9946236559

		21.343		0.01853		0.9924731183

		21.353		0.01853		0.9924731183

		21.363		0.01853		0.9924731183

		21.373		0.01853		0.9924731183

		21.383		0.01855		0.9946236559

		21.393		0.01853		0.9924731183

		21.403		0.01853		0.9924731183

		21.413		0.01855		0.9946236559

		21.423		0.01853		0.9924731183

		21.433		0.01851		0.9903225806

		21.443		0.01855		0.9946236559

		21.453		0.01853		0.9924731183

		21.463		0.01855		0.9946236559

		21.473		0.01853		0.9924731183

		21.483		0.01853		0.9924731183

		21.493		0.01853		0.9924731183

		21.503		0.01855		0.9946236559

		21.513		0.01855		0.9946236559

		21.523		0.01855		0.9946236559

		21.533		0.01855		0.9946236559

		21.543		0.01853		0.9924731183

		21.553		0.01855		0.9946236559

		21.563		0.0186		1

		21.573		0.01855		0.9946236559

		21.583		0.01858		0.9978494624

		21.593		0.01855		0.9946236559

		21.603		0.01858		0.9978494624

		21.613		0.01855		0.9946236559

		21.623		0.01855		0.9946236559

		21.633		0.0186		1

		21.643		0.01858		0.9978494624

		21.653		0.01853		0.9924731183

		21.663		0.01858		0.9978494624

		21.673		0.01853		0.9924731183

		21.683		0.01855		0.9946236559

		21.693		0.01855		0.9946236559

		21.703		0.01855		0.9946236559

		21.713		0.01855		0.9946236559

		21.723		0.01853		0.9924731183

		21.733		0.01858		0.9978494624

		21.743		0.01855		0.9946236559

		21.753		0.01858		0.9978494624

		21.763		0.01853		0.9924731183

		21.773		0.01855		0.9946236559

		21.783		0.01853		0.9924731183

		21.793		0.01855		0.9946236559

		21.803		0.01853		0.9924731183

		21.813		0.01858		0.9978494624

		21.823		0.01863		1.0032258065

		21.833		0.0186		1

		21.843		0.01855		0.9946236559

		21.853		0.0186		1

		21.863		0.01855		0.9946236559

		21.873		0.0186		1

		21.883		0.0186		1

		21.893		0.01853		0.9924731183

		21.903		0.01858		0.9978494624

		21.913		0.01855		0.9946236559

		21.923		0.01855		0.9946236559

		21.933		0.01855		0.9946236559

		21.943		0.01855		0.9946236559

		21.953		0.01855		0.9946236559

		21.963		0.01858		0.9978494624

		21.973		0.01853		0.9924731183

		21.983		0.01858		0.9978494624

		21.993		0.0186		1

		22.003		0.01858		0.9978494624

		22.013		0.0186		1

		22.023		0.01855		0.9946236559

		22.033		0.01855		0.9946236559

		22.043		0.01855		0.9946236559

		22.053		0.01855		0.9946236559

		22.063		0.01855		0.9946236559

		22.073		0.01865		1.0053763441

		22.083		0.01853		0.9924731183

		22.093		0.01855		0.9946236559

		22.103		0.0186		1

		22.113		0.01858		0.9978494624

		22.123		0.01853		0.9924731183

		22.133		0.01855		0.9946236559

		22.143		0.01855		0.9946236559

		22.153		0.0186		1

		22.163		0.01858		0.9978494624

		22.173		0.01855		0.9946236559

		22.183		0.01863		1.0032258065

		22.193		0.01858		0.9978494624

		22.203		0.01855		0.9946236559

		22.213		0.01858		0.9978494624

		22.223		0.01855		0.9946236559

		22.233		0.0186		1

		22.243		0.01855		0.9946236559

		22.253		0.01853		0.9924731183

		22.263		0.0186		1

		22.273		0.01858		0.9978494624

		22.283		0.01863		1.0032258065

		22.293		0.0186		1

		22.303		0.01858		0.9978494624

		22.313		0.01855		0.9946236559

		22.323		0.01855		0.9946236559

		22.333		0.0186		1

		22.343		0.01858		0.9978494624

		22.353		0.01855		0.9946236559

		22.363		0.01855		0.9946236559

		22.373		0.01858		0.9978494624

		22.383		0.01858		0.9978494624

		22.393		0.01858		0.9978494624

		22.403		0.01858		0.9978494624

		22.413		0.01855		0.9946236559

		22.423		0.0186		1

		22.433		0.01855		0.9946236559

		22.443		0.01858		0.9978494624

		22.453		0.01858		0.9978494624

		22.463		0.01855		0.9946236559

		22.473		0.0186		1

		22.483		0.01858		0.9978494624

		22.493		0.0186		1

		22.503		0.01855		0.9946236559

		22.513		0.0186		1

		22.523		0.01858		0.9978494624

		22.533		0.01858		0.9978494624

		22.543		0.01855		0.9946236559

		22.553		0.0186		1

		22.563		0.0186		1

		22.573		0.0186		1

		22.583		0.01858		0.9978494624

		22.593		0.01865		1.0053763441

		22.603		0.01853		0.9924731183

		22.613		0.01855		0.9946236559

		22.623		0.01863		1.0032258065

		22.633		0.0186		1

		22.643		0.01863		1.0032258065

		22.653		0.01858		0.9978494624

		22.663		0.01863		1.0032258065

		22.673		0.01855		0.9946236559

		22.683		0.0186		1

		22.693		0.01863		1.0032258065

		22.703		0.01858		0.9978494624

		22.713		0.0186		1

		22.723		0.01858		0.9978494624

		22.733		0.01858		0.9978494624

		22.743		0.01858		0.9978494624

		22.753		0.01858		0.9978494624

		22.763		0.01858		0.9978494624

		22.773		0.01855		0.9946236559

		22.783		0.0186		1

		22.793		0.0186		1

		22.803		0.01863		1.0032258065

		22.813		0.01865		1.0053763441

		22.823		0.0186		1

		22.833		0.01858		0.9978494624

		22.843		0.01855		0.9946236559

		22.853		0.0186		1

		22.863		0.01863		1.0032258065

		22.873		0.01858		0.9978494624

		22.883		0.0186		1

		22.893		0.01858		0.9978494624

		22.903		0.01863		1.0032258065

		22.913		0.0186		1

		22.923		0.0186		1

		22.933		0.01855		0.9946236559

		22.943		0.01858		0.9978494624

		22.953		0.01858		0.9978494624

		22.963		0.0186		1

		22.973		0.01858		0.9978494624

		22.983		0.0186		1

		22.997		0.0186		1

		23.003		0.0186		1

		23.013		0.0186		1

		23.023		0.01855		0.9946236559

		23.033		0.01865		1.0053763441

		23.043		0.01858		0.9978494624

		23.053		0.01865		1.0053763441

		23.064		0.01865		1.0053763441

		23.073		0.0186		1

		23.083		0.01858		0.9978494624

		23.093		0.01855		0.9946236559

		23.103		0.01858		0.9978494624

		23.113		0.01865		1.0053763441

		23.123		0.0186		1

		23.133		0.0186		1

		23.143		0.01863		1.0032258065

		23.153		0.01858		0.9978494624

		23.163		0.01858		0.9978494624

		23.174		0.01858		0.9978494624

		23.183		0.01858		0.9978494624

		23.193		0.01863		1.0032258065

		23.203		0.01858		0.9978494624

		23.213		0.01863		1.0032258065

		23.223		0.01858		0.9978494624

		23.233		0.0186		1

		23.243		0.01858		0.9978494624

		23.253		0.01858		0.9978494624

		23.263		0.01858		0.9978494624

		23.273		0.01858		0.9978494624

		23.284		0.01865		1.0053763441

		23.293		0.0186		1

		23.303		0.01863		1.0032258065

		23.313		0.0186		1

		23.323		0.01855		0.9946236559

		23.333		0.0186		1

		23.343		0.01865		1.0053763441

		23.353		0.01858		0.9978494624

		23.363		0.0186		1

		23.373		0.01858		0.9978494624

		23.383		0.01858		0.9978494624

		23.394		0.01855		0.9946236559

		23.403		0.0186		1

		23.413		0.01863		1.0032258065

		23.423		0.01858		0.9978494624

		23.433		0.01855		0.9946236559

		23.443		0.0186		1

		23.453		0.0186		1

		23.463		0.01858		0.9978494624

		23.473		0.0186		1

		23.483		0.01858		0.9978494624

		23.493		0.01865		1.0053763441

		23.504		0.0186		1

		23.513		0.0186		1

		23.523		0.0186		1

		23.533		0.0186		1

		23.543		0.01865		1.0053763441

		23.553		0.01858		0.9978494624

		23.563		0.0186		1

		23.573		0.01865		1.0053763441

		23.583		0.01858		0.9978494624

		23.593		0.01863		1.0032258065

		23.603		0.0186		1

		23.614		0.01858		0.9978494624

		23.623		0.0186		1

		23.633		0.0186		1

		23.643		0.01858		0.9978494624

		23.653		0.0186		1

		23.663		0.01858		0.9978494624

		23.673		0.01865		1.0053763441

		23.683		0.01858		0.9978494624

		23.693		0.0186		1

		23.703		0.0186		1

		23.713		0.01858		0.9978494624

		23.724		0.01858		0.9978494624

		23.733		0.01858		0.9978494624

		23.743		0.0186		1

		23.753		0.01858		0.9978494624

		23.763		0.01858		0.9978494624

		23.773		0.01858		0.9978494624

		23.783		0.0186		1

		23.793		0.0186		1

		23.803		0.0186		1

		23.813		0.0186		1

		23.823		0.0186		1

		23.833		0.0186		1

		23.843		0.0186		1

		23.853		0.0186		1

		23.863		0.01858		0.9978494624

		23.873		0.0186		1

		23.883		0.0186		1

		23.893		0.0186		1

		23.903		0.0186		1

		23.913		0.0186		1

		23.923		0.0186		1

		23.933		0.0186		1

		23.943		0.0186		1

		23.953		0.01863		1.0032258065

		23.963		0.01865		1.0053763441

		23.973		0.01858		0.9978494624

		23.983		0.0186		1

		23.993		0.01863		1.0032258065

		24.003		0.01858		0.9978494624

		24.013		0.0186		1

		24.023		0.0186		1

		24.033		0.0186		1

		24.043		0.0186		1

		24.053		0.01858		0.9978494624

		24.063		0.0186		1

		24.073		0.0186		1

		24.083		0.01865		1.0053763441

		24.093		0.01858		0.9978494624

		24.103		0.0186		1

		24.113		0.0186		1

		24.123		0.01865		1.0053763441

		24.133		0.0186		1

		24.143		0.01865		1.0053763441

		24.153		0.0186		1

		24.163		0.0186		1

		24.173		0.01858		0.9978494624

		24.183		0.01863		1.0032258065

		24.193		0.0186		1

		24.203		0.0186		1

		24.213		0.01858		0.9978494624

		24.223		0.01865		1.0053763441

		24.233		0.0186		1

		24.243		0.0186		1

		24.253		0.01858		0.9978494624

		24.263		0.0186		1

		24.273		0.01858		0.9978494624

		24.283		0.0187		1.0107526882

		24.293		0.01865		1.0053763441

		24.303		0.0186		1

		24.313		0.0186		1

		24.323		0.01858		0.9978494624

		24.333		0.01858		0.9978494624

		24.343		0.0186		1

		24.353		0.0186		1

		24.363		0.0186		1

		24.373		0.0186		1

		24.383		0.01865		1.0053763441

		24.393		0.01865		1.0053763441

		24.403		0.0186		1

		24.413		0.0186		1

		24.423		0.01858		0.9978494624

		24.433		0.01858		0.9978494624

		24.443		0.0186		1

		24.453		0.01863		1.0032258065

		24.463		0.0186		1

		24.473		0.01858		0.9978494624

		24.483		0.01858		0.9978494624

		24.493		0.0186		1

		24.503		0.0186		1

		24.513		0.01865		1.0053763441

		24.523		0.01858		0.9978494624

		24.533		0.0186		1

		24.543		0.0186		1

		24.553		0.0186		1

		24.563		0.01863		1.0032258065

		24.573		0.01858		0.9978494624

		24.583		0.0186		1

		24.593		0.01865		1.0053763441

		24.603		0.0186		1

		24.613		0.0186		1

		24.623		0.0186		1

		24.633		0.01858		0.9978494624

		24.643		0.0186		1

		24.653		0.01865		1.0053763441

		24.663		0.01865		1.0053763441

		24.673		0.0186		1

		24.683		0.01858		0.9978494624

		24.693		0.0186		1

		24.703		0.01858		0.9978494624

		24.713		0.01858		0.9978494624

		24.723		0.0186		1

		24.733		0.0186		1

		24.743		0.01865		1.0053763441

		24.753		0.01865		1.0053763441

		24.763		0.01865		1.0053763441

		24.773		0.01863		1.0032258065

		24.783		0.01865		1.0053763441

		24.793		0.01863		1.0032258065

		24.803		0.01863		1.0032258065

		24.813		0.0186		1

		24.823		0.0186		1

		24.833		0.0186		1

		24.843		0.0186		1

		24.853		0.0186		1

		24.863		0.0186		1

		24.873		0.0186		1

		24.883		0.0186		1

		24.893		0.0186		1

		24.903		0.01863		1.0032258065

		24.913		0.01863		1.0032258065

		24.923		0.01863		1.0032258065

		24.933		0.01865		1.0053763441

		24.943		0.0186		1

		24.953		0.01863		1.0032258065

		24.963		0.0186		1

		24.973		0.01863		1.0032258065

		24.983		0.0186		1

		24.993		0.01858		0.9978494624

		25.003		0.0186		1

		25.013		0.0186		1

		25.023		0.01863		1.0032258065

		25.033		0.0186		1

		25.043		0.0186		1

		25.053		0.01863		1.0032258065

		25.063		0.01863		1.0032258065

		25.073		0.01865		1.0053763441

		25.083		0.0186		1

		25.093		0.0186		1

		25.103		0.0186		1

		25.113		0.0186		1

		25.123		0.0187		1.0107526882

		25.133		0.01863		1.0032258065

		25.143		0.0186		1

		25.153		0.0186		1

		25.163		0.0186		1

		25.173		0.01858		0.9978494624

		25.183		0.0186		1

		25.193		0.0186		1

		25.203		0.01868		1.0086021505

		25.213		0.01865		1.0053763441

		25.223		0.01865		1.0053763441

		25.233		0.0186		1

		25.243		0.01868		1.0086021505

		25.253		0.01858		0.9978494624

		25.263		0.0186		1

		25.273		0.01868		1.0086021505

		25.283		0.01858		0.9978494624

		25.293		0.01863		1.0032258065

		25.303		0.0186		1

		25.313		0.01868		1.0086021505

		25.323		0.0186		1

		25.333		0.01865		1.0053763441

		25.343		0.01858		0.9978494624

		25.353		0.01865		1.0053763441

		25.363		0.01868		1.0086021505

		25.373		0.01858		0.9978494624

		25.383		0.01863		1.0032258065

		25.393		0.01863		1.0032258065

		25.403		0.01863		1.0032258065

		25.413		0.0186		1

		25.423		0.0186		1

		25.433		0.01858		0.9978494624

		25.443		0.01865		1.0053763441

		25.453		0.01858		0.9978494624

		25.463		0.01853		0.9924731183

		25.473		0.01848		0.9870967742

		25.483		0.01836		0.9741935484

		25.493		0.01831		0.9688172043

		25.503		0.01812		0.9483870968

		25.513		0.01787		0.9215053763

		25.523		0.01755		0.8870967742

		25.533		0.01736		0.8666666667

		25.543		0.01707		0.835483871

		25.553		0.0167		0.7956989247

		25.563		0.01638		0.7612903226

		25.573		0.01611		0.7322580645

		25.583		0.01577		0.6956989247

		25.593		0.01545		0.6612903226

		25.603		0.01516		0.6301075269

		25.613		0.01484		0.5956989247

		25.623		0.01455		0.564516129

		25.633		0.01428		0.535483871

		25.643		0.01404		0.5096774194

		25.653		0.01382		0.4860215054

		25.663		0.0136		0.4623655914

		25.673		0.01338		0.4387096774

		25.683		0.01326		0.4258064516

		25.693		0.01301		0.3989247312

		25.703		0.01279		0.3752688172

		25.713		0.01265		0.3602150538

		25.723		0.01255		0.3494623656

		25.733		0.01233		0.3258064516

		25.743		0.01221		0.3129032258

		25.753		0.01206		0.2967741935

		25.763		0.01196		0.2860215054

		25.773		0.01184		0.2731182796

		25.783		0.01174		0.2623655914

		25.793		0.01162		0.2494623656

		25.803		0.0115		0.2365591398

		25.813		0.01143		0.2290322581

		25.823		0.01135		0.2204301075

		25.833		0.0113		0.2150537634

		25.843		0.01113		0.1967741935

		25.853		0.01111		0.1946236559

		25.863		0.01106		0.1892473118

		25.873		0.01099		0.1817204301

		25.883		0.01094		0.176344086

		25.893		0.01084		0.1655913978

		25.903		0.01079		0.1602150538

		25.913		0.01074		0.1548387097

		25.923		0.01067		0.147311828

		25.933		0.01069		0.1494623656

		25.943		0.01062		0.1419354839

		25.953		0.01062		0.1419354839

		25.963		0.01052		0.1311827957

		25.973		0.01047		0.1258064516

		25.983		0.01045		0.123655914

		25.993		0.0104		0.1182795699

		26.003		0.0104		0.1182795699

		26.013		0.01038		0.1161290323

		26.023		0.01033		0.1107526882

		26.033		0.01028		0.1053763441

		26.043		0.01028		0.1053763441

		26.053		0.0103		0.1075268817

		26.063		0.01021		0.0978494624

		26.073		0.01018		0.0946236559

		26.083		0.01011		0.0870967742

		26.093		0.01011		0.0870967742

		26.103		0.01008		0.0838709677

		26.113		0.01008		0.0838709677

		26.123		0.01003		0.0784946237

		26.133		0.01006		0.0817204301

		26.143		0.01001		0.076344086

		26.153		0.00999		0.0741935484

		26.163		0.00999		0.0741935484

		26.173		0.00999		0.0741935484

		26.183		0.00999		0.0741935484

		26.193		0.00996		0.0709677419

		26.203		0.00994		0.0688172043

		26.213		0.00996		0.0709677419

		26.223		0.00989		0.0634408602

		26.233		0.00989		0.0634408602

		26.243		0.00989		0.0634408602

		26.253		0.00984		0.0580645161

		26.263		0.00984		0.0580645161

		26.273		0.00981		0.0548387097

		26.283		0.00984		0.0580645161

		26.293		0.00979		0.052688172

		26.303		0.00984		0.0580645161

		26.313		0.00981		0.0548387097

		26.323		0.00979		0.052688172

		26.333		0.00977		0.0505376344

		26.343		0.00974		0.047311828

		26.353		0.00974		0.047311828

		26.363		0.00974		0.047311828

		26.373		0.00974		0.047311828

		26.383		0.00974		0.047311828

		26.393		0.00972		0.0451612903

		26.403		0.00972		0.0451612903

		26.413		0.00974		0.047311828

		26.423		0.00969		0.0419354839

		26.433		0.00969		0.0419354839

		26.443		0.00972		0.0451612903

		26.453		0.00977		0.0505376344

		26.463		0.00986		0.0602150538

		26.473		0.00991		0.0655913978

		26.483		0.01003		0.0784946237

		26.493		0.01016		0.0924731183

		26.503		0.01035		0.1129032258

		26.513		0.0106		0.1397849462

		26.523		0.01082		0.1634408602

		26.533		0.01111		0.1946236559

		26.543		0.01138		0.223655914

		26.553		0.01174		0.2623655914

		26.563		0.01204		0.2946236559

		26.573		0.0123		0.3225806452

		26.583		0.01262		0.3569892473

		26.593		0.01294		0.3913978495

		26.603		0.01318		0.4172043011

		26.613		0.01348		0.4494623656

		26.623		0.01375		0.4784946237

		26.633		0.01399		0.5043010753

		26.643		0.01423		0.5301075269

		26.653		0.01445		0.5537634409

		26.663		0.0147		0.5806451613

		26.673		0.01487		0.5989247312

		26.683		0.01509		0.6225806452

		26.693		0.01526		0.6408602151

		26.703		0.01538		0.6537634409

		26.713		0.01563		0.6806451613

		26.723		0.01572		0.6903225806

		26.733		0.01587		0.7064516129

		26.743		0.01602		0.7225806452

		26.753		0.01611		0.7322580645

		26.763		0.01626		0.7483870968

		26.773		0.01633		0.7559139785

		26.783		0.01643		0.7666666667

		26.793		0.01655		0.7795698925

		26.803		0.01665		0.7903225806

		26.813		0.01672		0.7978494624

		26.823		0.01682		0.8086021505

		26.833		0.01689		0.8161290323

		26.843		0.01692		0.8193548387

		26.853		0.01702		0.8301075269

		26.863		0.01709		0.8376344086

		26.873		0.01714		0.8430107527

		26.883		0.01719		0.8483870968

		26.893		0.01724		0.8537634409

		26.903		0.01736		0.8666666667

		26.913		0.01738		0.8688172043

		26.923		0.01738		0.8688172043

		26.933		0.01743		0.8741935484

		26.943		0.01748		0.8795698925

		26.953		0.01753		0.8849462366

		26.963		0.01753		0.8849462366

		26.973		0.01758		0.8903225806

		26.983		0.01758		0.8903225806

		26.993		0.01748		0.8795698925

		27.003		0.01741		0.8720430108

		27.013		0.01733		0.8634408602

		27.023		0.01721		0.8505376344

		27.033		0.01699		0.8268817204

		27.043		0.01682		0.8086021505

		27.053		0.01653		0.7774193548

		27.063		0.01624		0.7462365591

		27.073		0.01599		0.7193548387

		27.083		0.0157		0.688172043

		27.093		0.01543		0.6591397849

		27.103		0.01509		0.6225806452

		27.113		0.01487		0.5989247312

		27.123		0.01455		0.564516129

		27.133		0.01431		0.5387096774

		27.143		0.01404		0.5096774194

		27.153		0.01375		0.4784946237

		27.163		0.0135		0.4516129032

		27.173		0.01335		0.435483871

		27.184		0.01306		0.4043010753

		27.193		0.01287		0.3838709677

		27.203		0.0127		0.3655913978

		27.213		0.01252		0.3462365591

		27.223		0.01235		0.3279569892

		27.233		0.01221		0.3129032258

		27.243		0.01208		0.2989247312

		27.253		0.01189		0.2784946237

		27.263		0.01179		0.2677419355

		27.273		0.01169		0.2569892473

		27.283		0.0116		0.247311828

		27.294		0.01145		0.2311827957

		27.303		0.01138		0.223655914

		27.313		0.01128		0.2129032258

		27.323		0.01118		0.2021505376

		27.333		0.01111		0.1946236559

		27.343		0.01099		0.1817204301

		27.353		0.01096		0.1784946237

		27.363		0.01091		0.1731182796

		27.373		0.01079		0.1602150538

		27.383		0.01074		0.1548387097

		27.393		0.01067		0.147311828

		27.404		0.01062		0.1419354839

		27.413		0.0106		0.1397849462

		27.423		0.01055		0.1344086022

		27.433		0.01055		0.1344086022

		27.443		0.0105		0.1290322581

		27.453		0.0105		0.1290322581

		27.463		0.01052		0.1311827957

		27.473		0.0106		0.1397849462

		27.483		0.01069		0.1494623656

		27.493		0.01079		0.1602150538

		27.503		0.01096		0.1784946237

		27.514		0.01123		0.2075268817

		27.523		0.01138		0.223655914

		27.533		0.01169		0.2569892473

		27.543		0.01191		0.2806451613

		27.553		0.01221		0.3129032258

		27.563		0.0125		0.3440860215

		27.573		0.01277		0.3731182796

		27.583		0.01311		0.4096774194

		27.593		0.01331		0.4311827957

		27.603		0.01357		0.4591397849

		27.613		0.01389		0.4935483871

		27.624		0.01418		0.5247311828

		27.633		0.01436		0.5440860215

		27.643		0.01458		0.5677419355

		27.653		0.01479		0.5903225806

		27.663		0.01499		0.611827957

		27.673		0.01519		0.6333333333

		27.683		0.01541		0.6569892473

		27.693		0.01553		0.6698924731

		27.703		0.01567		0.6849462366

		27.713		0.01584		0.7032258065

		27.723		0.01604		0.7247311828

		27.734		0.01611		0.7322580645

		27.743		0.01621		0.7430107527

		27.753		0.01636		0.7591397849

		27.763		0.01646		0.7698924731

		27.773		0.01655		0.7795698925

		27.783		0.01663		0.788172043

		27.793		0.01675		0.8010752688

		27.803		0.01682		0.8086021505

		27.813		0.01689		0.8161290323

		27.823		0.01702		0.8301075269

		27.833		0.01709		0.8376344086

		27.844		0.01711		0.8397849462

		27.853		0.01719		0.8483870968

		27.863		0.01724		0.8537634409

		27.873		0.01729		0.8591397849

		27.883		0.01738		0.8688172043

		27.893		0.01741		0.8720430108

		27.903		0.01743		0.8741935484

		27.913		0.0175		0.8817204301

		27.923		0.01753		0.8849462366

		27.933		0.01758		0.8903225806

		27.943		0.01758		0.8903225806

		27.954		0.01763		0.8956989247

		27.963		0.01758		0.8903225806

		27.973		0.01768		0.9010752688

		27.983		0.0176		0.8924731183

		27.993		0.01753		0.8849462366

		28.003		0.01743		0.8741935484

		28.013		0.01726		0.8559139785

		28.023		0.01709		0.8376344086

		28.033		0.01687		0.8139784946

		28.043		0.01663		0.788172043

		28.053		0.01638		0.7612903226

		28.063		0.01606		0.7268817204

		28.073		0.01577		0.6956989247

		28.083		0.01553		0.6698924731

		28.093		0.01523		0.6376344086

		28.103		0.01492		0.6043010753

		28.113		0.01462		0.5720430108

		28.123		0.01433		0.5408602151

		28.133		0.01406		0.511827957

		28.143		0.01384		0.488172043

		28.153		0.01362		0.464516129

		28.163		0.01335		0.435483871

		28.173		0.01313		0.411827957

		28.183		0.01292		0.3892473118

		28.193		0.01274		0.3698924731

		28.203		0.01257		0.3516129032

		28.213		0.0124		0.3333333333

		28.223		0.0123		0.3225806452

		28.233		0.01211		0.3021505376

		28.243		0.01201		0.2913978495

		28.253		0.01184		0.2731182796

		28.263		0.01172		0.2602150538

		28.273		0.01165		0.252688172

		28.283		0.01145		0.2311827957

		28.293		0.0114		0.2258064516

		28.303		0.0113		0.2150537634

		28.313		0.01121		0.2053763441

		28.323		0.01113		0.1967741935

		28.333		0.01106		0.1892473118

		28.343		0.01094		0.176344086

		28.353		0.01091		0.1731182796

		28.363		0.01086		0.1677419355

		28.373		0.01079		0.1602150538

		28.383		0.01074		0.1548387097

		28.393		0.01064		0.1440860215

		28.403		0.01062		0.1419354839

		28.414		0.01057		0.1365591398

		28.424		0.01057		0.1365591398

		28.434		0.01052		0.1311827957

		28.444		0.01052		0.1311827957

		28.454		0.01052		0.1311827957

		28.464		0.0106		0.1397849462

		28.474		0.01067		0.147311828

		28.484		0.01082		0.1634408602

		28.494		0.01094		0.176344086

		28.504		0.01113		0.1967741935

		28.514		0.01138		0.223655914

		28.524		0.01165		0.252688172

		28.534		0.01189		0.2784946237

		28.544		0.01216		0.3075268817

		28.554		0.01248		0.3419354839

		28.564		0.01274		0.3698924731

		28.574		0.01304		0.4021505376

		28.584		0.01333		0.4333333333

		28.594		0.01362		0.464516129

		28.604		0.01382		0.4860215054

		28.614		0.01409		0.5150537634

		28.624		0.01431		0.5387096774

		28.634		0.0146		0.5698924731

		28.644		0.01482		0.5935483871

		28.654		0.01499		0.611827957

		28.664		0.01516		0.6301075269

		28.674		0.01538		0.6537634409

		28.684		0.01553		0.6698924731

		28.694		0.01565		0.6827956989

		28.704		0.01582		0.7010752688

		28.714		0.01597		0.7172043011

		28.724		0.01611		0.7322580645

		28.734		0.01626		0.7483870968

		28.744		0.01631		0.7537634409

		28.754		0.01643		0.7666666667

		28.764		0.01653		0.7774193548

		28.774		0.01665		0.7903225806

		28.784		0.01672		0.7978494624

		28.794		0.0168		0.8064516129

		28.804		0.01689		0.8161290323

		28.814		0.01697		0.8247311828

		28.824		0.01704		0.8322580645

		28.834		0.01714		0.8430107527

		28.844		0.01714		0.8430107527

		28.854		0.01724		0.8537634409

		28.864		0.01726		0.8559139785

		28.874		0.01733		0.8634408602

		28.884		0.01743		0.8741935484

		28.894		0.01746		0.8774193548

		28.904		0.01748		0.8795698925

		28.914		0.01755		0.8870967742

		28.924		0.0176		0.8924731183

		28.934		0.01758		0.8903225806

		28.944		0.01758		0.8903225806

		28.954		0.01758		0.8903225806

		28.964		0.0176		0.8924731183

		28.974		0.0176		0.8924731183

		28.984		0.01753		0.8849462366

		28.994		0.01748		0.8795698925

		29.004		0.01729		0.8591397849

		29.014		0.01716		0.8451612903

		29.024		0.01689		0.8161290323

		29.034		0.0167		0.7956989247

		29.044		0.01638		0.7612903226

		29.054		0.01614		0.735483871

		29.064		0.01582		0.7010752688

		29.074		0.01553		0.6698924731

		29.084		0.01526		0.6408602151

		29.094		0.01492		0.6043010753

		29.104		0.01465		0.5752688172

		29.114		0.01438		0.5462365591

		29.124		0.01411		0.5172043011

		29.134		0.01384		0.488172043

		29.144		0.01362		0.464516129

		29.154		0.0134		0.4408602151

		29.164		0.01318		0.4172043011

		29.174		0.01296		0.3935483871

		29.184		0.01282		0.3784946237

		29.194		0.0126		0.3548387097

		29.204		0.01245		0.3387096774

		29.214		0.0123		0.3225806452

		29.224		0.01216		0.3075268817

		29.234		0.01201		0.2913978495

		29.244		0.01189		0.2784946237

		29.254		0.01174		0.2623655914

		29.264		0.01162		0.2494623656

		29.274		0.01152		0.2387096774

		29.284		0.0114		0.2258064516

		29.294		0.01133		0.2182795699

		29.304		0.01125		0.2096774194

		29.314		0.01116		0.2

		29.324		0.01108		0.1913978495

		29.334		0.01099		0.1817204301

		29.344		0.01091		0.1731182796

		29.354		0.01086		0.1677419355

		29.364		0.01084		0.1655913978

		29.374		0.01072		0.152688172

		29.384		0.01067		0.147311828

		29.394		0.01064		0.1440860215

		29.404		0.01057		0.1365591398

		29.414		0.0106		0.1397849462

		29.424		0.01052		0.1311827957

		29.434		0.01055		0.1344086022

		29.444		0.01052		0.1311827957

		29.454		0.01057		0.1365591398

		29.464		0.01062		0.1419354839

		29.474		0.01072		0.152688172

		29.484		0.01089		0.1709677419

		29.494		0.01108		0.1913978495

		29.504		0.01123		0.2075268817

		29.514		0.01147		0.2333333333

		29.524		0.01174		0.2623655914

		29.534		0.01206		0.2967741935

		29.544		0.01235		0.3279569892

		29.554		0.01262		0.3569892473

		29.564		0.01287		0.3838709677

		29.574		0.01318		0.4172043011

		29.584		0.01345		0.4462365591

		29.594		0.0137		0.4731182796

		29.604		0.01396		0.5010752688

		29.614		0.01428		0.535483871

		29.624		0.01448		0.5569892473

		29.634		0.01467		0.5774193548

		29.644		0.01487		0.5989247312

		29.654		0.01509		0.6225806452

		29.664		0.01528		0.6430107527

		29.674		0.01541		0.6569892473

		29.684		0.0156		0.6774193548

		29.694		0.01577		0.6956989247

		29.704		0.01587		0.7064516129

		29.714		0.01606		0.7268817204

		29.724		0.01621		0.7430107527

		29.734		0.01631		0.7537634409

		29.744		0.01646		0.7698924731

		29.754		0.0165		0.7741935484

		29.764		0.0166		0.7849462366

		29.774		0.01667		0.7924731183

		29.784		0.01685		0.811827957

		29.794		0.01685		0.811827957

		29.804		0.01694		0.8215053763

		29.814		0.01699		0.8268817204

		29.824		0.01707		0.835483871

		29.834		0.01716		0.8451612903

		29.844		0.01719		0.8483870968

		29.854		0.01729		0.8591397849

		29.864		0.01733		0.8634408602

		29.874		0.01738		0.8688172043

		29.884		0.01741		0.8720430108

		29.894		0.01748		0.8795698925

		29.904		0.0175		0.8817204301

		29.914		0.01755		0.8870967742

		29.924		0.0176		0.8924731183

		29.934		0.0176		0.8924731183

		29.944		0.0176		0.8924731183

		29.954		0.01765		0.8978494624

		29.964		0.0177		0.9032258065

		29.974		0.01763		0.8956989247

		29.984		0.0175		0.8817204301

		29.994		0.01738		0.8688172043

		30.004		0.01724		0.8537634409

		30.014		0.01704		0.8322580645

		30.024		0.01682		0.8086021505

		30.034		0.01655		0.7795698925

		30.044		0.01628		0.7505376344

		30.054		0.01599		0.7193548387

		30.064		0.0157		0.688172043

		30.074		0.01545		0.6612903226

		30.084		0.01511		0.6247311828

		30.094		0.01479		0.5903225806

		30.104		0.01453		0.5623655914

		30.114		0.01426		0.5333333333

		30.124		0.01399		0.5043010753

		30.134		0.01375		0.4784946237

		30.144		0.01355		0.4569892473

		30.154		0.01328		0.4279569892

		30.164		0.01311		0.4096774194

		30.174		0.01289		0.3860215054

		30.184		0.01272		0.3677419355

		30.194		0.01252		0.3462365591

		30.204		0.01238		0.3311827957

		30.214		0.01218		0.3096774194

		30.224		0.01208		0.2989247312

		30.234		0.01191		0.2806451613

		30.244		0.01182		0.2709677419

		30.254		0.01167		0.2548387097

		30.264		0.01157		0.2440860215

		30.274		0.01145		0.2311827957

		30.284		0.01138		0.223655914

		30.294		0.01128		0.2129032258

		30.304		0.01118		0.2021505376

		30.314		0.01111		0.1946236559

		30.324		0.01106		0.1892473118

		30.334		0.01101		0.1838709677

		30.344		0.01094		0.176344086

		30.354		0.01079		0.1602150538

		30.364		0.01074		0.1548387097

		30.374		0.01072		0.152688172

		30.384		0.01064		0.1440860215

		30.394		0.01064		0.1440860215

		30.404		0.01057		0.1365591398

		30.414		0.01057		0.1365591398

		30.424		0.01052		0.1311827957

		30.434		0.0105		0.1290322581

		30.444		0.01057		0.1365591398

		30.454		0.0106		0.1397849462

		30.464		0.01069		0.1494623656

		30.474		0.01086		0.1677419355

		30.484		0.01099		0.1817204301

		30.494		0.01118		0.2021505376

		30.506		0.01147		0.2333333333

		30.516		0.01169		0.2569892473

		30.526		0.01199		0.2892473118

		30.536		0.01226		0.3182795699

		30.546		0.01255		0.3494623656

		30.556		0.01289		0.3860215054

		30.566		0.01313		0.411827957

		30.576		0.01343		0.4440860215

		30.586		0.01367		0.4698924731

		30.596		0.01399		0.5043010753

		30.606		0.01416		0.5225806452

		30.616		0.01438		0.5462365591

		30.626		0.01462		0.5720430108

		30.636		0.01482		0.5935483871

		30.646		0.01504		0.6172043011

		30.656		0.01523		0.6376344086

		30.666		0.01538		0.6537634409

		30.676		0.01555		0.6720430108

		30.686		0.01577		0.6956989247

		30.696		0.01584		0.7032258065

		30.706		0.01599		0.7193548387

		30.716		0.01616		0.7376344086

		30.726		0.01626		0.7483870968

		30.736		0.01641		0.764516129

		30.746		0.01648		0.7720430108

		30.756		0.01655		0.7795698925

		30.766		0.01665		0.7903225806

		30.776		0.01677		0.8032258065

		30.786		0.01685		0.811827957

		30.796		0.01694		0.8215053763

		30.806		0.01704		0.8322580645

		30.816		0.01707		0.835483871

		30.826		0.01711		0.8397849462

		30.836		0.01719		0.8483870968

		30.846		0.01724		0.8537634409

		30.856		0.01731		0.8612903226

		30.866		0.01736		0.8666666667

		30.876		0.01741		0.8720430108

		30.886		0.01746		0.8774193548

		30.896		0.01748		0.8795698925

		30.906		0.0175		0.8817204301

		30.916		0.01758		0.8903225806

		30.926		0.01758		0.8903225806

		30.936		0.01763		0.8956989247

		30.946		0.01768		0.9010752688

		30.956		0.01763		0.8956989247

		30.966		0.0176		0.8924731183

		30.976		0.01755		0.8870967742

		30.986		0.01743		0.8741935484

		30.996		0.01731		0.8612903226

		31.006		0.01709		0.8376344086

		31.016		0.01694		0.8215053763

		31.026		0.01667		0.7924731183

		31.036		0.01643		0.7666666667

		31.046		0.01614		0.735483871

		31.056		0.01584		0.7032258065

		31.066		0.01555		0.6720430108

		31.076		0.01523		0.6376344086

		31.086		0.01494		0.6064516129

		31.096		0.01467		0.5774193548

		31.106		0.01438		0.5462365591

		31.116		0.01411		0.5172043011

		31.126		0.01387		0.4913978495

		31.136		0.01362		0.464516129

		31.146		0.01338		0.4387096774

		31.156		0.01318		0.4172043011

		31.166		0.01299		0.3967741935

		31.176		0.01279		0.3752688172

		31.186		0.0126		0.3548387097

		31.196		0.0124		0.3333333333

		31.206		0.01228		0.3204301075

		31.216		0.01213		0.3043010753

		31.226		0.01201		0.2913978495

		31.236		0.01191		0.2806451613

		31.246		0.01174		0.2623655914

		31.256		0.01162		0.2494623656

		31.266		0.01147		0.2333333333

		31.276		0.01143		0.2290322581

		31.286		0.01133		0.2182795699

		31.296		0.01123		0.2075268817

		31.306		0.01116		0.2

		31.316		0.01106		0.1892473118

		31.326		0.01099		0.1817204301

		31.336		0.01091		0.1731182796

		31.346		0.01086		0.1677419355

		31.356		0.01079		0.1602150538

		31.366		0.01074		0.1548387097

		31.376		0.01067		0.147311828

		31.386		0.01064		0.1440860215

		31.396		0.01057		0.1365591398

		31.406		0.01052		0.1311827957

		31.416		0.0105		0.1290322581

		31.426		0.01047		0.1258064516

		31.436		0.0105		0.1290322581

		31.446		0.01055		0.1344086022

		31.456		0.01055		0.1344086022

		31.466		0.01072		0.152688172

		31.476		0.01079		0.1602150538

		31.486		0.01091		0.1731182796

		31.496		0.01116		0.2

		31.506		0.0113		0.2150537634

		31.516		0.01157		0.2440860215

		31.526		0.01184		0.2731182796

		31.536		0.01213		0.3043010753

		31.546		0.0124		0.3333333333

		31.556		0.0127		0.3655913978

		31.566		0.01296		0.3935483871

		31.576		0.01326		0.4258064516

		31.586		0.01353		0.4548387097

		31.596		0.01379		0.4827956989

		31.606		0.01409		0.5150537634

		31.616		0.01428		0.535483871

		31.626		0.0145		0.5591397849

		31.636		0.01475		0.5860215054

		31.646		0.01492		0.6043010753

		31.656		0.01514		0.6279569892

		31.666		0.01531		0.6462365591

		31.676		0.0155		0.6666666667

		31.686		0.01565		0.6827956989

		31.696		0.0158		0.6989247312

		31.706		0.01592		0.711827957

		31.716		0.01609		0.7301075269

		31.726		0.01624		0.7462365591

		31.736		0.01633		0.7559139785

		31.746		0.01641		0.764516129

		31.756		0.01653		0.7774193548

		31.766		0.01663		0.788172043

		31.776		0.0167		0.7956989247

		31.786		0.01685		0.811827957

		31.796		0.01685		0.811827957

		31.806		0.01699		0.8268817204

		31.816		0.01702		0.8301075269

		31.826		0.01709		0.8376344086

		31.836		0.01714		0.8430107527

		31.846		0.01724		0.8537634409

		31.856		0.01729		0.8591397849

		31.866		0.01731		0.8612903226

		31.876		0.01736		0.8666666667

		31.886		0.01748		0.8795698925

		31.896		0.01748		0.8795698925

		31.906		0.01748		0.8795698925

		31.916		0.01753		0.8849462366

		31.926		0.0176		0.8924731183

		31.936		0.0176		0.8924731183

		31.946		0.01765		0.8978494624

		31.956		0.01763		0.8956989247

		31.966		0.01758		0.8903225806

		31.976		0.01753		0.8849462366

		31.986		0.01743		0.8741935484

		31.996		0.01731		0.8612903226

		32.006		0.01714		0.8430107527

		32.016		0.01694		0.8215053763

		32.026		0.0167		0.7956989247

		32.036		0.01641		0.764516129

		32.046		0.01614		0.735483871

		32.055		0.01584		0.7032258065

		32.066		0.01553		0.6698924731

		32.076		0.01526		0.6408602151

		32.086		0.01494		0.6064516129

		32.096		0.0147		0.5806451613

		32.106		0.0144		0.5483870968

		32.116		0.01414		0.5204301075

		32.126		0.01392		0.4967741935

		32.136		0.01362		0.464516129

		32.145		0.0134		0.4408602151

		32.156		0.01316		0.4150537634

		32.166		0.01299		0.3967741935

		32.176		0.01274		0.3698924731

		32.186		0.01262		0.3569892473

		32.196		0.01245		0.3387096774

		32.206		0.0123		0.3225806452

		32.216		0.01218		0.3096774194

		32.226		0.01201		0.2913978495

		32.236		0.01187		0.276344086

		32.246		0.01177		0.2655913978

		32.256		0.01162		0.2494623656

		32.266		0.0115		0.2365591398

		32.276		0.01145		0.2311827957

		32.286		0.0113		0.2150537634

		32.296		0.01121		0.2053763441

		32.305		0.01113		0.1967741935

		32.316		0.01111		0.1946236559

		32.326		0.01101		0.1838709677

		32.336		0.01091		0.1731182796

		32.346		0.01089		0.1709677419

		32.356		0.01082		0.1634408602

		32.366		0.01079		0.1602150538

		32.376		0.01067		0.147311828

		32.386		0.01062		0.1419354839

		32.395		0.01055		0.1344086022

		32.406		0.01057		0.1365591398

		32.416		0.01047		0.1258064516

		32.426		0.0105		0.1290322581

		32.436		0.01047		0.1258064516

		32.446		0.01052		0.1311827957

		32.456		0.0106		0.1397849462

		32.466		0.01069		0.1494623656

		32.476		0.01079		0.1602150538

		32.486		0.01096		0.1784946237

		32.496		0.01116		0.2

		32.506		0.0114		0.2258064516

		32.516		0.0116		0.247311828

		32.526		0.01187		0.276344086

		32.536		0.01218		0.3096774194

		32.546		0.01245		0.3387096774

		32.555		0.01274		0.3698924731

		32.566		0.01304		0.4021505376

		32.576		0.01333		0.4333333333

		32.586		0.01362		0.464516129

		32.596		0.01384		0.488172043

		32.606		0.01409		0.5150537634

		32.616		0.01431		0.5387096774

		32.626		0.01455		0.564516129

		32.636		0.01479		0.5903225806

		32.645		0.01501		0.6139784946

		32.656		0.01516		0.6301075269

		32.666		0.01538		0.6537634409

		32.676		0.01553		0.6698924731

		32.686		0.01567		0.6849462366

		32.696		0.01582		0.7010752688

		32.706		0.01599		0.7193548387

		32.716		0.01611		0.7322580645

		32.726		0.01621		0.7430107527

		32.736		0.01633		0.7559139785

		32.746		0.01646		0.7698924731

		32.756		0.01655		0.7795698925

		32.766		0.01665		0.7903225806

		32.776		0.01672		0.7978494624

		32.786		0.01685		0.811827957

		32.796		0.01689		0.8161290323

		32.805		0.01699		0.8268817204

		32.816		0.01704		0.8322580645

		32.826		0.01711		0.8397849462

		32.836		0.01721		0.8505376344

		32.846		0.01724		0.8537634409

		32.856		0.01726		0.8559139785

		32.866		0.01733		0.8634408602

		32.876		0.01738		0.8688172043

		32.886		0.01746		0.8774193548

		32.895		0.01748		0.8795698925

		32.906		0.01753		0.8849462366

		32.916		0.01755		0.8870967742

		32.926		0.01758		0.8903225806

		32.936		0.01763		0.8956989247

		32.946		0.0176		0.8924731183

		32.956		0.01768		0.9010752688

		32.966		0.0176		0.8924731183

		32.976		0.01755		0.8870967742

		32.986		0.01746		0.8774193548

		32.996		0.01736		0.8666666667

		33.01		0.01709		0.8376344086

		33.016		0.01694		0.8215053763

		33.026		0.0167		0.7956989247

		33.036		0.01638		0.7612903226

		33.046		0.01614		0.735483871

		33.055		0.01582		0.7010752688

		33.066		0.01555		0.6720430108

		33.076		0.01523		0.6376344086

		33.086		0.01494		0.6064516129

		33.096		0.0147		0.5806451613

		33.106		0.01436		0.5440860215

		33.117		0.01409		0.5150537634

		33.126		0.01384		0.488172043

		33.136		0.01362		0.464516129

		33.145		0.01338		0.4387096774

		33.156		0.01321		0.4204301075

		33.166		0.01296		0.3935483871

		33.176		0.01277		0.3731182796

		33.186		0.01257		0.3516129032

		33.196		0.01248		0.3419354839

		33.206		0.01233		0.3258064516

		33.216		0.01218		0.3096774194

		33.227		0.01199		0.2892473118

		33.236		0.01184		0.2731182796

		33.246		0.01172		0.2602150538

		33.256		0.0116		0.247311828

		33.266		0.01155		0.2419354839

		33.276		0.01138		0.223655914

		33.286		0.0113		0.2150537634

		33.296		0.01123		0.2075268817

		33.305		0.01113		0.1967741935

		33.316		0.01104		0.1870967742

		33.326		0.01104		0.1870967742

		33.337		0.01096		0.1784946237

		33.346		0.01084		0.1655913978

		33.356		0.01077		0.1580645161

		33.366		0.01072		0.152688172

		33.376		0.01067		0.147311828

		33.386		0.01069		0.1494623656

		33.395		0.01057		0.1365591398

		33.406		0.01052		0.1311827957

		33.416		0.0105		0.1290322581

		33.426		0.01047		0.1258064516

		33.436		0.01047		0.1258064516

		33.447		0.0105		0.1290322581

		33.456		0.01055		0.1344086022

		33.466		0.01064		0.1440860215

		33.476		0.01072		0.152688172

		33.486		0.01089		0.1709677419

		33.496		0.01111		0.1946236559

		33.506		0.01128		0.2129032258

		33.516		0.01152		0.2387096774

		33.526		0.01182		0.2709677419

		33.536		0.01213		0.3043010753

		33.546		0.01238		0.3311827957

		33.557		0.01267		0.3623655914

		33.566		0.01292		0.3892473118

		33.576		0.01326		0.4258064516

		33.586		0.0135		0.4516129032

		33.596		0.01377		0.4806451613

		33.606		0.01401		0.5064516129

		33.616		0.01431		0.5387096774

		33.626		0.0145		0.5591397849

		33.636		0.01475		0.5860215054

		33.645		0.01494		0.6064516129

		33.656		0.01514		0.6279569892

		33.666		0.01531		0.6462365591

		33.676		0.01553		0.6698924731

		33.686		0.01572		0.6903225806

		33.696		0.01575		0.6935483871

		33.706		0.01599		0.7193548387

		33.716		0.01609		0.7301075269

		33.726		0.01619		0.7408602151

		33.736		0.01628		0.7505376344

		33.746		0.01638		0.7612903226

		33.756		0.01658		0.7827956989

		33.766		0.0166		0.7849462366

		33.777		0.0167		0.7956989247

		33.786		0.0168		0.8064516129

		33.796		0.01689		0.8161290323

		33.805		0.01697		0.8247311828

		33.816		0.01704		0.8322580645

		33.826		0.01714		0.8430107527

		33.836		0.01719		0.8483870968

		33.846		0.01724		0.8537634409

		33.856		0.01726		0.8559139785

		33.866		0.01729		0.8591397849

		33.876		0.01738		0.8688172043

		33.886		0.01741		0.8720430108

		33.895		0.01753		0.8849462366

		33.906		0.01748		0.8795698925

		33.916		0.01753		0.8849462366

		33.926		0.01758		0.8903225806

		33.936		0.01763		0.8956989247

		33.946		0.0177		0.9032258065

		33.956		0.01763		0.8956989247

		33.966		0.01763		0.8956989247

		33.976		0.01753		0.8849462366

		33.986		0.01748		0.8795698925

		33.996		0.01733		0.8634408602

		34.006		0.01719		0.8483870968

		34.016		0.01702		0.8301075269

		34.026		0.01675		0.8010752688

		34.036		0.0165		0.7741935484

		34.046		0.01621		0.7430107527

		34.055		0.01592		0.711827957

		34.066		0.01563		0.6806451613

		34.076		0.01536		0.6516129032

		34.086		0.01509		0.6225806452

		34.096		0.01475		0.5860215054

		34.106		0.01448		0.5569892473

		34.116		0.01418		0.5247311828

		34.126		0.01394		0.4989247312

		34.136		0.01367		0.4698924731

		34.145		0.01353		0.4548387097

		34.156		0.01328		0.4279569892

		34.166		0.01306		0.4043010753

		34.176		0.01284		0.3806451613

		34.186		0.01267		0.3623655914

		34.196		0.01248		0.3419354839

		34.206		0.0123		0.3225806452

		34.216		0.01223		0.3150537634

		34.226		0.01206		0.2967741935

		34.236		0.01189		0.2784946237

		34.246		0.01177		0.2655913978

		34.256		0.01165		0.252688172

		34.266		0.01157		0.2440860215

		34.276		0.01147		0.2333333333

		34.286		0.01135		0.2204301075

		34.296		0.01125		0.2096774194

		34.305		0.01116		0.2

		34.316		0.01113		0.1967741935

		34.326		0.01101		0.1838709677

		34.336		0.01094		0.176344086

		34.346		0.01086		0.1677419355

		34.356		0.01084		0.1655913978

		34.366		0.01074		0.1548387097

		34.376		0.01072		0.152688172

		34.386		0.01062		0.1419354839

		34.395		0.0106		0.1397849462

		34.406		0.01052		0.1311827957

		34.416		0.01047		0.1258064516

		34.426		0.0105		0.1290322581

		34.436		0.01045		0.123655914

		34.446		0.0105		0.1290322581

		34.456		0.01055		0.1344086022

		34.466		0.01062		0.1419354839

		34.476		0.01074		0.1548387097

		34.486		0.01086		0.1677419355

		34.496		0.01106		0.1892473118

		34.506		0.01123		0.2075268817

		34.516		0.01147		0.2333333333

		34.526		0.01172		0.2602150538

		34.536		0.01201		0.2913978495

		34.546		0.0123		0.3225806452

		34.555		0.01257		0.3516129032

		34.566		0.01287		0.3838709677

		34.576		0.01321		0.4204301075

		34.586		0.01343		0.4440860215

		34.596		0.0137		0.4731182796

		34.606		0.01401		0.5064516129

		34.616		0.01423		0.5301075269

		34.626		0.01445		0.5537634409

		34.636		0.01465		0.5752688172

		34.645		0.01487		0.5989247312

		34.656		0.01509		0.6225806452

		34.666		0.01528		0.6430107527

		34.676		0.0155		0.6666666667

		34.686		0.0156		0.6774193548

		34.696		0.01577		0.6956989247

		34.706		0.01589		0.7086021505

		34.716		0.01606		0.7268817204

		34.726		0.01609		0.7301075269

		34.736		0.01626		0.7483870968

		34.746		0.01638		0.7612903226

		34.756		0.01648		0.7720430108

		34.766		0.01663		0.788172043

		34.776		0.0167		0.7956989247

		34.786		0.01677		0.8032258065

		34.796		0.01687		0.8139784946

		34.805		0.01697		0.8247311828

		34.816		0.01699		0.8268817204

		34.826		0.01709		0.8376344086

		34.836		0.01711		0.8397849462

		34.846		0.01716		0.8451612903

		34.856		0.01729		0.8591397849

		34.866		0.01733		0.8634408602

		34.876		0.01736		0.8666666667

		34.886		0.01743		0.8741935484

		34.895		0.0175		0.8817204301

		34.906		0.01748		0.8795698925

		34.916		0.01753		0.8849462366

		34.926		0.0176		0.8924731183

		34.936		0.01763		0.8956989247

		34.946		0.0176		0.8924731183

		34.956		0.0176		0.8924731183

		34.966		0.01758		0.8903225806

		34.976		0.01755		0.8870967742

		34.986		0.01748		0.8795698925

		34.996		0.01736		0.8666666667

		35.006		0.01719		0.8483870968

		35.016		0.01704		0.8322580645

		35.026		0.01675		0.8010752688

		35.036		0.0165		0.7741935484

		35.046		0.01619		0.7408602151

		35.056		0.01592		0.711827957

		35.066		0.0156		0.6774193548

		35.076		0.01533		0.6483870968

		35.086		0.01504		0.6172043011

		35.096		0.01475		0.5860215054

		35.106		0.01445		0.5537634409

		35.116		0.01418		0.5247311828

		35.126		0.01399		0.5043010753

		35.136		0.01367		0.4698924731

		35.145		0.01345		0.4462365591

		35.156		0.01323		0.4225806452

		35.166		0.01306		0.4043010753

		35.176		0.01282		0.3784946237

		35.186		0.01265		0.3602150538

		35.196		0.01248		0.3419354839

		35.206		0.01233		0.3258064516

		35.216		0.01221		0.3129032258

		35.226		0.01201		0.2913978495

		35.236		0.01194		0.2838709677

		35.246		0.01177		0.2655913978

		35.256		0.01165		0.252688172

		35.266		0.01152		0.2387096774

		35.276		0.01147		0.2333333333

		35.286		0.01133		0.2182795699

		35.296		0.0113		0.2150537634

		35.306		0.01116		0.2

		35.316		0.01106		0.1892473118

		35.326		0.01096		0.1784946237

		35.336		0.01091		0.1731182796

		35.346		0.01084		0.1655913978

		35.356		0.01084		0.1655913978

		35.366		0.01074		0.1548387097

		35.376		0.01069		0.1494623656

		35.386		0.01064		0.1440860215

		35.395		0.0106		0.1397849462

		35.406		0.01055		0.1344086022

		35.416		0.01057		0.1365591398

		35.426		0.01052		0.1311827957

		35.436		0.01055		0.1344086022

		35.446		0.0106		0.1397849462

		35.456		0.01062		0.1419354839

		35.466		0.01072		0.152688172

		35.476		0.01089		0.1709677419

		35.486		0.01108		0.1913978495

		35.496		0.01125		0.2096774194

		35.506		0.01152		0.2387096774

		35.516		0.01174		0.2623655914

		35.526		0.01204		0.2946236559

		35.536		0.01235		0.3279569892

		35.546		0.0126		0.3548387097

		35.556		0.01289		0.3860215054

		35.566		0.01318		0.4172043011

		35.576		0.01345		0.4462365591

		35.586		0.01372		0.4752688172

		35.596		0.01396		0.5010752688

		35.606		0.01423		0.5301075269

		35.616		0.0145		0.5591397849

		35.626		0.0147		0.5806451613

		35.636		0.01492		0.6043010753

		35.645		0.01506		0.6193548387

		35.656		0.01533		0.6483870968

		35.666		0.01543		0.6591397849

		35.676		0.0156		0.6774193548

		35.686		0.01577		0.6956989247

		35.696		0.01594		0.7139784946

		35.706		0.01606		0.7268817204

		35.716		0.01616		0.7376344086

		35.726		0.01628		0.7505376344

		35.736		0.01641		0.764516129

		35.746		0.0165		0.7741935484

		35.756		0.01658		0.7827956989

		35.766		0.01667		0.7924731183

		35.776		0.01677		0.8032258065

		35.786		0.01689		0.8161290323

		35.796		0.01694		0.8215053763

		35.806		0.01702		0.8301075269

		35.816		0.01707		0.835483871

		35.826		0.01714		0.8430107527

		35.836		0.01721		0.8505376344

		35.846		0.01726		0.8559139785

		35.856		0.01731		0.8612903226

		35.866		0.01736		0.8666666667

		35.876		0.01738		0.8688172043

		35.886		0.01746		0.8774193548

		35.895		0.0175		0.8817204301

		35.906		0.01753		0.8849462366

		35.916		0.0176		0.8924731183

		35.926		0.01763		0.8956989247

		35.936		0.01768		0.9010752688

		35.946		0.0176		0.8924731183

		35.956		0.01763		0.8956989247

		35.966		0.01765		0.8978494624

		35.976		0.01753		0.8849462366

		35.986		0.01746		0.8774193548

		35.996		0.01731		0.8612903226

		36.006		0.01714		0.8430107527

		36.016		0.01689		0.8161290323

		36.026		0.01665		0.7903225806

		36.036		0.01643		0.7666666667

		36.046		0.01614		0.735483871

		36.056		0.01584		0.7032258065

		36.066		0.01555		0.6720430108

		36.076		0.01523		0.6376344086

		36.086		0.01497		0.6096774194

		36.096		0.01465		0.5752688172

		36.106		0.01443		0.5516129032

		36.116		0.01416		0.5225806452

		36.126		0.01387		0.4913978495

		36.136		0.01362		0.464516129

		36.145		0.0134		0.4408602151

		36.156		0.01323		0.4225806452

		36.166		0.01301		0.3989247312

		36.176		0.01277		0.3731182796

		36.186		0.01257		0.3516129032

		36.196		0.01243		0.3365591398

		36.206		0.01226		0.3182795699

		36.216		0.01216		0.3075268817

		36.226		0.01199		0.2892473118

		36.236		0.01184		0.2731182796

		36.246		0.01172		0.2602150538

		36.256		0.01162		0.2494623656

		36.266		0.0115		0.2365591398

		36.276		0.01145		0.2311827957

		36.286		0.0113		0.2150537634

		36.296		0.01123		0.2075268817

		36.306		0.01111		0.1946236559

		36.316		0.01104		0.1870967742

		36.326		0.01099		0.1817204301

		36.336		0.01091		0.1731182796

		36.346		0.01089		0.1709677419

		36.356		0.01084		0.1655913978

		36.366		0.01077		0.1580645161

		36.376		0.01064		0.1440860215

		36.386		0.0106		0.1397849462

		36.395		0.01055		0.1344086022

		36.406		0.01052		0.1311827957

		36.416		0.01052		0.1311827957

		36.426		0.0105		0.1290322581

		36.436		0.01052		0.1311827957

		36.446		0.0106		0.1397849462

		36.456		0.01069		0.1494623656

		36.466		0.01082		0.1634408602

		36.476		0.01099		0.1817204301

		36.486		0.01121		0.2053763441

		36.496		0.01143		0.2290322581

		36.506		0.01169		0.2569892473

		36.516		0.01196		0.2860215054

		36.526		0.01223		0.3150537634

		36.536		0.01252		0.3462365591

		36.546		0.01282		0.3784946237

		36.556		0.01309		0.4075268817

		36.566		0.01338		0.4387096774

		36.576		0.01362		0.464516129

		36.586		0.01394		0.4989247312

		36.596		0.01414		0.5204301075

		36.606		0.01438		0.5462365591

		36.616		0.01462		0.5720430108

		36.626		0.01482		0.5935483871

		36.636		0.01504		0.6172043011

		36.645		0.01526		0.6408602151

		36.656		0.01543		0.6591397849

		36.666		0.01555		0.6720430108

		36.676		0.01572		0.6903225806

		36.686		0.01584		0.7032258065

		36.696		0.01599		0.7193548387

		36.706		0.01614		0.735483871

		36.716		0.01626		0.7483870968

		36.726		0.01636		0.7591397849

		36.736		0.01648		0.7720430108

		36.746		0.01655		0.7795698925

		36.756		0.01663		0.788172043

		36.766		0.01675		0.8010752688

		36.776		0.01685		0.811827957

		36.786		0.01694		0.8215053763

		36.796		0.01697		0.8247311828

		36.806		0.01702		0.8301075269

		36.816		0.01709		0.8376344086

		36.826		0.01716		0.8451612903

		36.836		0.01724		0.8537634409

		36.846		0.01731		0.8612903226

		36.856		0.01736		0.8666666667

		36.866		0.01738		0.8688172043

		36.876		0.01741		0.8720430108

		36.886		0.01748		0.8795698925

		36.895		0.0175		0.8817204301

		36.906		0.01755		0.8870967742

		36.916		0.0176		0.8924731183

		36.926		0.01768		0.9010752688

		36.936		0.01765		0.8978494624

		36.946		0.01763		0.8956989247

		36.956		0.01768		0.9010752688

		36.966		0.01763		0.8956989247

		36.976		0.01755		0.8870967742

		36.986		0.01743		0.8741935484

		36.996		0.01731		0.8612903226

		37.006		0.01719		0.8483870968

		37.016		0.01689		0.8161290323

		37.026		0.0167		0.7956989247

		37.036		0.01641		0.764516129

		37.046		0.01611		0.7322580645

		37.056		0.0158		0.6989247312

		37.066		0.01553		0.6698924731

		37.076		0.01521		0.635483871

		37.086		0.01494		0.6064516129

		37.096		0.01465		0.5752688172

		37.106		0.01436		0.5440860215

		37.116		0.01411		0.5172043011

		37.127		0.01384		0.488172043

		37.136		0.01357		0.4591397849

		37.145		0.01338		0.4387096774

		37.156		0.01318		0.4172043011

		37.166		0.01294		0.3913978495

		37.176		0.01277		0.3731182796

		37.186		0.0126		0.3548387097

		37.196		0.01243		0.3365591398

		37.206		0.01228		0.3204301075

		37.216		0.01213		0.3043010753

		37.226		0.01199		0.2892473118

		37.237		0.01184		0.2731182796

		37.246		0.01172		0.2602150538

		37.256		0.01162		0.2494623656

		37.266		0.01152		0.2387096774

		37.276		0.01143		0.2290322581

		37.286		0.0113		0.2150537634

		37.296		0.01121		0.2053763441

		37.306		0.01113		0.1967741935

		37.316		0.01106		0.1892473118

		37.326		0.01099		0.1817204301

		37.336		0.01094		0.176344086

		37.347		0.01082		0.1634408602

		37.356		0.01079		0.1602150538

		37.366		0.01072		0.152688172

		37.376		0.01069		0.1494623656

		37.386		0.01062		0.1419354839

		37.395		0.01057		0.1365591398

		37.406		0.01055		0.1344086022

		37.416		0.0106		0.1397849462

		37.426		0.0106		0.1397849462

		37.436		0.01067		0.147311828

		37.446		0.01077		0.1580645161

		37.457		0.01094		0.176344086

		37.466		0.01108		0.1913978495

		37.476		0.0113		0.2150537634

		37.486		0.01152		0.2387096774

		37.496		0.01179		0.2677419355

		37.506		0.01211		0.3021505376

		37.516		0.01233		0.3258064516

		37.526		0.01265		0.3602150538

		37.536		0.01294		0.3913978495

		37.546		0.01318		0.4172043011

		37.556		0.0135		0.4516129032

		37.567		0.01379		0.4827956989

		37.576		0.01404		0.5096774194

		37.586		0.01426		0.5333333333

		37.596		0.01445		0.5537634409

		37.606		0.01472		0.5827956989

		37.616		0.01489		0.6010752688

		37.626		0.01514		0.6279569892

		37.636		0.01528		0.6430107527

		37.645		0.01548		0.664516129

		37.656		0.0156		0.6774193548

		37.666		0.0158		0.6989247312

		37.677		0.01587		0.7064516129

		37.686		0.01594		0.7139784946

		37.696		0.01602		0.7225806452

		37.706		0.01604		0.7247311828

		37.716		0.01602		0.7225806452

		37.726		0.01599		0.7193548387

		37.736		0.01592		0.711827957

		37.746		0.0158		0.6989247312

		37.756		0.0156		0.6774193548

		37.766		0.01538		0.6537634409

		37.776		0.01514		0.6279569892

		37.787		0.01489		0.6010752688

		37.796		0.01465		0.5752688172

		37.806		0.01445		0.5537634409

		37.816		0.01421		0.5279569892

		37.826		0.01396		0.5010752688

		37.836		0.0137		0.4731182796

		37.846		0.01345		0.4462365591

		37.856		0.01328		0.4279569892

		37.866		0.01301		0.3989247312

		37.876		0.01282		0.3784946237

		37.886		0.01265		0.3602150538

		37.897		0.01248		0.3419354839

		37.906		0.01233		0.3258064516

		37.916		0.01221		0.3129032258

		37.926		0.01218		0.3096774194

		37.936		0.01211		0.3021505376

		37.946		0.01211		0.3021505376

		37.956		0.01216		0.3075268817

		37.966		0.01226		0.3182795699

		37.976		0.01235		0.3279569892

		37.986		0.01252		0.3462365591

		37.996		0.01267		0.3623655914

		38.006		0.01289		0.3860215054

		38.016		0.01311		0.4096774194

		38.026		0.01338		0.4387096774

		38.036		0.01357		0.4591397849

		38.046		0.01382		0.4860215054

		38.056		0.01406		0.511827957

		38.066		0.01431		0.5387096774

		38.076		0.01453		0.5623655914

		38.086		0.01477		0.588172043

		38.096		0.01494		0.6064516129

		38.106		0.01514		0.6279569892

		38.116		0.01533		0.6483870968

		38.126		0.01558		0.6752688172

		38.136		0.01567		0.6849462366

		38.145		0.01582		0.7010752688

		38.156		0.01592		0.711827957

		38.166		0.01606		0.7268817204

		38.176		0.01614		0.735483871

		38.186		0.01621		0.7430107527

		38.196		0.01626		0.7483870968

		38.206		0.01631		0.7537634409

		38.216		0.01626		0.7483870968

		38.226		0.01619		0.7408602151

		38.236		0.01614		0.735483871

		38.246		0.01597		0.7172043011

		38.256		0.0158		0.6989247312

		38.266		0.0156		0.6774193548

		38.276		0.01536		0.6516129032

		38.286		0.01511		0.6247311828

		38.296		0.01484		0.5956989247

		38.306		0.01462		0.5720430108

		38.316		0.01438		0.5462365591

		38.326		0.01411		0.5172043011

		38.336		0.01384		0.488172043

		38.346		0.0136		0.4623655914

		38.356		0.01338		0.4387096774

		38.366		0.01318		0.4172043011

		38.376		0.01299		0.3967741935

		38.386		0.01279		0.3752688172

		38.396		0.0126		0.3548387097

		38.406		0.01245		0.3387096774

		38.416		0.0123		0.3225806452

		38.426		0.01223		0.3150537634

		38.436		0.01221		0.3129032258

		38.446		0.01218		0.3096774194

		38.456		0.01216		0.3075268817

		38.466		0.01226		0.3182795699

		38.476		0.01238		0.3311827957

		38.486		0.01248		0.3419354839

		38.496		0.01272		0.3677419355

		38.506		0.01289		0.3860215054

		38.516		0.01316		0.4150537634

		38.526		0.01333		0.4333333333

		38.536		0.0136		0.4623655914

		38.546		0.01384		0.488172043

		38.556		0.01409		0.5150537634

		38.566		0.01433		0.5408602151

		38.576		0.01453		0.5623655914

		38.586		0.01472		0.5827956989

		38.596		0.01494		0.6064516129

		38.606		0.01514		0.6279569892

		38.616		0.01533		0.6483870968

		38.626		0.01553		0.6698924731

		38.636		0.01567		0.6849462366

		38.646		0.01582		0.7010752688

		38.656		0.01597		0.7172043011

		38.666		0.01606		0.7268817204

		38.676		0.01621		0.7430107527

		38.686		0.01621		0.7430107527

		38.696		0.01631		0.7537634409

		38.706		0.01631		0.7537634409

		38.716		0.01628		0.7505376344

		38.726		0.01624		0.7462365591

		38.736		0.01611		0.7322580645

		38.746		0.01592		0.711827957

		38.756		0.0158		0.6989247312

		38.766		0.01563		0.6806451613

		38.776		0.01531		0.6462365591

		38.786		0.01506		0.6193548387

		38.796		0.01482		0.5935483871

		38.806		0.0146		0.5698924731

		38.816		0.01431		0.5387096774

		38.826		0.01404		0.5096774194

		38.836		0.01384		0.488172043

		38.846		0.01357		0.4591397849

		38.856		0.01335		0.435483871

		38.866		0.01318		0.4172043011

		38.876		0.01294		0.3913978495

		38.886		0.0127		0.3655913978

		38.896		0.01262		0.3569892473

		38.906		0.01245		0.3387096774

		38.916		0.0123		0.3225806452

		38.926		0.01226		0.3182795699

		38.936		0.01221		0.3129032258

		38.946		0.01221		0.3129032258

		38.956		0.01223		0.3150537634

		38.966		0.0123		0.3225806452

		38.976		0.01243		0.3365591398

		38.986		0.01265		0.3602150538

		38.996		0.01284		0.3806451613

		39.006		0.01299		0.3967741935

		39.016		0.01318		0.4172043011

		39.026		0.01343		0.4440860215

		39.036		0.01367		0.4698924731

		39.046		0.01387		0.4913978495

		39.056		0.01416		0.5225806452

		39.066		0.01436		0.5440860215

		39.076		0.01458		0.5677419355

		39.086		0.01479		0.5903225806

		39.096		0.01499		0.611827957

		39.106		0.01519		0.6333333333

		39.116		0.01543		0.6591397849

		39.126		0.01553		0.6698924731

		39.136		0.01572		0.6903225806

		39.146		0.01584		0.7032258065

		39.156		0.01597		0.7172043011

		39.166		0.01609		0.7301075269

		39.176		0.01621		0.7430107527

		39.186		0.01628		0.7505376344

		39.196		0.01628		0.7505376344

		39.206		0.01626		0.7483870968

		39.216		0.01621		0.7430107527

		39.226		0.01614		0.735483871

		39.236		0.01599		0.7193548387

		39.246		0.01582		0.7010752688

		39.256		0.01563		0.6806451613

		39.266		0.01545		0.6612903226

		39.276		0.01519		0.6333333333

		39.286		0.01492		0.6043010753

		39.296		0.01475		0.5860215054

		39.306		0.01443		0.5516129032

		39.316		0.01416		0.5225806452

		39.326		0.01389		0.4935483871

		39.336		0.0137		0.4731182796

		39.346		0.01345		0.4462365591

		39.356		0.01323		0.4225806452

		39.366		0.01304		0.4021505376

		39.376		0.01284		0.3806451613

		39.386		0.01265		0.3602150538

		39.396		0.0125		0.3440860215

		39.406		0.0124		0.3333333333

		39.416		0.0123		0.3225806452

		39.426		0.01228		0.3204301075

		39.436		0.01226		0.3182795699

		39.446		0.01226		0.3182795699

		39.456		0.0123		0.3225806452

		39.466		0.01243		0.3365591398

		39.476		0.01262		0.3569892473

		39.486		0.01277		0.3731182796

		39.496		0.01301		0.3989247312

		39.506		0.01316		0.4150537634

		39.516		0.0134		0.4408602151

		39.526		0.01365		0.4677419355

		39.536		0.01389		0.4935483871

		39.546		0.01411		0.5172043011

		39.556		0.01436		0.5440860215

		39.566		0.01455		0.564516129

		39.576		0.01479		0.5903225806

		39.586		0.01499		0.611827957

		39.596		0.01516		0.6301075269

		39.606		0.01541		0.6569892473

		39.616		0.01553		0.6698924731

		39.626		0.0157		0.688172043

		39.636		0.01584		0.7032258065

		39.646		0.01602		0.7225806452

		39.656		0.01614		0.735483871

		39.666		0.01621		0.7430107527

		39.676		0.01631		0.7537634409

		39.686		0.01636		0.7591397849

		39.696		0.01638		0.7612903226

		39.706		0.01636		0.7591397849

		39.716		0.01628		0.7505376344

		39.726		0.01624		0.7462365591

		39.736		0.01609		0.7301075269

		39.746		0.01594		0.7139784946

		39.756		0.0157		0.688172043

		39.766		0.01545		0.6612903226

		39.776		0.01519		0.6333333333

		39.786		0.01492		0.6043010753

		39.796		0.0147		0.5806451613

		39.806		0.01445		0.5537634409

		39.816		0.01416		0.5225806452

		39.826		0.01399		0.5043010753

		39.836		0.0137		0.4731182796

		39.846		0.01348		0.4494623656

		39.856		0.01328		0.4279569892

		39.866		0.01304		0.4021505376

		39.876		0.01284		0.3806451613

		39.886		0.01267		0.3623655914

		39.896		0.01255		0.3494623656

		39.906		0.0124		0.3333333333

		39.916		0.01233		0.3258064516

		39.926		0.01228		0.3204301075

		39.936		0.01228		0.3204301075

		39.946		0.01233		0.3258064516

		39.956		0.0124		0.3333333333

		39.966		0.01255		0.3494623656

		39.976		0.01267		0.3623655914

		39.986		0.01282		0.3784946237

		39.996		0.01306		0.4043010753

		40.006		0.01328		0.4279569892

		40.016		0.01355		0.4569892473

		40.026		0.01379		0.4827956989

		40.036		0.01396		0.5010752688

		40.046		0.01423		0.5301075269

		40.056		0.01445		0.5537634409

		40.066		0.01465		0.5752688172

		40.076		0.01489		0.6010752688

		40.086		0.01506		0.6193548387

		40.096		0.01528		0.6430107527

		40.106		0.01548		0.664516129

		40.116		0.01565		0.6827956989

		40.126		0.01577		0.6956989247

		40.136		0.01589		0.7086021505

		40.146		0.01602		0.7225806452

		40.156		0.01616		0.7376344086

		40.166		0.01624		0.7462365591

		40.176		0.01628		0.7505376344

		40.186		0.01631		0.7537634409

		40.196		0.01624		0.7462365591

		40.206		0.01624		0.7462365591

		40.216		0.01616		0.7376344086

		40.226		0.01606		0.7268817204

		40.236		0.01582		0.7010752688

		40.246		0.01563		0.6806451613

		40.256		0.01536		0.6516129032

		40.266		0.01514		0.6279569892

		40.276		0.01484		0.5956989247

		40.286		0.01465		0.5752688172

		40.296		0.01436		0.5440860215

		40.306		0.01411		0.5172043011

		40.316		0.01389		0.4935483871

		40.326		0.0136		0.4623655914

		40.336		0.0134		0.4408602151

		40.346		0.01321		0.4204301075

		40.356		0.01301		0.3989247312

		40.366		0.01279		0.3752688172

		40.376		0.0126		0.3548387097

		40.386		0.01248		0.3419354839

		40.396		0.01233		0.3258064516

		40.406		0.0123		0.3225806452

		40.416		0.01228		0.3204301075

		40.426		0.01218		0.3096774194

		40.436		0.01221		0.3129032258

		40.446		0.01228		0.3204301075

		40.456		0.01243		0.3365591398

		40.466		0.01252		0.3462365591

		40.476		0.01274		0.3698924731

		40.486		0.01296		0.3935483871

		40.496		0.01321		0.4204301075

		40.506		0.01338		0.4387096774

		40.52		0.01372		0.4752688172

		40.526		0.01387		0.4913978495

		40.536		0.01411		0.5172043011

		40.546		0.01433		0.5408602151

		40.556		0.01458		0.5677419355

		40.566		0.01475		0.5860215054

		40.576		0.01497		0.6096774194

		40.586		0.01516		0.6301075269

		40.596		0.01536		0.6516129032

		40.606		0.01553		0.6698924731

		40.616		0.01567		0.6849462366

		40.626		0.01582		0.7010752688

		40.636		0.01597		0.7172043011

		40.646		0.01609		0.7301075269

		40.656		0.01624		0.7462365591

		40.666		0.01624		0.7462365591

		40.676		0.01624		0.7462365591

		40.686		0.01628		0.7505376344

		40.696		0.01624		0.7462365591

		40.706		0.01614		0.735483871

		40.716		0.01599		0.7193548387

		40.726		0.01587		0.7064516129

		40.736		0.01565		0.6827956989

		40.746		0.01541		0.6569892473

		40.756		0.01521		0.635483871

		40.766		0.01497		0.6096774194

		40.776		0.01467		0.5774193548

		40.786		0.0144		0.5483870968

		40.796		0.01411		0.5172043011

		40.806		0.01392		0.4967741935

		40.816		0.01367		0.4698924731

		40.826		0.01343		0.4440860215

		40.836		0.01321		0.4204301075

		40.846		0.01304		0.4021505376

		40.856		0.01279		0.3752688172

		40.866		0.01265		0.3602150538

		40.876		0.0125		0.3440860215

		40.886		0.0124		0.3333333333

		40.896		0.0123		0.3225806452

		40.906		0.01223		0.3150537634

		40.916		0.01223		0.3150537634

		40.926		0.01228		0.3204301075

		40.936		0.01233		0.3258064516

		40.946		0.01243		0.3365591398

		40.956		0.0127		0.3655913978

		40.966		0.01282		0.3784946237

		40.976		0.01304		0.4021505376

		40.986		0.01323		0.4225806452

		40.996		0.01345		0.4462365591

		41.006		0.01367		0.4698924731

		41.016		0.01394		0.4989247312

		41.026		0.01416		0.5225806452

		41.036		0.01436		0.5440860215

		41.046		0.01462		0.5720430108

		41.056		0.01479		0.5903225806

		41.066		0.01499		0.611827957

		41.076		0.01521		0.635483871

		41.086		0.01541		0.6569892473

		41.096		0.01555		0.6720430108

		41.106		0.01572		0.6903225806

		41.116		0.01584		0.7032258065

		41.126		0.01604		0.7247311828

		41.136		0.01611		0.7322580645

		41.146		0.01619		0.7408602151

		41.156		0.01631		0.7537634409

		41.166		0.01636		0.7591397849

		41.176		0.01638		0.7612903226

		41.186		0.01633		0.7559139785

		41.196		0.01624		0.7462365591

		41.206		0.01616		0.7376344086

		41.216		0.01599		0.7193548387

		41.226		0.0158		0.6989247312

		41.236		0.01558		0.6752688172

		41.246		0.01533		0.6483870968

		41.256		0.01509		0.6225806452

		41.266		0.01482		0.5935483871

		41.276		0.01455		0.564516129

		41.286		0.01431		0.5387096774

		41.296		0.01406		0.511827957

		41.306		0.01382		0.4860215054

		41.316		0.01357		0.4591397849

		41.326		0.01335		0.435483871

		41.336		0.01321		0.4204301075

		41.346		0.01294		0.3913978495

		41.357		0.01277		0.3731182796

		41.366		0.0126		0.3548387097

		41.376		0.01248		0.3419354839

		41.386		0.01235		0.3279569892

		41.396		0.01228		0.3204301075

		41.406		0.01228		0.3204301075

		41.416		0.01226		0.3182795699

		41.426		0.0123		0.3225806452

		41.436		0.01243		0.3365591398

		41.446		0.01255		0.3494623656

		41.456		0.01274		0.3698924731

		41.467		0.01292		0.3892473118

		41.476		0.01313		0.411827957

		41.486		0.01335		0.435483871

		41.496		0.01357		0.4591397849

		41.506		0.01377		0.4806451613

		41.516		0.01404		0.5096774194

		41.526		0.01428		0.535483871

		41.536		0.01448		0.5569892473

		41.546		0.01475		0.5860215054

		41.556		0.01494		0.6064516129

		41.566		0.01516		0.6301075269

		41.577		0.01531		0.6462365591

		41.586		0.01548		0.664516129

		41.596		0.01563		0.6806451613

		41.606		0.01577		0.6956989247

		41.616		0.01589		0.7086021505

		41.626		0.01611		0.7322580645

		41.636		0.01616		0.7376344086

		41.646		0.01631		0.7537634409

		41.656		0.01636		0.7591397849

		41.666		0.01636		0.7591397849

		41.676		0.01631		0.7537634409

		41.687		0.01626		0.7483870968

		41.696		0.01614		0.735483871

		41.706		0.01606		0.7268817204

		41.716		0.01584		0.7032258065

		41.726		0.01565		0.6827956989

		41.736		0.01538		0.6537634409

		41.746		0.01514		0.6279569892

		41.756		0.01494		0.6064516129

		41.766		0.01462		0.5720430108

		41.776		0.01438		0.5462365591

		41.786		0.01414		0.5204301075

		41.797		0.01384		0.488172043

		41.806		0.01365		0.4677419355

		41.816		0.01343		0.4440860215

		41.826		0.01326		0.4258064516

		41.836		0.01299		0.3967741935

		41.846		0.01279		0.3752688172

		41.856		0.0127		0.3655913978

		41.866		0.01248		0.3419354839

		41.876		0.01235		0.3279569892

		41.886		0.01235		0.3279569892

		41.896		0.01228		0.3204301075

		41.907		0.01223		0.3150537634

		41.916		0.01223		0.3150537634

		41.926		0.01238		0.3311827957

		41.936		0.01245		0.3387096774

		41.946		0.0126		0.3548387097

		41.956		0.01279		0.3752688172

		41.966		0.01301		0.3989247312

		41.976		0.01323		0.4225806452

		41.986		0.01345		0.4462365591

		41.996		0.0137		0.4731182796

		42.006		0.01394		0.4989247312

		42.017		0.01416		0.5225806452

		42.026		0.01438		0.5462365591

		42.036		0.0146		0.5698924731

		42.046		0.01484		0.5956989247

		42.056		0.01501		0.6139784946

		42.066		0.01521		0.635483871

		42.076		0.01541		0.6569892473

		42.086		0.01558		0.6752688172

		42.096		0.01572		0.6903225806

		42.106		0.01589		0.7086021505

		42.116		0.01602		0.7225806452

		42.126		0.01619		0.7408602151

		42.136		0.01626		0.7483870968

		42.146		0.01628		0.7505376344

		42.156		0.01633		0.7559139785

		42.166		0.01631		0.7537634409

		42.176		0.01628		0.7505376344

		42.186		0.01619		0.7408602151

		42.196		0.01606		0.7268817204

		42.206		0.01589		0.7086021505

		42.216		0.01572		0.6903225806

		42.226		0.0155		0.6666666667

		42.236		0.01526		0.6408602151

		42.246		0.01501		0.6139784946

		42.256		0.01477		0.588172043

		42.266		0.01448		0.5569892473

		42.276		0.01431		0.5387096774

		42.286		0.01401		0.5064516129

		42.296		0.01372		0.4752688172

		42.306		0.0135		0.4516129032

		42.316		0.01333		0.4333333333

		42.326		0.01309		0.4075268817

		42.336		0.01287		0.3838709677

		42.346		0.0127		0.3655913978

		42.356		0.01255		0.3494623656

		42.366		0.01243		0.3365591398

		42.376		0.01235		0.3279569892

		42.386		0.01228		0.3204301075

		42.396		0.01223		0.3150537634

		42.406		0.01223		0.3150537634

		42.416		0.01233		0.3258064516

		42.426		0.01245		0.3387096774

		42.436		0.01255		0.3494623656

		42.446		0.0127		0.3655913978

		42.456		0.01289		0.3860215054

		42.466		0.01311		0.4096774194

		42.476		0.01333		0.4333333333

		42.486		0.01357		0.4591397849

		42.496		0.01382		0.4860215054

		42.506		0.01411		0.5172043011

		42.516		0.01426		0.5333333333

		42.526		0.01453		0.5623655914

		42.536		0.0147		0.5806451613

		42.546		0.01494		0.6064516129

		42.556		0.01514		0.6279569892

		42.566		0.01528		0.6430107527

		42.576		0.01553		0.6698924731

		42.586		0.01567		0.6849462366

		42.596		0.0158		0.6989247312

		42.606		0.01594		0.7139784946

		42.616		0.01609		0.7301075269

		42.626		0.01621		0.7430107527

		42.636		0.01631		0.7537634409

		42.646		0.01626		0.7483870968

		42.656		0.01631		0.7537634409

		42.666		0.01626		0.7483870968

		42.676		0.01624		0.7462365591

		42.686		0.01606		0.7268817204

		42.696		0.01589		0.7086021505

		42.706		0.01572		0.6903225806

		42.716		0.01553		0.6698924731

		42.726		0.01528		0.6430107527

		42.736		0.01501		0.6139784946

		42.746		0.01475		0.5860215054

		42.756		0.0145		0.5591397849

		42.766		0.01426		0.5333333333

		42.776		0.01406		0.511827957

		42.786		0.01375		0.4784946237

		42.796		0.0135		0.4516129032

		42.806		0.01328		0.4279569892

		42.816		0.01311		0.4096774194

		42.826		0.01287		0.3838709677

		42.836		0.0127		0.3655913978

		42.846		0.0126		0.3548387097

		42.856		0.01243		0.3365591398

		42.866		0.0123		0.3225806452

		42.876		0.01228		0.3204301075

		42.886		0.01223		0.3150537634

		42.896		0.01223		0.3150537634

		42.906		0.01228		0.3204301075

		42.916		0.0124		0.3333333333

		42.926		0.01255		0.3494623656

		42.936		0.0127		0.3655913978

		42.946		0.01289		0.3860215054

		42.956		0.01311		0.4096774194

		42.966		0.01338		0.4387096774

		42.976		0.01353		0.4548387097

		42.986		0.01382		0.4860215054

		42.996		0.01404		0.5096774194

		43.006		0.01423		0.5301075269

		43.016		0.01448		0.5569892473

		43.026		0.01477		0.588172043

		43.036		0.01492		0.6043010753

		43.046		0.01511		0.6247311828

		43.056		0.01533		0.6483870968

		43.066		0.0155		0.6666666667

		43.076		0.01567		0.6849462366

		43.086		0.01584		0.7032258065

		43.096		0.01594		0.7139784946

		43.106		0.01606		0.7268817204

		43.116		0.01619		0.7408602151

		43.126		0.01626		0.7483870968

		43.136		0.01633		0.7559139785

		43.146		0.01633		0.7559139785

		43.156		0.01636		0.7591397849

		43.166		0.01624		0.7462365591

		43.176		0.01616		0.7376344086

		43.186		0.01597		0.7172043011

		43.196		0.0158		0.6989247312

		43.206		0.01563		0.6806451613

		43.216		0.01536		0.6516129032

		43.226		0.01514		0.6279569892

		43.236		0.01487		0.5989247312

		43.246		0.01462		0.5720430108

		43.256		0.01443		0.5516129032

		43.266		0.01411		0.5172043011

		43.276		0.01384		0.488172043

		43.286		0.0136		0.4623655914

		43.296		0.0134		0.4408602151

		43.306		0.01323		0.4225806452

		43.316		0.01299		0.3967741935

		43.326		0.01279		0.3752688172

		43.336		0.01262		0.3569892473

		43.346		0.01248		0.3419354839

		43.356		0.01235		0.3279569892

		43.366		0.01228		0.3204301075

		43.376		0.01226		0.3182795699

		43.386		0.01228		0.3204301075

		43.396		0.01228		0.3204301075

		43.406		0.01243		0.3365591398

		43.416		0.01252		0.3462365591

		43.426		0.01272		0.3677419355

		43.436		0.01287		0.3838709677

		43.446		0.01311		0.4096774194

		43.456		0.01333		0.4333333333

		43.466		0.01353		0.4548387097

		43.476		0.01377		0.4806451613

		43.486		0.01404		0.5096774194

		43.496		0.01428		0.535483871

		43.506		0.01445		0.5537634409

		43.516		0.01467		0.5774193548

		43.526		0.01492		0.6043010753

		43.536		0.01514		0.6279569892

		43.546		0.01526		0.6408602151

		43.556		0.01548		0.664516129

		43.566		0.0156		0.6774193548

		43.576		0.01577		0.6956989247

		43.586		0.01594		0.7139784946

		43.596		0.01604		0.7247311828

		43.606		0.01621		0.7430107527

		43.616		0.01626		0.7483870968

		43.626		0.01628		0.7505376344

		43.636		0.01636		0.7591397849

		43.646		0.01633		0.7559139785

		43.656		0.01631		0.7537634409

		43.666		0.01624		0.7462365591

		43.676		0.01604		0.7247311828

		43.686		0.01589		0.7086021505

		43.696		0.0157		0.688172043

		43.706		0.01543		0.6591397849

		43.716		0.01521		0.635483871

		43.726		0.01499		0.611827957

		43.736		0.01472		0.5827956989

		43.746		0.01445		0.5537634409

		43.756		0.01418		0.5247311828

		43.766		0.01394		0.4989247312

		43.776		0.01375		0.4784946237

		43.786		0.01345		0.4462365591

		43.796		0.01323		0.4225806452

		43.806		0.01306		0.4043010753

		43.816		0.01292		0.3892473118

		43.826		0.0127		0.3655913978

		43.836		0.01252		0.3462365591

		43.846		0.0124		0.3333333333

		43.856		0.0123		0.3225806452

		43.866		0.01223		0.3150537634

		43.876		0.01223		0.3150537634

		43.886		0.01228		0.3204301075

		43.896		0.01235		0.3279569892

		43.906		0.01238		0.3311827957

		43.916		0.01255		0.3494623656

		43.926		0.01267		0.3623655914

		43.936		0.01287		0.3838709677

		43.946		0.01311		0.4096774194

		43.956		0.01333		0.4333333333

		43.966		0.01353		0.4548387097

		43.976		0.01382		0.4860215054

		43.986		0.01401		0.5064516129

		43.996		0.01423		0.5301075269

		44.006		0.01443		0.5516129032

		44.016		0.01467		0.5774193548

		44.026		0.01492		0.6043010753

		44.036		0.01511		0.6247311828

		44.046		0.01528		0.6430107527

		44.056		0.01545		0.6612903226

		44.066		0.01565		0.6827956989

		44.076		0.01577		0.6956989247

		44.086		0.01592		0.711827957

		44.096		0.01604		0.7247311828

		44.106		0.01616		0.7376344086

		44.116		0.01628		0.7505376344

		44.126		0.01631		0.7537634409

		44.136		0.01633		0.7559139785

		44.146		0.01628		0.7505376344

		44.156		0.01631		0.7537634409

		44.166		0.01616		0.7376344086

		44.176		0.01599		0.7193548387

		44.186		0.01587		0.7064516129

		44.196		0.01565		0.6827956989

		44.206		0.01538		0.6537634409

		44.216		0.01514		0.6279569892

		44.226		0.01494		0.6064516129

		44.236		0.01467		0.5774193548

		44.246		0.0144		0.5483870968

		44.256		0.01409		0.5150537634

		44.266		0.01387		0.4913978495

		44.276		0.01367		0.4698924731

		44.286		0.0134		0.4408602151

		44.296		0.01318		0.4172043011

		44.306		0.01301		0.3989247312

		44.316		0.01284		0.3806451613

		44.326		0.01267		0.3623655914

		44.336		0.0125		0.3440860215

		44.346		0.0124		0.3333333333

		44.356		0.01228		0.3204301075

		44.366		0.01226		0.3182795699

		44.376		0.01223		0.3150537634

		44.386		0.01226		0.3182795699

		44.396		0.0123		0.3225806452

		44.406		0.01245		0.3387096774

		44.416		0.01262		0.3569892473

		44.426		0.01284		0.3806451613

		44.436		0.01299		0.3967741935

		44.446		0.01321		0.4204301075

		44.456		0.01345		0.4462365591

		44.466		0.01367		0.4698924731

		44.476		0.01389		0.4935483871

		44.486		0.01414		0.5204301075

		44.496		0.01436		0.5440860215

		44.506		0.01462		0.5720430108

		44.516		0.01479		0.5903225806

		44.526		0.01501		0.6139784946

		44.536		0.01516		0.6301075269

		44.546		0.01538		0.6537634409

		44.556		0.01555		0.6720430108

		44.566		0.0157		0.688172043

		44.576		0.01587		0.7064516129

		44.586		0.01599		0.7193548387

		44.596		0.01611		0.7322580645

		44.606		0.01621		0.7430107527

		44.616		0.01628		0.7505376344

		44.626		0.01636		0.7591397849

		44.636		0.01636		0.7591397849

		44.646		0.01631		0.7537634409

		44.656		0.01624		0.7462365591

		44.666		0.01611		0.7322580645

		44.676		0.01597		0.7172043011

		44.686		0.01575		0.6935483871

		44.696		0.01553		0.6698924731

		44.706		0.01536		0.6516129032

		44.716		0.01506		0.6193548387

		44.726		0.01479		0.5903225806

		44.736		0.0146		0.5698924731

		44.746		0.01428		0.535483871

		44.756		0.01404		0.5096774194

		44.766		0.01377		0.4806451613

		44.776		0.01357		0.4591397849

		44.786		0.01335		0.435483871

		44.796		0.01311		0.4096774194

		44.806		0.01292		0.3892473118

		44.816		0.01274		0.3698924731

		44.826		0.01257		0.3516129032

		44.836		0.01248		0.3419354839

		44.846		0.01238		0.3311827957

		44.856		0.0123		0.3225806452

		44.866		0.01228		0.3204301075

		44.876		0.0123		0.3225806452

		44.886		0.01235		0.3279569892

		44.896		0.01245		0.3387096774

		44.906		0.01262		0.3569892473

		44.916		0.01279		0.3752688172

		44.926		0.01299		0.3967741935

		44.936		0.01318		0.4172043011

		44.946		0.01338		0.4387096774

		44.956		0.01362		0.464516129

		44.966		0.01387		0.4913978495

		44.976		0.01414		0.5204301075

		44.986		0.01431		0.5387096774

		44.996		0.01455		0.564516129

		45.006		0.01475		0.5860215054

		45.016		0.01497		0.6096774194

		45.026		0.01521		0.635483871

		45.036		0.01533		0.6483870968

		45.046		0.01553		0.6698924731

		45.056		0.0157		0.688172043

		45.066		0.01584		0.7032258065

		45.076		0.01597		0.7172043011

		45.086		0.01614		0.735483871

		45.096		0.01619		0.7408602151

		45.106		0.01628		0.7505376344

		45.116		0.01631		0.7537634409

		45.126		0.01633		0.7559139785

		45.136		0.01628		0.7505376344

		45.146		0.01624		0.7462365591

		45.156		0.01611		0.7322580645

		45.166		0.01594		0.7139784946

		45.176		0.0158		0.6989247312

		45.186		0.0155		0.6666666667

		45.196		0.01531		0.6462365591

		45.206		0.01504		0.6172043011

		45.216		0.01482		0.5935483871

		45.226		0.01453		0.5623655914

		45.236		0.01436		0.5440860215

		45.246		0.01404		0.5096774194

		45.256		0.01379		0.4827956989

		45.266		0.01355		0.4569892473

		45.276		0.01335		0.435483871

		45.286		0.01311		0.4096774194

		45.296		0.01292		0.3892473118

		45.306		0.01277		0.3731182796

		45.316		0.01262		0.3569892473

		45.326		0.01243		0.3365591398

		45.336		0.01233		0.3258064516

		45.346		0.01223		0.3150537634

		45.356		0.01221		0.3129032258

		45.366		0.01218		0.3096774194

		45.376		0.01226		0.3182795699

		45.386		0.01235		0.3279569892

		45.396		0.0125		0.3440860215

		45.406		0.01265		0.3602150538

		45.416		0.01292		0.3892473118

		45.426		0.01309		0.4075268817

		45.436		0.01328		0.4279569892

		45.446		0.0136		0.4623655914

		45.456		0.01377		0.4806451613

		45.466		0.01406		0.511827957

		45.477		0.01433		0.5408602151

		45.486		0.01445		0.5537634409

		45.496		0.01467		0.5774193548

		45.506		0.01487		0.5989247312

		45.516		0.01509		0.6225806452

		45.526		0.01533		0.6483870968

		45.536		0.01545		0.6612903226

		45.546		0.01563		0.6806451613

		45.556		0.01577		0.6956989247

		45.566		0.01589		0.7086021505

		45.576		0.01606		0.7268817204

		45.587		0.01619		0.7408602151

		45.596		0.01621		0.7430107527

		45.606		0.01626		0.7483870968

		45.616		0.01628		0.7505376344

		45.626		0.01631		0.7537634409

		45.636		0.01621		0.7430107527

		45.646		0.01609		0.7301075269

		45.656		0.01592		0.711827957

		45.666		0.01575		0.6935483871

		45.676		0.0155		0.6666666667

		45.686		0.01531		0.6462365591

		45.697		0.01506		0.6193548387

		45.706		0.01482		0.5935483871

		45.716		0.01455		0.564516129

		45.726		0.01431		0.5387096774

		45.736		0.01404		0.5096774194

		45.746		0.01379		0.4827956989

		45.756		0.01357		0.4591397849

		45.766		0.01333		0.4333333333

		45.776		0.01313		0.411827957

		45.786		0.01292		0.3892473118

		45.796		0.01274		0.3698924731

		45.807		0.01255		0.3494623656

		45.816		0.01243		0.3365591398

		45.826		0.01233		0.3258064516

		45.836		0.0123		0.3225806452

		45.846		0.01221		0.3129032258

		45.856		0.01221		0.3129032258

		45.866		0.01233		0.3258064516

		45.876		0.01233		0.3258064516

		45.886		0.01245		0.3387096774

		45.896		0.01262		0.3569892473

		45.906		0.01279		0.3752688172

		45.917		0.01299		0.3967741935

		45.926		0.01321		0.4204301075

		45.936		0.0134		0.4408602151

		45.946		0.01375		0.4784946237

		45.956		0.01394		0.4989247312

		45.966		0.01414		0.5204301075

		45.976		0.01445		0.5537634409

		45.986		0.0146		0.5698924731

		45.996		0.01479		0.5903225806

		46.006		0.01509		0.6225806452

		46.016		0.01519		0.6333333333

		46.027		0.01543		0.6591397849

		46.036		0.01563		0.6806451613

		46.046		0.01575		0.6935483871

		46.056		0.01587		0.7064516129

		46.066		0.01604		0.7247311828

		46.076		0.01611		0.7322580645

		46.086		0.01621		0.7430107527

		46.096		0.01626		0.7483870968

		46.106		0.01631		0.7537634409

		46.116		0.01633		0.7559139785

		46.126		0.01626		0.7483870968

		46.136		0.01614		0.735483871

		46.146		0.01604		0.7247311828

		46.156		0.0158		0.6989247312

		46.166		0.01563		0.6806451613

		46.176		0.01536		0.6516129032

		46.186		0.01514		0.6279569892

		46.196		0.01489		0.6010752688

		46.206		0.01462		0.5720430108

		46.216		0.01436		0.5440860215

		46.226		0.01414		0.5204301075

		46.236		0.01392		0.4967741935

		46.247		0.01365		0.4677419355

		46.256		0.0134		0.4408602151

		46.266		0.01321		0.4204301075

		46.276		0.01299		0.3967741935

		46.286		0.01279		0.3752688172

		46.296		0.01262		0.3569892473

		46.306		0.01248		0.3419354839

		46.316		0.01243		0.3365591398

		46.326		0.01226		0.3182795699

		46.336		0.01221		0.3129032258

		46.346		0.01221		0.3129032258

		46.357		0.01221		0.3129032258

		46.366		0.01233		0.3258064516

		46.376		0.0124		0.3333333333

		46.386		0.01255		0.3494623656

		46.396		0.0127		0.3655913978

		46.406		0.01289		0.3860215054

		46.416		0.01313		0.411827957

		46.426		0.01335		0.435483871

		46.436		0.0136		0.4623655914

		46.446		0.01387		0.4913978495

		46.456		0.01409		0.5150537634

		46.467		0.01438		0.5462365591

		46.476		0.01453		0.5623655914

		46.486		0.01479		0.5903225806

		46.496		0.01497		0.6096774194

		46.506		0.01514		0.6279569892

		46.516		0.01533		0.6483870968

		46.526		0.01548		0.664516129

		46.536		0.01567		0.6849462366

		46.546		0.0158		0.6989247312

		46.556		0.01594		0.7139784946

		46.566		0.01604		0.7247311828

		46.577		0.01619		0.7408602151

		46.586		0.01621		0.7430107527

		46.596		0.01624		0.7462365591

		46.606		0.01624		0.7462365591

		46.616		0.01621		0.7430107527

		46.626		0.01611		0.7322580645

		46.636		0.01602		0.7225806452

		46.646		0.01584		0.7032258065

		46.656		0.01563		0.6806451613

		46.666		0.01543		0.6591397849

		46.676		0.01521		0.635483871

		46.687		0.01489		0.6010752688

		46.696		0.0147		0.5806451613

		46.706		0.01443		0.5516129032

		46.716		0.01414		0.5204301075

		46.726		0.01392		0.4967741935

		46.736		0.01375		0.4784946237

		46.746		0.01348		0.4494623656

		46.756		0.01323		0.4225806452

		46.766		0.01299		0.3967741935

		46.776		0.01284		0.3806451613

		46.786		0.01265		0.3602150538

		46.797		0.01248		0.3419354839

		46.806		0.01238		0.3311827957

		46.816		0.01223		0.3150537634

		46.826		0.01221		0.3129032258

		46.836		0.01218		0.3096774194

		46.846		0.01218		0.3096774194

		46.856		0.01226		0.3182795699

		46.866		0.01228		0.3204301075

		46.876		0.01243		0.3365591398

		46.886		0.0126		0.3548387097

		46.896		0.01279		0.3752688172

		46.907		0.01304		0.4021505376

		46.916		0.01323		0.4225806452

		46.926		0.01345		0.4462365591

		46.936		0.01372		0.4752688172

		46.946		0.01394		0.4989247312

		46.956		0.01418		0.5247311828

		46.966		0.01443		0.5516129032

		46.976		0.01465		0.5752688172

		46.986		0.01489		0.6010752688

		46.996		0.01504		0.6172043011

		47.006		0.01528		0.6430107527

		47.017		0.01548		0.664516129

		47.026		0.0156		0.6774193548

		47.036		0.01572		0.6903225806

		47.046		0.01587		0.7064516129

		47.056		0.01602		0.7225806452

		47.066		0.01611		0.7322580645

		47.076		0.01616		0.7376344086

		47.086		0.01624		0.7462365591

		47.096		0.01628		0.7505376344

		47.106		0.01624		0.7462365591

		47.116		0.01621		0.7430107527

		47.126		0.01606		0.7268817204

		47.136		0.01589		0.7086021505

		47.146		0.01575		0.6935483871

		47.156		0.01553		0.6698924731

		47.166		0.01531		0.6462365591

		47.176		0.01501		0.6139784946

		47.186		0.01484		0.5956989247

		47.196		0.01453		0.5623655914

		47.206		0.01433		0.5408602151

		47.216		0.01401		0.5064516129

		47.226		0.01387		0.4913978495

		47.236		0.01355		0.4569892473

		47.246		0.01338		0.4387096774

		47.256		0.01321		0.4204301075

		47.266		0.01301		0.3989247312

		47.276		0.01274		0.3698924731

		47.286		0.01265		0.3602150538

		47.296		0.01255		0.3494623656

		47.306		0.0124		0.3333333333

		47.316		0.01235		0.3279569892

		47.326		0.01238		0.3311827957

		47.336		0.0124		0.3333333333

		47.346		0.01252		0.3462365591

		47.356		0.01265		0.3602150538

		47.366		0.01274		0.3698924731

		47.376		0.01294		0.3913978495

		47.386		0.01313		0.411827957

		47.396		0.01333		0.4333333333

		47.406		0.01362		0.464516129

		47.416		0.01384		0.488172043

		47.426		0.01409		0.5150537634

		47.436		0.01431		0.5387096774

		47.446		0.0145		0.5591397849

		47.456		0.01467		0.5774193548

		47.466		0.01489		0.6010752688

		47.476		0.01504		0.6172043011

		47.486		0.01521		0.635483871

		47.496		0.01533		0.6483870968

		47.506		0.01533		0.6483870968

		47.516		0.01541		0.6569892473

		47.526		0.01533		0.6483870968

		47.536		0.01523		0.6376344086

		47.546		0.01509		0.6225806452

		47.556		0.01497		0.6096774194

		47.566		0.01482		0.5935483871

		47.576		0.0146		0.5698924731

		47.586		0.01438		0.5462365591

		47.596		0.01416		0.5225806452

		47.606		0.01399		0.5043010753

		47.616		0.01375		0.4784946237

		47.626		0.01357		0.4591397849

		47.636		0.0134		0.4408602151

		47.646		0.01328		0.4279569892

		47.656		0.01316		0.4150537634

		47.666		0.01316		0.4150537634

		47.676		0.01311		0.4096774194

		47.686		0.01318		0.4172043011

		47.696		0.01323		0.4225806452

		47.706		0.01335		0.435483871

		47.716		0.01353		0.4548387097

		47.726		0.01372		0.4752688172

		47.736		0.01392		0.4967741935

		47.746		0.01406		0.511827957

		47.756		0.01426		0.5333333333

		47.766		0.01445		0.5537634409

		47.776		0.01472		0.5827956989

		47.786		0.01484		0.5956989247

		47.796		0.01499		0.611827957

		47.806		0.01516		0.6301075269

		47.816		0.01533		0.6483870968

		47.826		0.01541		0.6569892473

		47.836		0.01548		0.664516129

		47.846		0.01555		0.6720430108

		47.856		0.0155		0.6666666667

		47.866		0.01548		0.664516129

		47.876		0.01536		0.6516129032

		47.886		0.01523		0.6376344086

		47.896		0.01504		0.6172043011

		47.906		0.01489		0.6010752688

		47.916		0.01465		0.5752688172

		47.926		0.0144		0.5483870968

		47.936		0.01418		0.5247311828

		47.946		0.01399		0.5043010753

		47.956		0.01379		0.4827956989

		47.966		0.0136		0.4623655914

		47.976		0.01345		0.4462365591

		47.986		0.01335		0.435483871

		47.996		0.01328		0.4279569892

		48.006		0.01326		0.4258064516

		48.016		0.01326		0.4258064516

		48.029		0.01333		0.4333333333

		48.039		0.01343		0.4440860215

		48.049		0.01362		0.464516129

		48.06		0.01377		0.4806451613

		48.069		0.01394		0.4989247312

		48.079		0.01414		0.5204301075

		48.089		0.01436		0.5440860215

		48.099		0.0145		0.5591397849

		48.109		0.0147		0.5806451613

		48.119		0.01494		0.6064516129

		48.129		0.01506		0.6193548387

		48.139		0.01521		0.635483871

		48.149		0.01533		0.6483870968

		48.159		0.01545		0.6612903226

		48.17		0.01548		0.664516129

		48.179		0.01553		0.6698924731

		48.189		0.0155		0.6666666667

		48.199		0.01543		0.6591397849

		48.209		0.01531		0.6462365591

		48.219		0.01514		0.6279569892

		48.229		0.01501		0.6139784946

		48.239		0.01482		0.5935483871

		48.249		0.01455		0.564516129

		48.259		0.01436		0.5440860215

		48.269		0.01418		0.5247311828

		48.28		0.01389		0.4935483871

		48.289		0.01372		0.4752688172

		48.299		0.01357		0.4591397849

		48.309		0.01345		0.4462365591

		48.319		0.01333		0.4333333333

		48.329		0.01328		0.4279569892

		48.339		0.01331		0.4311827957

		48.349		0.01331		0.4311827957

		48.359		0.0134		0.4408602151

		48.369		0.01353		0.4548387097

		48.379		0.01365		0.4677419355

		48.39		0.01382		0.4860215054

		48.399		0.01396		0.5010752688

		48.409		0.01421		0.5279569892

		48.419		0.0144		0.5483870968

		48.429		0.01462		0.5720430108

		48.439		0.01477		0.588172043

		48.449		0.01494		0.6064516129

		48.459		0.01511		0.6247311828

		48.469		0.01533		0.6483870968

		48.479		0.01545		0.6612903226

		48.489		0.01548		0.664516129

		48.5		0.01558		0.6752688172

		48.509		0.01558		0.6752688172

		48.519		0.01558		0.6752688172

		48.529		0.01548		0.664516129

		48.539		0.01533		0.6483870968

		48.549		0.01519		0.6333333333

		48.559		0.01504		0.6172043011

		48.569		0.01479		0.5903225806

		48.579		0.01458		0.5677419355

		48.589		0.01436		0.5440860215

		48.599		0.01414		0.5204301075

		48.61		0.01392		0.4967741935

		48.619		0.01375		0.4784946237

		48.629		0.01362		0.464516129

		48.639		0.01345		0.4462365591

		48.649		0.0134		0.4408602151

		48.659		0.01335		0.435483871

		48.669		0.0134		0.4408602151

		48.679		0.0134		0.4408602151

		48.689		0.01348		0.4494623656

		48.699		0.01355		0.4569892473

		48.709		0.0137		0.4731182796

		48.72		0.01389		0.4935483871

		48.729		0.01404		0.5096774194

		48.739		0.01423		0.5301075269

		48.749		0.0144		0.5483870968

		48.759		0.01465		0.5752688172

		48.769		0.01479		0.5903225806

		48.779		0.01504		0.6172043011

		48.789		0.01519		0.6333333333

		48.799		0.01533		0.6483870968

		48.809		0.01541		0.6569892473

		48.819		0.0155		0.6666666667

		48.829		0.01558		0.6752688172

		48.839		0.01558		0.6752688172

		48.849		0.01558		0.6752688172

		48.859		0.0155		0.6666666667

		48.869		0.01538		0.6537634409

		48.879		0.01516		0.6301075269

		48.889		0.01499		0.611827957

		48.899		0.01484		0.5956989247

		48.909		0.0146		0.5698924731

		48.919		0.0144		0.5483870968

		48.929		0.01414		0.5204301075

		48.939		0.01392		0.4967741935

		48.949		0.01377		0.4806451613

		48.959		0.01357		0.4591397849

		48.969		0.01345		0.4462365591

		48.979		0.01338		0.4387096774

		48.989		0.01333		0.4333333333

		48.999		0.01328		0.4279569892

		49.009		0.01333		0.4333333333

		49.019		0.01348		0.4494623656

		49.029		0.01355		0.4569892473

		49.039		0.01367		0.4698924731

		49.049		0.01387		0.4913978495

		49.059		0.01404		0.5096774194

		49.069		0.01423		0.5301075269

		49.079		0.01443		0.5516129032

		49.089		0.01462		0.5720430108

		49.099		0.01482		0.5935483871

		49.109		0.01499		0.611827957

		49.119		0.01519		0.6333333333

		49.129		0.01531		0.6462365591

		49.139		0.01541		0.6569892473

		49.149		0.0155		0.6666666667

		49.159		0.01563		0.6806451613

		49.169		0.01558		0.6752688172

		49.179		0.0155		0.6666666667

		49.189		0.0155		0.6666666667

		49.199		0.01533		0.6483870968

		49.209		0.01526		0.6408602151

		49.219		0.01501		0.6139784946

		49.229		0.01482		0.5935483871

		49.239		0.0146		0.5698924731

		49.249		0.0144		0.5483870968

		49.259		0.01418		0.5247311828

		49.269		0.01399		0.5043010753

		49.279		0.01379		0.4827956989

		49.289		0.01362		0.464516129

		49.299		0.0135		0.4516129032

		49.309		0.0134		0.4408602151

		49.319		0.01333		0.4333333333

		49.329		0.01335		0.435483871

		49.339		0.01338		0.4387096774

		49.349		0.01348		0.4494623656

		49.359		0.01357		0.4591397849

		49.369		0.01377		0.4806451613

		49.379		0.01392		0.4967741935

		49.389		0.01406		0.511827957

		49.399		0.01426		0.5333333333

		49.409		0.01443		0.5516129032

		49.419		0.01462		0.5720430108

		49.429		0.01487		0.5989247312

		49.439		0.01501		0.6139784946

		49.449		0.01519		0.6333333333

		49.459		0.01531		0.6462365591

		49.469		0.01543		0.6591397849

		49.479		0.01558		0.6752688172

		49.489		0.0156		0.6774193548

		49.499		0.0156		0.6774193548

		49.509		0.01558		0.6752688172

		49.519		0.01548		0.664516129

		49.529		0.01536		0.6516129032

		49.539		0.01519		0.6333333333

		49.549		0.01504		0.6172043011

		49.559		0.01482		0.5935483871

		49.569		0.0146		0.5698924731

		49.579		0.01438		0.5462365591

		49.589		0.01418		0.5247311828

		49.599		0.01396		0.5010752688

		49.609		0.01377		0.4806451613

		49.619		0.01362		0.464516129

		49.629		0.0135		0.4516129032

		49.639		0.0134		0.4408602151

		49.649		0.01338		0.4387096774

		49.659		0.01343		0.4440860215

		49.669		0.01343		0.4440860215

		49.679		0.0135		0.4516129032

		49.689		0.01357		0.4591397849

		49.699		0.01377		0.4806451613

		49.709		0.01394		0.4989247312

		49.719		0.01411		0.5172043011

		49.729		0.01431		0.5387096774

		49.739		0.01448		0.5569892473

		49.749		0.01467		0.5774193548

		49.759		0.01492		0.6043010753

		49.769		0.01514		0.6279569892

		49.779		0.01521		0.635483871

		49.789		0.01538		0.6537634409

		49.799		0.01548		0.664516129

		49.809		0.01565		0.6827956989

		49.819		0.01567		0.6849462366

		49.829		0.01567		0.6849462366

		49.839		0.0157		0.688172043

		49.849		0.01558		0.6752688172

		49.859		0.01545		0.6612903226

		49.869		0.01531		0.6462365591

		49.879		0.01519		0.6333333333

		49.889		0.01494		0.6064516129

		49.899		0.01472		0.5827956989

		49.909		0.0145		0.5591397849

		49.919		0.01426		0.5333333333

		49.929		0.01404		0.5096774194

		49.939		0.01384		0.488172043

		49.949		0.01367		0.4698924731

		49.959		0.0136		0.4623655914

		49.969		0.0135		0.4516129032

		49.979		0.0135		0.4516129032

		49.989		0.01353		0.4548387097

		49.999		0.01355		0.4569892473

		50.009		0.01367		0.4698924731

		50.019		0.01377		0.4806451613

		50.029		0.01396		0.5010752688

		50.039		0.01406		0.511827957

		50.049		0.01428		0.535483871

		50.059		0.01445		0.5537634409

		50.069		0.01482		0.5935483871

		50.079		0.01484		0.5956989247

		50.089		0.01504		0.6172043011

		50.099		0.01523		0.6376344086

		50.109		0.01538		0.6537634409

		50.119		0.01548		0.664516129

		50.129		0.0157		0.688172043

		50.139		0.01575		0.6935483871

		50.149		0.0158		0.6989247312

		50.159		0.01582		0.7010752688

		50.169		0.01577		0.6956989247

		50.179		0.01567		0.6849462366

		50.189		0.01558		0.6752688172

		50.199		0.01543		0.6591397849

		50.209		0.01521		0.635483871

		50.219		0.01499		0.611827957

		50.229		0.01482		0.5935483871

		50.239		0.01453		0.5623655914

		50.249		0.01431		0.5387096774

		50.259		0.01418		0.5247311828

		50.269		0.01392		0.4967741935

		50.279		0.01377		0.4806451613

		50.289		0.01362		0.464516129

		50.299		0.0136		0.4623655914

		50.309		0.0135		0.4516129032

		50.319		0.01345		0.4462365591

		50.329		0.01355		0.4569892473

		50.339		0.01372		0.4752688172

		50.349		0.01375		0.4784946237

		50.359		0.01389		0.4935483871

		50.369		0.01404		0.5096774194

		50.379		0.01423		0.5301075269

		50.389		0.01445		0.5537634409

		50.399		0.01462		0.5720430108

		50.409		0.01477		0.588172043

		50.419		0.01497		0.6096774194

		50.429		0.01516		0.6301075269

		50.439		0.01536		0.6516129032

		50.449		0.01545		0.6612903226

		50.459		0.01555		0.6720430108

		50.469		0.0156		0.6774193548

		50.479		0.01567		0.6849462366

		50.489		0.0157		0.688172043

		50.499		0.01563		0.6806451613

		50.509		0.01555		0.6720430108

		50.519		0.01538		0.6537634409

		50.533		0.01514		0.6279569892

		50.539		0.01504		0.6172043011

		50.549		0.01484		0.5956989247

		50.559		0.01462		0.5720430108

		50.569		0.01438		0.5462365591

		50.579		0.01421		0.5279569892

		50.589		0.01401		0.5064516129

		50.599		0.01377		0.4806451613

		50.609		0.0137		0.4731182796

		50.619		0.01353		0.4548387097

		50.629		0.01348		0.4494623656

		50.639		0.01348		0.4494623656

		50.649		0.01362		0.464516129

		50.659		0.0136		0.4623655914

		50.669		0.01372		0.4752688172

		50.679		0.01392		0.4967741935

		50.689		0.01401		0.5064516129

		50.699		0.01423		0.5301075269

		50.709		0.0144		0.5483870968

		50.719		0.0146		0.5698924731

		50.729		0.01477		0.588172043

		50.739		0.01494		0.6064516129

		50.749		0.01516		0.6301075269

		50.759		0.01533		0.6483870968

		50.769		0.01543		0.6591397849

		50.779		0.0156		0.6774193548

		50.789		0.01567		0.6849462366

		50.799		0.01572		0.6903225806

		50.809		0.01577		0.6956989247

		50.819		0.01575		0.6935483871

		50.829		0.01567		0.6849462366

		50.839		0.01555		0.6720430108

		50.849		0.01543		0.6591397849

		50.859		0.01521		0.635483871

		50.869		0.01504		0.6172043011

		50.879		0.01484		0.5956989247

		50.889		0.01458		0.5677419355

		50.899		0.01443		0.5516129032

		50.909		0.01414		0.5204301075

		50.919		0.01396		0.5010752688

		50.929		0.01379		0.4827956989

		50.939		0.01365		0.4677419355

		50.949		0.01353		0.4548387097

		50.959		0.01345		0.4462365591

		50.969		0.01348		0.4494623656

		50.979		0.01348		0.4494623656

		50.989		0.01353		0.4548387097

		50.999		0.01362		0.464516129

		51.009		0.01377		0.4806451613

		51.019		0.01392		0.4967741935

		51.029		0.01411		0.5172043011

		51.039		0.01428		0.535483871

		51.049		0.01445		0.5537634409

		51.059		0.01467		0.5774193548

		51.069		0.01484		0.5956989247

		51.079		0.01504		0.6172043011

		51.089		0.01519		0.6333333333

		51.099		0.01536		0.6516129032

		51.109		0.01553		0.6698924731

		51.119		0.0156		0.6774193548

		51.129		0.01565		0.6827956989

		51.139		0.01572		0.6903225806

		51.149		0.01565		0.6827956989

		51.159		0.01565		0.6827956989

		51.169		0.01545		0.6612903226

		51.179		0.01538		0.6537634409

		51.189		0.01521		0.635483871

		51.199		0.01497		0.6096774194

		51.209		0.01475		0.5860215054

		51.219		0.0145		0.5591397849

		51.229		0.01426		0.5333333333

		51.239		0.01414		0.5204301075

		51.249		0.01387		0.4913978495

		51.259		0.0137		0.4731182796

		51.269		0.01357		0.4591397849

		51.279		0.01348		0.4494623656

		51.289		0.01343		0.4440860215

		51.299		0.0134		0.4408602151

		51.309		0.01345		0.4462365591

		51.319		0.01353		0.4548387097

		51.329		0.01362		0.464516129

		51.339		0.01375		0.4784946237

		51.349		0.01396		0.5010752688

		51.359		0.01411		0.5172043011

		51.369		0.01436		0.5440860215

		51.379		0.0145		0.5591397849

		51.389		0.01472		0.5827956989

		51.399		0.01489		0.6010752688

		51.409		0.01509		0.6225806452

		51.419		0.01521		0.635483871

		51.429		0.01533		0.6483870968

		51.439		0.01545		0.6612903226

		51.449		0.0156		0.6774193548

		51.459		0.0156		0.6774193548

		51.469		0.01563		0.6806451613

		51.479		0.01563		0.6806451613

		51.489		0.01553		0.6698924731

		51.499		0.01541		0.6569892473

		51.509		0.01528		0.6430107527

		51.519		0.01501		0.6139784946

		51.529		0.01484		0.5956989247

		51.539		0.01462		0.5720430108

		51.549		0.0144		0.5483870968

		51.56		0.01411		0.5172043011

		51.569		0.01396		0.5010752688

		51.579		0.01377		0.4806451613

		51.589		0.0136		0.4623655914

		51.599		0.01345		0.4462365591

		51.609		0.01335		0.435483871

		51.619		0.01333		0.4333333333

		51.629		0.01333		0.4333333333

		51.639		0.01338		0.4387096774

		51.649		0.01345		0.4462365591

		51.659		0.01357		0.4591397849

		51.669		0.0137		0.4731182796

		51.679		0.01389		0.4935483871

		51.689		0.01411		0.5172043011

		51.699		0.01423		0.5301075269

		51.709		0.01445		0.5537634409

		51.719		0.01462		0.5720430108

		51.729		0.01482		0.5935483871

		51.739		0.01504		0.6172043011

		51.749		0.01519		0.6333333333

		51.759		0.01533		0.6483870968

		51.769		0.01538		0.6537634409

		51.779		0.01545		0.6612903226

		51.789		0.01553		0.6698924731

		51.799		0.01563		0.6806451613

		51.809		0.0155		0.6666666667

		51.819		0.01541		0.6569892473

		51.829		0.01531		0.6462365591

		51.839		0.01516		0.6301075269

		51.849		0.01499		0.611827957

		51.859		0.01479		0.5903225806

		51.869		0.0145		0.5591397849

		51.879		0.01428		0.535483871

		51.889		0.01409		0.5150537634

		51.899		0.01392		0.4967741935

		51.909		0.0137		0.4731182796

		51.919		0.01353		0.4548387097

		51.929		0.01343		0.4440860215

		51.939		0.01338		0.4387096774

		51.949		0.01328		0.4279569892

		51.959		0.01333		0.4333333333

		51.969		0.01338		0.4387096774

		51.979		0.01348		0.4494623656

		51.989		0.01362		0.464516129

		51.999		0.01375		0.4784946237

		52.009		0.01392		0.4967741935

		52.019		0.01409		0.5150537634

		52.029		0.01426		0.5333333333

		52.039		0.01445		0.5537634409

		52.049		0.01467		0.5774193548

		52.059		0.01487		0.5989247312

		52.07		0.01504		0.6172043011

		52.079		0.01519		0.6333333333

		52.089		0.01533		0.6483870968

		52.099		0.01543		0.6591397849

		52.11		0.01553		0.6698924731

		52.119		0.01553		0.6698924731

		52.129		0.01555		0.6720430108

		52.139		0.01548		0.664516129

		52.149		0.01538		0.6537634409

		52.159		0.01523		0.6376344086

		52.169		0.01506		0.6193548387

		52.182		0.01482		0.5935483871

		52.189		0.0147		0.5806451613

		52.199		0.01445		0.5537634409

		52.209		0.01423		0.5301075269

		52.219		0.01401		0.5064516129

		52.229		0.01382		0.4860215054

		52.239		0.0137		0.4731182796

		52.249		0.01348		0.4494623656

		52.259		0.01338		0.4387096774

		52.269		0.01335		0.435483871

		52.279		0.01326		0.4258064516

		52.29		0.01335		0.435483871

		52.299		0.01338		0.4387096774

		52.309		0.01343		0.4440860215

		52.319		0.01357		0.4591397849

		52.329		0.0137		0.4731182796

		52.339		0.01389		0.4935483871

		52.349		0.01409		0.5150537634

		52.359		0.01428		0.535483871

		52.369		0.01445		0.5537634409

		52.379		0.01467		0.5774193548

		52.389		0.01484		0.5956989247

		52.401		0.01506		0.6193548387

		52.409		0.01514		0.6279569892

		52.419		0.01538		0.6537634409

		52.429		0.01548		0.664516129

		52.439		0.01553		0.6698924731

		52.449		0.0155		0.6666666667

		52.459		0.01553		0.6698924731

		52.469		0.01553		0.6698924731

		52.479		0.01538		0.6537634409

		52.489		0.01528		0.6430107527

		52.499		0.01506		0.6193548387

		52.51		0.01487		0.5989247312

		52.519		0.01465		0.5752688172

		52.529		0.01445		0.5537634409

		52.539		0.01421		0.5279569892

		52.549		0.01401		0.5064516129

		52.559		0.01384		0.488172043

		52.569		0.01365		0.4677419355

		52.579		0.01348		0.4494623656

		52.589		0.01343		0.4440860215

		52.599		0.01338		0.4387096774

		52.609		0.01335		0.435483871

		52.62		0.01338		0.4387096774

		52.629		0.01343		0.4440860215

		52.639		0.01353		0.4548387097

		52.649		0.01365		0.4677419355

		52.659		0.01387		0.4913978495

		52.669		0.01401		0.5064516129

		52.679		0.01418		0.5247311828

		52.689		0.0144		0.5483870968

		52.699		0.01462		0.5720430108

		52.709		0.01475		0.5860215054

		52.719		0.01494		0.6064516129

		52.73		0.01514		0.6279569892

		52.739		0.01528		0.6430107527

		52.749		0.01541		0.6569892473

		52.759		0.01553		0.6698924731

		52.769		0.01558		0.6752688172

		52.779		0.01563		0.6806451613

		52.789		0.01558		0.6752688172

		52.799		0.01548		0.664516129

		52.809		0.01541		0.6569892473

		52.819		0.01523		0.6376344086

		52.829		0.01506		0.6193548387

		52.84		0.01484		0.5956989247

		52.849		0.01465		0.5752688172

		52.859		0.01445		0.5537634409

		52.869		0.01421		0.5279569892

		52.879		0.01401		0.5064516129

		52.889		0.01379		0.4827956989

		52.899		0.01365		0.4677419355

		52.909		0.01353		0.4548387097

		52.919		0.01343		0.4440860215

		52.929		0.01333		0.4333333333

		52.939		0.01338		0.4387096774

		52.949		0.0134		0.4408602151

		52.959		0.0134		0.4408602151

		52.969		0.0135		0.4516129032

		52.979		0.01367		0.4698924731

		52.989		0.01382		0.4860215054

		52.999		0.01399		0.5043010753

		53.009		0.01418		0.5247311828

		53.019		0.01436		0.5440860215

		53.029		0.01455		0.564516129

		53.039		0.01475		0.5860215054

		53.049		0.01494		0.6064516129

		53.059		0.01506		0.6193548387

		53.069		0.01523		0.6376344086

		53.079		0.01541		0.6569892473

		53.089		0.01548		0.664516129

		53.099		0.01555		0.6720430108

		53.109		0.01555		0.6720430108

		53.119		0.01553		0.6698924731

		53.129		0.01543		0.6591397849

		53.139		0.01536		0.6516129032

		53.149		0.01521		0.635483871

		53.159		0.01504		0.6172043011

		53.169		0.01482		0.5935483871

		53.179		0.0146		0.5698924731

		53.189		0.01445		0.5537634409

		53.199		0.01416		0.5225806452

		53.209		0.01392		0.4967741935

		53.219		0.01375		0.4784946237

		53.229		0.01357		0.4591397849

		53.239		0.01348		0.4494623656

		53.249		0.01338		0.4387096774

		53.259		0.01338		0.4387096774

		53.269		0.01331		0.4311827957

		53.279		0.01338		0.4387096774

		53.289		0.01343		0.4440860215

		53.299		0.01357		0.4591397849

		53.309		0.0137		0.4731182796

		53.319		0.01384		0.488172043

		53.329		0.01406		0.511827957

		53.339		0.01426		0.5333333333

		53.349		0.0144		0.5483870968

		53.359		0.01462		0.5720430108

		53.369		0.01479		0.5903225806

		53.379		0.01499		0.611827957

		53.389		0.01516		0.6301075269

		53.399		0.01533		0.6483870968

		53.409		0.01543		0.6591397849

		53.419		0.0155		0.6666666667

		53.429		0.01558		0.6752688172

		53.439		0.01553		0.6698924731

		53.449		0.0156		0.6774193548

		53.459		0.01545		0.6612903226

		53.469		0.01536		0.6516129032

		53.479		0.01521		0.635483871

		53.489		0.01499		0.611827957

		53.499		0.01487		0.5989247312

		53.509		0.01458		0.5677419355

		53.519		0.01436		0.5440860215

		53.529		0.01418		0.5247311828

		53.539		0.01396		0.5010752688

		53.549		0.01375		0.4784946237

		53.559		0.01357		0.4591397849

		53.569		0.01348		0.4494623656

		53.579		0.0134		0.4408602151

		53.589		0.01338		0.4387096774

		53.599		0.01335		0.435483871

		53.609		0.01338		0.4387096774

		53.619		0.01348		0.4494623656

		53.629		0.01357		0.4591397849

		53.639		0.01375		0.4784946237

		53.649		0.01392		0.4967741935

		53.659		0.01409		0.5150537634

		53.669		0.01431		0.5387096774

		53.679		0.01445		0.5537634409

		53.689		0.01465		0.5752688172

		53.699		0.01489		0.6010752688

		53.709		0.01501		0.6139784946

		53.719		0.01516		0.6301075269

		53.729		0.01528		0.6430107527

		53.739		0.01543		0.6591397849

		53.749		0.01548		0.664516129

		53.759		0.01558		0.6752688172

		53.769		0.01553		0.6698924731

		53.779		0.0155		0.6666666667

		53.789		0.01543		0.6591397849

		53.799		0.01531		0.6462365591

		53.809		0.01514		0.6279569892

		53.819		0.01499		0.611827957

		53.829		0.01477		0.588172043

		53.839		0.01455		0.564516129

		53.849		0.01433		0.5408602151

		53.859		0.01414		0.5204301075

		53.869		0.01392		0.4967741935

		53.879		0.01372		0.4752688172

		53.889		0.01357		0.4591397849

		53.899		0.01348		0.4494623656

		53.909		0.01331		0.4311827957

		53.919		0.01328		0.4279569892

		53.929		0.01331		0.4311827957

		53.939		0.01338		0.4387096774

		53.949		0.01343		0.4440860215

		53.959		0.01357		0.4591397849

		53.969		0.01372		0.4752688172

		53.979		0.01384		0.488172043

		53.989		0.01406		0.511827957

		53.999		0.01423		0.5301075269

		54.009		0.01448		0.5569892473

		54.019		0.01462		0.5720430108

		54.029		0.01479		0.5903225806

		54.039		0.01501		0.6139784946

		54.049		0.01516		0.6301075269

		54.059		0.01528		0.6430107527

		54.069		0.01543		0.6591397849

		54.079		0.01553		0.6698924731

		54.089		0.01558		0.6752688172

		54.099		0.01555		0.6720430108

		54.109		0.01553		0.6698924731

		54.119		0.01545		0.6612903226

		54.129		0.01531		0.6462365591

		54.139		0.01516		0.6301075269

		54.149		0.01499		0.611827957

		54.159		0.01475		0.5860215054

		54.169		0.01458		0.5677419355

		54.179		0.01433		0.5408602151

		54.189		0.01416		0.5225806452

		54.199		0.01396		0.5010752688

		54.209		0.01375		0.4784946237

		54.219		0.01357		0.4591397849

		54.229		0.01345		0.4462365591

		54.239		0.01335		0.435483871

		54.249		0.01331		0.4311827957

		54.259		0.0134		0.4408602151

		54.269		0.01343		0.4440860215

		54.279		0.01348		0.4494623656

		54.289		0.01365		0.4677419355

		54.299		0.01375		0.4784946237

		54.309		0.01392		0.4967741935

		54.319		0.01406		0.511827957

		54.329		0.01426		0.5333333333

		54.339		0.01448		0.5569892473

		54.349		0.01465		0.5752688172

		54.359		0.01487		0.5989247312

		54.369		0.01504		0.6172043011

		54.379		0.01519		0.6333333333

		54.389		0.01533		0.6483870968

		54.399		0.01543		0.6591397849

		54.409		0.01553		0.6698924731

		54.419		0.0156		0.6774193548

		54.429		0.01555		0.6720430108

		54.439		0.0155		0.6666666667

		54.449		0.01545		0.6612903226

		54.459		0.01538		0.6537634409

		54.469		0.01523		0.6376344086

		54.479		0.01497		0.6096774194

		54.489		0.01479		0.5903225806

		54.499		0.01455		0.564516129

		54.509		0.01436		0.5440860215

		54.519		0.01411		0.5172043011

		54.529		0.01392		0.4967741935

		54.539		0.01367		0.4698924731

		54.549		0.01362		0.464516129

		54.559		0.01345		0.4462365591

		54.569		0.01333		0.4333333333

		54.579		0.01331		0.4311827957

		54.589		0.01331		0.4311827957

		54.599		0.01333		0.4333333333

		54.609		0.01343		0.4440860215

		54.619		0.01357		0.4591397849

		54.629		0.0137		0.4731182796

		54.639		0.01389		0.4935483871

		54.649		0.01411		0.5172043011

		54.659		0.01426		0.5333333333

		54.669		0.01445		0.5537634409

		54.679		0.01467		0.5774193548

		54.689		0.01482		0.5935483871

		54.699		0.01509		0.6225806452

		54.709		0.01514		0.6279569892

		54.719		0.01536		0.6516129032

		54.729		0.01543		0.6591397849

		54.739		0.0155		0.6666666667

		54.749		0.0155		0.6666666667

		54.759		0.01553		0.6698924731

		54.769		0.01548		0.664516129

		54.779		0.01538		0.6537634409

		54.789		0.01528		0.6430107527

		54.799		0.01506		0.6193548387

		54.809		0.01492		0.6043010753

		54.819		0.01472		0.5827956989

		54.829		0.01448		0.5569892473

		54.839		0.01431		0.5387096774

		54.849		0.01404		0.5096774194

		54.859		0.01384		0.488172043

		54.869		0.01365		0.4677419355

		54.879		0.0135		0.4516129032

		54.889		0.01338		0.4387096774

		54.899		0.01335		0.435483871

		54.909		0.01323		0.4225806452

		54.919		0.01321		0.4204301075

		54.929		0.01331		0.4311827957

		54.939		0.01343		0.4440860215

		54.949		0.0135		0.4516129032

		54.959		0.01365		0.4677419355

		54.969		0.01382		0.4860215054

		54.979		0.01401		0.5064516129

		54.989		0.01423		0.5301075269

		54.999		0.01438		0.5462365591

		55.009		0.01458		0.5677419355

		55.019		0.01475		0.5860215054

		55.029		0.01492		0.6043010753

		55.039		0.01514		0.6279569892

		55.049		0.01523		0.6376344086

		55.059		0.01531		0.6462365591

		55.069		0.01545		0.6612903226

		55.079		0.01548		0.664516129

		55.089		0.01541		0.6569892473

		55.099		0.01536		0.6516129032

		55.109		0.01531		0.6462365591

		55.119		0.01519		0.6333333333

		55.129		0.01504		0.6172043011

		55.139		0.01482		0.5935483871

		55.149		0.01465		0.5752688172

		55.159		0.0144		0.5483870968

		55.169		0.01421		0.5279569892

		55.179		0.01396		0.5010752688

		55.189		0.01377		0.4806451613

		55.199		0.01362		0.464516129

		55.209		0.0134		0.4408602151

		55.219		0.01328		0.4279569892

		55.229		0.01321		0.4204301075

		55.239		0.01321		0.4204301075

		55.249		0.01318		0.4172043011

		55.259		0.01323		0.4225806452

		55.269		0.01331		0.4311827957

		55.279		0.0134		0.4408602151

		55.289		0.01357		0.4591397849

		55.299		0.01375		0.4784946237

		55.309		0.01389		0.4935483871

		55.319		0.01414		0.5204301075

		55.329		0.01433		0.5408602151

		55.339		0.01453		0.5623655914

		55.349		0.01467		0.5774193548

		55.359		0.01487		0.5989247312

		55.369		0.01504		0.6172043011

		55.379		0.01523		0.6376344086

		55.389		0.01526		0.6408602151

		55.399		0.01536		0.6516129032

		55.409		0.01541		0.6569892473

		55.419		0.01538		0.6537634409

		55.429		0.01541		0.6569892473

		55.439		0.01531		0.6462365591

		55.449		0.01514		0.6279569892

		55.459		0.01501		0.6139784946

		55.469		0.01477		0.588172043

		55.479		0.01455		0.564516129

		55.489		0.01438		0.5462365591

		55.499		0.01421		0.5279569892

		55.509		0.01396		0.5010752688

		55.519		0.01372		0.4752688172

		55.529		0.01355		0.4569892473

		55.539		0.0134		0.4408602151

		55.549		0.01331		0.4311827957

		55.559		0.01321		0.4204301075

		55.569		0.01316		0.4150537634

		55.579		0.01321		0.4204301075

		55.589		0.01321		0.4204301075

		55.599		0.01331		0.4311827957

		55.609		0.01345		0.4462365591

		55.619		0.0136		0.4623655914

		55.629		0.01377		0.4806451613

		55.639		0.01394		0.4989247312

		55.649		0.01416		0.5225806452

		55.659		0.01433		0.5408602151

		55.669		0.01453		0.5623655914

		55.679		0.01477		0.588172043

		55.689		0.01494		0.6064516129

		55.699		0.01504		0.6172043011

		55.709		0.01516		0.6301075269

		55.719		0.01528		0.6430107527

		55.729		0.01538		0.6537634409

		55.739		0.01536		0.6516129032

		55.749		0.01536		0.6516129032

		55.759		0.01531		0.6462365591

		55.769		0.01519		0.6333333333

		55.779		0.01509		0.6225806452

		55.789		0.01489		0.6010752688

		55.799		0.01475		0.5860215054

		55.809		0.01448		0.5569892473

		55.819		0.01428		0.535483871

		55.829		0.01404		0.5096774194

		55.839		0.01384		0.488172043

		55.849		0.01367		0.4698924731

		55.859		0.01345		0.4462365591

		55.869		0.0134		0.4408602151

		55.879		0.01326		0.4258064516

		55.889		0.01316		0.4150537634

		55.899		0.01309		0.4075268817

		55.909		0.01318		0.4172043011

		55.919		0.01318		0.4172043011

		55.929		0.01328		0.4279569892

		55.939		0.01343		0.4440860215

		55.949		0.01357		0.4591397849

		55.959		0.01375		0.4784946237

		55.969		0.01396		0.5010752688

		55.979		0.01416		0.5225806452

		55.989		0.01436		0.5440860215

		55.999		0.01455		0.564516129

		56.009		0.01472		0.5827956989

		56.019		0.01489		0.6010752688

		56.029		0.01504		0.6172043011

		56.039		0.01521		0.635483871

		56.049		0.01531		0.6462365591

		56.059		0.01536		0.6516129032

		56.069		0.01536		0.6516129032

		56.079		0.01538		0.6537634409

		56.089		0.01536		0.6516129032

		56.099		0.01521		0.635483871

		56.109		0.01509		0.6225806452

		56.119		0.01489		0.6010752688

		56.129		0.01465		0.5752688172

		56.139		0.01443		0.5516129032

		56.149		0.01421		0.5279569892

		56.159		0.01406		0.511827957

		56.169		0.01384		0.488172043

		56.179		0.01365		0.4677419355

		56.19		0.01345		0.4462365591

		56.199		0.01335		0.435483871

		56.209		0.01326		0.4258064516

		56.219		0.01321		0.4204301075

		56.229		0.01321		0.4204301075

		56.239		0.01326		0.4258064516

		56.249		0.01335		0.435483871

		56.259		0.01343		0.4440860215

		56.269		0.0136		0.4623655914

		56.279		0.01375		0.4784946237

		56.289		0.01392		0.4967741935

		56.3		0.01418		0.5247311828

		56.309		0.01431		0.5387096774

		56.319		0.01448		0.5569892473

		56.329		0.0147		0.5806451613

		56.34		0.01487		0.5989247312

		56.35		0.01501		0.6139784946

		56.36		0.01521		0.635483871

		56.37		0.01528		0.6430107527

		56.38		0.01541		0.6569892473

		56.39		0.01545		0.6612903226

		56.4		0.01548		0.664516129

		56.41		0.01545		0.6612903226

		56.42		0.01536		0.6516129032

		56.43		0.01523		0.6376344086

		56.44		0.01506		0.6193548387

		56.45		0.01487		0.5989247312

		56.46		0.01467		0.5774193548

		56.47		0.01445		0.5537634409

		56.48		0.01426		0.5333333333

		56.49		0.01406		0.511827957

		56.5		0.01384		0.488172043

		56.51		0.01365		0.4677419355

		56.52		0.0135		0.4516129032

		56.53		0.0134		0.4408602151

		56.54		0.01335		0.435483871

		56.55		0.01326		0.4258064516

		56.561		0.01331		0.4311827957

		56.57		0.01331		0.4311827957

		56.58		0.0134		0.4408602151

		56.59		0.01353		0.4548387097

		56.6		0.01365		0.4677419355

		56.611		0.01392		0.4967741935

		56.62		0.01401		0.5064516129

		56.63		0.01418		0.5247311828

		56.64		0.01436		0.5440860215

		56.65		0.01458		0.5677419355

		56.66		0.01477		0.588172043

		56.67		0.01499		0.611827957

		56.68		0.01509		0.6225806452

		56.69		0.01531		0.6462365591

		56.7		0.01536		0.6516129032

		56.71		0.01545		0.6612903226

		56.72		0.01548		0.664516129

		56.73		0.01545		0.6612903226

		56.74		0.01543		0.6591397849

		56.75		0.01531		0.6462365591

		56.761		0.01519		0.6333333333

		56.77		0.01504		0.6172043011

		56.78		0.01484		0.5956989247

		56.79		0.0146		0.5698924731

		56.8		0.0145		0.5591397849

		56.81		0.01421		0.5279569892

		56.82		0.01396		0.5010752688

		56.83		0.01384		0.488172043

		56.84		0.0136		0.4623655914

		56.85		0.01348		0.4494623656

		56.86		0.01333		0.4333333333

		56.87		0.01326		0.4258064516

		56.88		0.01323		0.4225806452

		56.89		0.01323		0.4225806452

		56.9		0.01333		0.4333333333

		56.91		0.01343		0.4440860215

		56.92		0.01355		0.4569892473

		56.93		0.01362		0.464516129

		56.94		0.01382		0.4860215054

		56.95		0.01399		0.5043010753

		56.96		0.01426		0.5333333333

		56.97		0.01438		0.5462365591

		56.98		0.01458		0.5677419355

		56.99		0.01479		0.5903225806

		57		0.01494		0.6064516129

		57.01		0.01511		0.6247311828

		57.02		0.01526		0.6408602151

		57.03		0.01533		0.6483870968

		57.041		0.01548		0.664516129

		57.05		0.01548		0.664516129

		57.06		0.01553		0.6698924731

		57.07		0.01538		0.6537634409

		57.08		0.01536		0.6516129032

		57.09		0.01516		0.6301075269

		57.1		0.01499		0.611827957

		57.11		0.01482		0.5935483871

		57.12		0.01465		0.5752688172

		57.13		0.01443		0.5516129032

		57.14		0.01418		0.5247311828

		57.15		0.01401		0.5064516129

		57.16		0.01382		0.4860215054

		57.17		0.0136		0.4623655914

		57.18		0.01348		0.4494623656

		57.19		0.01333		0.4333333333

		57.2		0.01328		0.4279569892

		57.21		0.01326		0.4258064516

		57.22		0.01328		0.4279569892

		57.23		0.01335		0.435483871

		57.24		0.01353		0.4548387097

		57.25		0.01362		0.464516129

		57.26		0.01372		0.4752688172

		57.271		0.01392		0.4967741935

		57.28		0.01409		0.5150537634

		57.29		0.01426		0.5333333333

		57.3		0.0145		0.5591397849

		57.31		0.01467		0.5774193548

		57.32		0.01482		0.5935483871

		57.33		0.01501		0.6139784946

		57.34		0.01521		0.635483871

		57.35		0.01531		0.6462365591

		57.36		0.01543		0.6591397849

		57.37		0.01548		0.664516129

		57.38		0.01558		0.6752688172

		57.39		0.01558		0.6752688172

		57.4		0.01545		0.6612903226

		57.41		0.01538		0.6537634409

		57.42		0.01523		0.6376344086

		57.43		0.01504		0.6172043011

		57.441		0.01482		0.5935483871

		57.45		0.01462		0.5720430108

		57.46		0.01443		0.5516129032

		57.47		0.01428		0.535483871

		57.48		0.01404		0.5096774194

		57.491		0.01377		0.4806451613

		57.5		0.01362		0.464516129

		57.51		0.01348		0.4494623656

		57.52		0.01338		0.4387096774

		57.53		0.01335		0.435483871

		57.54		0.01333		0.4333333333

		57.55		0.01331		0.4311827957

		57.56		0.0134		0.4408602151

		57.57		0.01348		0.4494623656

		57.58		0.01367		0.4698924731

		57.59		0.01382		0.4860215054

		57.6		0.01396		0.5010752688

		57.61		0.01411		0.5172043011

		57.62		0.01431		0.5387096774

		57.63		0.01453		0.5623655914

		57.64		0.0147		0.5806451613

		57.65		0.01487		0.5989247312

		57.661		0.01506		0.6193548387

		57.67		0.01526		0.6408602151

		57.68		0.01536		0.6516129032

		57.69		0.01543		0.6591397849

		57.7		0.01553		0.6698924731

		57.711		0.01553		0.6698924731

		57.72		0.01558		0.6752688172

		57.73		0.0155		0.6666666667

		57.74		0.01536		0.6516129032

		57.75		0.01521		0.635483871

		57.76		0.01509		0.6225806452

		57.77		0.01482		0.5935483871

		57.78		0.01467		0.5774193548

		57.79		0.01445		0.5537634409

		57.8		0.01418		0.5247311828

		57.81		0.01399		0.5043010753

		57.82		0.01379		0.4827956989

		57.83		0.01362		0.464516129

		57.84		0.01353		0.4548387097

		57.85		0.01345		0.4462365591

		57.86		0.0134		0.4408602151

		57.87		0.01328		0.4279569892

		57.88		0.0134		0.4408602151

		57.89		0.01338		0.4387096774

		57.9		0.01348		0.4494623656

		57.91		0.01365		0.4677419355

		57.92		0.01384		0.488172043

		57.93		0.01399		0.5043010753

		57.94		0.01416		0.5225806452

		57.95		0.01438		0.5462365591

		57.96		0.01455		0.564516129

		57.97		0.01477		0.588172043

		57.98		0.01492		0.6043010753

		57.99		0.01509		0.6225806452

		58		0.01528		0.6430107527

		58.01		0.01538		0.6537634409

		58.02		0.01545		0.6612903226

		58.03		0.01553		0.6698924731

		58.044		0.01553		0.6698924731

		58.053		0.0155		0.6666666667

		58.063		0.01548		0.664516129

		58.073		0.01531		0.6462365591

		58.083		0.01514		0.6279569892

		58.093		0.01499		0.611827957

		58.103		0.01475		0.5860215054

		58.113		0.01453		0.5623655914

		58.123		0.01436		0.5440860215

		58.133		0.01414		0.5204301075

		58.143		0.01392		0.4967741935

		58.153		0.01372		0.4752688172

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617

						-1.0109289617





FO134

		



at 0.1Hz

at 1Hz

at 2Hz

at 3Hz

Time/sec.

Percentage Response




_940232946

_932796188.unknown

