[image: image1.png]Vandagraph
>ensor Technologies






Specification Datasheet - 8010019

SubGravity

SG-22 Oxygen Sensor
v1.1

21st July 2014

	Specifications


	VST part number
	8010019

	OEM
	SubGravity

	OEM model number
	SG-22 Oxygen Sensor

	Output
	9 – 13 mV

	Output connector
	3 pin Molex

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Range
	0 – 100% oxygen

	Pressure range
	600 to 2500 mbar

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds

	Cross interference
	<0.5% vol. O2 response to 5% CO2 balance N2

	Response time
	< 5 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 50 °C, 5% relative error

	Nominal sensor life
	> 500,000 % oxygen hours

	Long term output drift
	< 1% vol oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	> 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes

	Conformal coating on PCB
	Yes

	Pressure equalization - PCB
	Yes 

	Pressure equalization - output cap
	None

	Weight
	25 g

	Flow diverter
	No

	Packaging
	Gas barrier bag

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.



Page 2 of 2

	Drawings/images


[image: image2.jpg]



[image: image3.jpg]@145
|
T
~ i
|
|
| w0
| 3 -
T
|
\
X
! M16 x 1 (5/8-24 UNEF)
@28
1 _—Locking Laich
(-)f%/ TN
h==a
$29

%\N MOLEX CONNECTOR



[image: image4.jpg]



Page 3 of 4

	Labelling


Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics – black

[image: image5.png]Draper, UT 84020, USA
www.sub-gravity.com x
SUBDRAVITY info@sub-gravity.com

SubGravity
9 12356 S. 900 E. Suite 103 C E SG-22 Oxygen Sensor

M SIN SG123456
2014-07  INNTIMINALINIH




Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

[image: image6.png]SubGravity

12356 S. 900 E. Suite 103
Draper, UT 84020, USA
www.sub-gravity.com

+1 801.256.6608
info@sub-gravity.com

SUBORAVITY

SG-22 Oxygen Sensor

(] 2014-07 [1i] C€
SIN SG123456 —
JRUREE TR A

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor
may require 30 minutes to reach signal stability after installation.





Page 4 of 4






Vandagraph Sensor Technologies Ltd.

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

www.vandagraphst.com

