	[image: image3.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid





	
	[image: image2.png]


 Ltd.
15 Station Road, Cross Hills, Keighley, West Yorkshire, BD20 7DT, UK.

Website: www.viamed.co.uk.  Email: info@viamed.co.uk.
Tel: +44 (0)1535 634542.  Fax: +44 (0)1535 635582.


.
[image: image1.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid




 Tom Thumb Resuscitation Unit Servicing Manual (TT490).
Version 1.1.  16/07/04



Part no: 0390023.

TOM THUMB CALIBRATION / SERVICE SHEET.
	Description.
	Tom Thumb.
	Serial No.
	

	Model.
	TT 490-15.
	Time & Date of Service / Calibration.
	


Record the manometer reading in millibars below, ensuring the TT490-15 meets the limits specified.

	Test Equipment
	Test
	Specification
	Reading
	P/F

	
	
	
	Calibration
	

	
	
	
	
	

	Snoop liquid.
	Check all ports / connections for leaks.
	No bubbling.
	
	p

	
	
	
	
	

	Manometer
	Adjustable Valve:

@ 15 Lpm
	Minimum:

( 8 cmH2O.
	
	p

	
	
	
	
	

	Manometer
	Adjustable Valve:

@ 15 Lpm.
	Tolerance:

( 43 & ( 47 cmH2O.
	
	p

	
	
	
	
	

	Manometer
	Precision Valve:

@ 15 Lpm.
	Maximum:

+3.0 cmH2O over the adjustable valve setting.
	
	p

	
	
	
	
	

	Manometer
	Pressure Gauge Test:

@ 30 cmH2O.  
	Tolerance:                             (29 & (31 cm H(O
	
	p

	
	
	
	
	

	Visual check.
	Gauge cannot be removed without tools.
	
	
	p

	
	
	
	
	

	Visual check.


	Gauge appears straight.
	
	
	p

	
	
	
	
	

	Visual check.
	All adjustable settings are set to a minimum.
	
	
	p

	
	
	
	
	

	Visual check.
	Labels are attached.
	CE Label.
	
	p

	
	
	Viamed Flowmeter Label.
	
	

	
	
	Serial No. Label.
	
	

	
	
	Tom Thumb Label
	
	


	Service.
	Parts Replaced.

	Changed “O” Rings
	(
	Diamond Copper Washer
	(

	Re-greased Adjustable Valve Internal Parts
	(
	Adjustable Valve
	

	Re-calibrated Tom Thumb
	(
	Adjustable Valve Internal Parts
	

	
	
	Safety Valve
	

	
	
	Safety Valve Internal Parts
	

	
	
	Pressure gauge
	

	
	
	Pressure Gauge Face
	

	
	
	“O” Rings
	(

	
	
	Flow Gauge
	

	
	
	Flow Gauge Cover
	

	
	
	Outlet
	

	
	
	Inlet 
	

	
	
	Blanking Bolts
	

	
	
	Rail Clamp
	

	
	
	Main Body Block
	


Calibration: …………………………………
Signed:  …………………………………           















[image: image2.png][image: image3.jpg]