[image: image24.png]


 

                                             Tom Thumb.

	COMPANY OPERATING PROCEDURE.


0310030 TT480.
VM3/COP/50.02
	Date:
  13-Jul-04.
	Revision date: 5-Apr-07.
	Issue:   1.



Important: Use only oxygen-compatable grease during assembly of Tom Thumbs.  Do not use or allow organic greases to enter the Tom Thumb.  Ensure all parts are clean before assembly.

	Equipment required.
	12mm open-ended spanner, pick, isopropyl alcohol, kitchen tissue,,adjustable spanner, PTFE tape, 3mm Allen key, 7/8 A/F spanner, 5/8 – 9/16 A/F spanner, O-Rings, diamond copper washer.


	Parts list.

	Qty.
	Description.
	Part No.

	1
	Body block
	0330201

	1
	Diamond copper washer
	0330212

	1
	Pressure gauge
	0330203

	6
	Thin O ring – 1/16 section
	0330214

	1
	Flowmeter blocking bolt & nut
	0330206

	1
	Precision valve
	0330210

	1
	Blanking bolt
	0330207

	1
	Adjustable valve
	0330211

	1
	10mm inlet
	0330208

	1
	15mm outlet
	0330209  

	1
	Serial no. Label
	0390030

	1
	Spacer block
	0330215

	1
	Rail clamp
	0330055

	4
	M4x10mm hex drive bolts
	0330216

	1
	“Tom Thumb” label
	0390015

	1
	Adjustable valve label
	0390010

	1
	Address Label
	0390016

	1
	Gauge Face Label
	0390017

	---
	Fomblin grease
	0330220


	Assembly.


1. [image: image24.png]a). Clean the surface of the body block with isopropyl alcohol. Clean the threads of the body block to remove residual swarf.  b). Affix Tom Thumb label onto the front face of the block.

	[image: image1.jpg]



	[image: image2.jpg]





2. a). Insert diamond copper washer into the single threaded hole on the label side of body block.  b). Screw in the pressure gauge in the single threaded hole on the label face of the body block; the gauge should appear straight in comparison to the block and not able to be removed without the aid of tools. Use 9/16 AF open-ended spanner to tighten.

	[image: image3.jpg]



	[image: image4.jpg]





3. Place 2x thin O rings onto the “Flowmeter” blocking bolt.  Push flowmeter blocking bolt / O rings from bottom to top into the left hand (non threaded) hole and screw the flowmeter blocking bolt nut into position. Use 2x 5/8 AF spanner to tighten.

	[image: image5.jpg]



	[image: image6.jpg]



	[image: image7.jpg]





4. Place a thin O ring over the threaded end of the adjustable valve. Screw the adjustable valve into the remaining threaded hole in the top face of the body block. Use the 7/8 AF spanner to tighten.

	[image: image8.jpg]



	[image: image9.jpg]





5. Place a thin O ring over the thread, of the precision valve and screw into the centre bottom threaded hole. Use 7/8 AF spanner to tighten.

	[image: image10.jpg]



	[image: image11.jpg]





6. Place a thin O ring onto the 10mm inlet and screw into the threaded hole on the left hand face using a 10mm open-ended spanner until tight.

	[image: image12.jpg]



	[image: image13.jpg]





7. Place a thin O ring over the thread of a blanking bolt and screw into the remaining threaded hole on the bottom face of the body block. Use 9/16 AF open-ended spanner to tighten.

	[image: image14.jpg]



	[image: image15.jpg]





8. Place a thin O ring over the threaded end of the 15mm outlet.  Screw the 15mm outlet into the threaded hole in the right face of the body block. Use 5/8 AF spanner to tighten.

	[image: image16.jpg]



	[image: image17.jpg]oG

Se¥







9. Affix Viamed details / serial number label across the Left Hand rear face of the body block.


[image: image18.jpg]



10. Pierce two holes through serial no. label and attach the spacer block to the body block using 2x M4x10mm hex drive bolts.

[image: image19.jpg]



11. Attach rail clamp to the spacer block using 2x M4x10mm hex drive bolts.


[image: image20.jpg]



	Testing & Calibration: TT480.


1. Check the TT480 for damage i.e. scratches to plastics, poor chroming, etc.

2. Set the adjustable valve to minimum (fully counter clockwise).

3. Connect the inlet to the oxygen supply at a pressure of 4 bar.

4. Connect the digital manometer (on 0-100.0 mbar range) with T piece adapter to the Tom Thumb outlet.

5. Set the flow of oxygen to 10 Lpm

6. Check all mechanical connections between parts for leaks; use Snoop fluid and check for bubbling.

7. Check for minimum pressure out: Cover the T piece adapter port. Check the TT pressure gauge is less than 8.0 cmH2O. Check that the digital manometer reading is less than 7.8 mbar and record on the Tom Thumb Calibration / Test & QA Sheet.

8. Set up the adjustable valve: Set the adjustable valve control to maximum. Cover the holes in the precision valve and the T piece adapter port. Alter the adjustable valve collar until a reading of 43 - 47 cmH2O is achieved on the TT pressure gauge. Ensure a reading of 42.2 – 46.1 mbar on the digital manometer and record on the Tom Thumb Calibration / Test & QA Sheet. Slightly unscrew the adjustable valve and apply Oxygen-Compatible threadlock to the exposed threads. Retighten the adjustable valve with the adjustable spanner.

[image: image21.jpg]



9. Check for TT pressure gauge sticking: Release and cover the T piece adapter port a number of times, ensuring that TT pressure gauge returns to 43 - 47 cmH2O. Uncover the precision valve holes.

10. Set up the precision valve: Slightly unscrew the precision valve adjustable screw and apply Oxygen-Compatible threadlock to the exposed threads.  Screw into the precision valve body until flush. Cover the holes in the adjustable valve and the T piece adapter port.  Carefully adjust the precision valve screw in quarter turns to achieve a reading on the TT pressure gauge of 3.0 cmH2O greater than the reading recorded at (8) (clockwise adjustment on precision valve screw to increase). Record the digital manometer reading on the Tom Thumb Calibration / Test & QA Sheet.

[image: image22.jpg]



11. TT pressure gauge accuracy check: Cover the T piece adapter port and set flow to achieve a displayed reading of 30 cmH2O on the TT pressure gauge. Ensure a reading of 29-31 mbar on the digital manometer and record
12. Turn the flowmeter and adjustable valve to minimum. Disconnect from oxygen supply and disconnect the digital manometer.

13. Clean the Tom Thumb with isopropyl alcohol and apply an adjustable valve label to the body block.

[image: image23.jpg]



14. Attach the TT480 Instructions for Use card.
15. Ensure stock sheets are generated and submitted together with the Tom Thumb Calibration / Test & QA Sheet for QA checking. 

	Quality Checks: TT480.


Do not carry out QA checks on Tom Thumbs until at least 1 hour after testing – this allows Oxygen-Compatible threadlock, applied during testing, to be totally dry.

1. Check the TT480 for damage i.e. scratches to plastics, poor chroming, etc.

2. Set the adjustable valve to minimum (fully counter clockwise).

3. Connect the inlet to the oxygen supply at a pressure of 4 bar. 

4. Connect the digital manometer (on 0-100.0 mbar range) with T piece adapter to the Tom Thumb outlet.

5. Set the flow of oxygen to 10 Lpm.

6. Check all mechanical connections between parts for leaks; use Snoop fluid and check for bubbling.

7. Check for minimum pressure out: Cover the T piece adapter port. Check the TT pressure gauge is less than 8.0 cmH2O. Check that the digital manometer reading is less than 7.8mbar and record on the Tom Thumb Calibration / Test & QA Sheet. 

8. Set up the adjustable valve: Set the adjustable valve control to maximum. Cover the holes in the precision valve and the T piece adapter port. Ensure a reading of 42.2 – 46.1 mbar on the digital manometer and record on the Tom Thumb Calibration / Test & QA Sheet. 

9. Check for TT pressure gauge sticking: Release and cover the T piece adapter port a number of times, ensuring that TT pressure gauge returns to 43 - 47 cmH2O. Uncover the precision valve holes.

10. Set up the precision valve: Cover the holes in the adjustable valve and the T piece adapter port.  Ensure a reading on the TT pressure gauge of 3.0 cmH2O greater than the reading recorded at (8). Record the digital manometer reading on the Tom Thumb Calibration / Test & QA Sheet.

11. TT pressure gauge accuracy check: Cover the T piece adapter port and increase the flow of gas to achieve a displayed reading of 30 cmH2O on the TT pressure gauge. Ensure a reading of  29 – 31 mbar on the digital manometer and record
12. Turn the flowmeter and adjustable valve to minimum. Disconnect from oxygen supply and disconnect the digital manometer.

13. Clean the Tom Thumb with isopropyl alcohol.

14. Ensure the appropriate labels have been fitted and the Instructions for Use card is attached.  Complete stock sheets and file.  Ensure the unit is booked onto the computerised stock system.

	Packaging.


Wrap thoroughly in bubble packaging. 

Add One Patient Circuit (if required)

Pack appropriately for despatch by insured post, courier or company delivery.

Pt No. 0330201





Pt No. 0330206





Pt No. 0330210





Pt No. 0330211





Pt No. 0330208





Pt No. 0330207





Pt No. 0330209





Pt No. 0330215





Pt No. 0330055








Page 1 of 7

