VM3COP61.00 Microstim 2510000 QA Procedures

Microstims DB's are supplied to Viamed fully assembled and Tested,

· Ensure there is no physical damage to the unit or marks to labels.

· Remove sticker from the cable

· Ensure that the battery is in good condition. Without the patient lead connected press one of the function buttons, (in this case P.T.C is better). If the battery condition is good then the LED indicator will be green. The LED indicator changes from green to amber, then red as the battery output decreases. Replace the battery if the LED indicator is red.

· Connect the DB3 tester to the output, red (+ve) and black (-ve).  Set the intensity control to maximum (fully clockwise).

· Depress each of the function buttons in turn.  When a function button is pressed and held there should be audio output in conjunction with illumination of the green LED indicator on the Microstim DB3. The output can be verified by observing if the red LED indicator on the DB3 is also illuminated in synchronization. 


Test result:

a) DB3 Tester indicator illuminated red – Microstim output test passed.

b) DB3 Tester indicator not illuminated:




Battery level maybe low, check battery condition 

If still not illuminated check intensity control set to maximum.
If still not illuminated then the Microstim has developed a fault.
· Disconnect the Microstim, Remove the battery and refit the battery cover.

VM3COP99 Microstim 2510000 QA Procedures _24_11_14.odt
24/11/2014
Page 1 of 1

