	VANDAGRAPH
  COMPANY OPERATING PROCEDURE

	Calibration Procedures  O2 Analyser. VN202/Tek-Ox

	VM3/COP/40.58

	Date: 25/11/14
	Revision Date: 27-Nov-14
	Issue 1



	[image: image1.jpg]






Table 1: 

1. Connect analyser to gas lines

2. Open oxygen cylinder

3. When stable set 100% ( at least 1 minute)

4. Switch  gas to air

5. Allow to stabilise ( at least 1 minute)

6. Remove from gas.

7. Remove sensor and connect to test rig Microcal & DVM

8. Set microcal ro approx. 10mV

9. Set cal on analyse  until analyser reads  20.9%

10. Set 20,30,40,50,60 70 80 90 and note readings

11. Fill in and sign calibration certificate.

12. Transfer results to electronic calibration certificate

13. Print out and sign
	
	


J.S.Lamb 

27/11/14
VM3COP40.58 Calibration VN202 Tek-OX.doc   27/11/2014      Page 1 of 1

