OBL Sensor Specification

Type: OOD103

Änderungsregister / Change log

| Revision | Date | Author | Description |
| --- | --- | --- | --- |
| 0 | 04/2016 | Maik Lingies | Create |
| 1 | 07/2019 | Maik Lingies | Change |
| 2 | 08.12.2020 | Maik Lingies | Start with Revisions Log Add Label: OOD103: E1002856 (REF 8010060 / OSGA-20) remove Label: OOD103: 1001570 (REF 8010015 / UDT-22) OOD103: E1002736 (REF 8010018 / M3S), OOD103: E1001570 (REF 8010019 / SG-22), Change old article number to new SAP number |
| 3 | 17.01.2024 | Maik Lingies | Add new Label OOD103: E1002931 (REF 8010029 / mCCR) |

| Measurement Range: | 0 % ... 100 % oxygen |
| --- | --- |
| Output in Ambient Air: | 9 to 13 mV |
| Electrical Interface: | 3 pin (Molex 22-11-1031) |
| Repeatability: | < 1% vol. O2 @ constant temperature and pressure |
| Linearity error: | 0.1 to 1.0 bar ppO2 < 3% relative |
|  | 0.1 to 2.5 bar ppO2 < 5% relative (at ppO2 > 1.0 bar after 1 h equilibration time) |
| Response time: | < 5 sec. to 90% of final value |
| Zero Offset Voltage: | < 0.5% oxygen reading in 100% nitrogen @ 25°C after 36 seconds |
| Cross Interference: | < 0.5% vol. O2 response to: 5% CO2 balance N2 |
| Influence of Humidity: | - 0.03% rel. per % RH at 25°C |
| Pressure Range: | 600 to 2500 mbar |
| Influence of Mechanical Shock: | < 1% relative after a fall from 1m |
| Operating Temperature: | 0°C to 50°C |
| Temperature Compensation: | Built-in NTC compensation |
| Effect of Temperature Compensation (steady state): | Between 0°C and +50°C: 5% relative error |
| Operating Humidity: | 0 - 99% RH non-condensing |
| Long Term Output Drift: | < 1% vol. oxygen per month |
|  | Typically < -15% relative over lifetime |
| Storage Temperature: | -20°C to +50°C |
| Recommended Storage: | +5°C to +15°C |
| Recommended Load: | ≥10 kOhms |
| Warm-Up Time: | < 30 minutes, after replacement of sensor |
| Nominal Sensor Lifetime: | 500,000% vol oxygen hours |
| Weight: | Approximately 28 grams |
| Warranty Period: | 15 months (including 3-month shelf life) |
| Part No.: | OOD103: E1001570 (REF 8010000 / VST2200) |
|  | OOD103: E1002733 (REF 8010006 / CR22D) |
|  | OOD103: E1002734 (REF 8010008 / NaNM01) |
|  | OOD103: E1002735 (REF 8010014 / K-22D) |
|  | OOD103: E1002736 (REF 8010018 / M3S) |
|  | OOD103: E1002737 (REF 8010020 / 22D) |
|  | OOD103: E1002738 (REF 8010021 / DST-2200) |
|  | OOD103: E1002749 (REF 8010026 / MSD100) |
|  | OOD103: E1002856 (REF 8010060 / OSGA-20) |
|  | OOD103: E1002931 (REF 8010029 / mCCR) |
| All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air. |  |

Mechanical draft:

General tolerances ISO 2768-c

*Intermediate thread: Metric / Unified Extra Fine

Customized Design Specification:

- Integration of additional Zitex membrane in sensor gas diffusion way

- Conformal coating on the sensor PCB with PLASTIK 70

- 2 x 2mm holes drilled in the sensor PCB

- No information about EnviteC on PCB top / bottom sides, information for ordering process and traceability only.

PCB Layout:

| Top | Bottom |
| --- | --- |

Product Labeling

| OOD103: E1001570 (REF 8010000 / VST2200) |  |
| --- | --- |
| Label on Sensor (90x15) | Label on Blister- Card (63,5x46,6) |
| OOD103: E1002733 (REF 8010006 / CR22D) |  |
| Label on Sensor (90x15) | Label on Blister- Card (63,5x46,6) |
| OOD103: E1002734 (REF 8010008 / NaNM01) |  |
| Label on Sensor (90x15) | Label on Blister- Card (63,5x46,6) |
| OOD103: E1002735 (REF 8010014 / K-22D) |  |
| Label on Sensor (90x15) | Label on Blister- Card (63,5x46,6) |
| OOD103: E1002736 (REF 8010018 / M3S) |  |
| Label on Sensor (90x15) | Label on Blister- Card (63,5x46,6) |
| OOD103: E1002737 (REF 8010020 / 22D) |  |
| Label on Sensor (90x15) | Label on Blister- Card (63,5x46,6) |
| OOD103: E1002738 (REF 8010021 / DST-2200) |  |
| Label on Sensor (90x15) | Label on Blister- Card (63,5x46,6) |
| OOD103: E1002749 (REF 8010026 / MSD100) |  |
| Label on Sensor (90x15) | Label on Blister- Card (63,5x46,6) |
| OOD103: E1002856 (REF 8010060 / OSGA-20) |  |
| Label on Sensor (90x15) | Label on Blister- Card (63,5x46,6) |
| OOD103: E1002856 (REF 8010060 / OSGA-20) |  |
| Label on Sensor (90x15) | Label on Blister- Card (63,5x46,6) |

Customer Approval

Date:

Signature: