[image: image1.png]


 

                   O2 Sensors
	COMPANY OPERATING PROCEDURES


Teledyne R22-2BUD (0110222), R22A (0110122), R22AH (0110142), 
R22DHO (0110213), R22DN (0110216), R22MED (0110022)
O2 Sensor Test Procedure
VM3/COP/37.37
	Date:
  3-Feb-03
	Revision date: 30-Apr-15
	Issue:   3



Test & Equipment Required
Functional test required using R22 tester and R22 test lead, (digital multimeter and 3 pin molex adapter). 200mV range

Output specification as in document: (VM3COP37.99, column (i)).

Method.
1. Open the packaging and remove the O2 sensor


2. Check the O2 sensor damage and signs of leaking electrolyte.

3. Check for an output  Ensure that  reading lies within the output specification range,
10mV +/- 3mV

4. Check for an output see VM3COP37.99 or Intrastats QA form for Values  

5. If no, or very low output, or sensor is returned as faulty leave the O2 sensor exposed to air for 1 Hour then re-test. 

6. If the O2 sensor fails scan them to return to supplier shelf (182560) they are sealed, initialed and output/fault written on bar code . 

7. Disconnect the O2 sensor from the R22 tester and label with Viamed stickers if appropriate. 

8.   Scan the Barcode into the QA form in intrastats, and complete the form.

9.   Reseal in serial numbered packet, 

VM3COP37.37 R-22 Oxygen Sensor 0110122_30_04_15.doc         30/04/2015      Page 1 of 1

[image: image1.png]