Parts required for annual service of Tom Thumb variants TT480 and TT490 v1.0

Steve Hardaker – 22nd May 2015


Parts required for Annual Service of TT480 and TT490 Tom Thumb

For the annual service of a TT480 or TT490 Tom Thumb infant resuscitator, the following parts are required: 


	Qty
	Part No
	Description

	1
	0330213
	O-ring, thick

	6
	0330214
	O-ring, thin

	The following may be required if the gauge is not sufficiently tight:

	1
	0330212
	Gauge seal (diamond copper washer)


Notes: 

Please see separate price list for current pricing. 

For any other variants of Tom Thumb, please contact Viamed. 

The gauge seal is only required if the gauge is loose or has been removed.

This seal is a diamond-profile copper washer that it designed to compress as the gauge is tightened, allowing the gauge to be rotated up to 1 full turn from initial contact with the washer so as to allow correct alignment. 

Do not re-use a gauge seal that has been previously compressed.

For variants of Tom Thumb that have an oxygen or air hose, Viamed recommends that these be replaced every 4 years.

Hoses are available in different lengths; please contact Viamed with the required gas type and hose length for a part number and price.

