Corrosion in Analox EII analysers

[image: image1.jpg]


[image: image2.jpg]


As the removal of the back plate allows the electronics to be exposed. Several cases have been seen where the battery wires are corroded (black) The corrosion travels up the wires and ounto the pcb. This could be cased by sea air.
