The display is a four digit read out.
The range is 0-100%
The resolution is +/-0.1%
Accuracy for the system is quoted as +/- 1.0%
Initial limitations 0.1 error at 20 equates to a 0.5 error at 100
The potentiometer has been chosen for reliability but is extremely sensitive when setting to 0.1.
This is not a problem on an analyser with a +/-1% specification.
The device is totally software controlled so the electronics cannot be inaccurate
If it works it will be within specification unless the sensor is inaccurate. Therefore a quick electronic 20% 100% check is adequate.
Electronic Calibration equipment

The VN202 mkII is designed to work in conjunction with an Oxygen sensor.
Dry Oxygen = 100%
Dry Air = 20.9%
Ratio is 4.7846
The Microcal and the Digital Voltmeter are calibrated and traceable to Namas.
The Microcal is set to 10mV as near as possible . The VN202 calibration is adjusted to 20.9% as near as possible
The Micocal is set to 47.846 and the VN202 should read 100.0
There will always be an error in setting the Microcal and the VN202 accurately below 1%

Final test of VN202 mkII before despatch

It is usually sufficient to add an Oxygen sensor and to check it can be set to 20.9%.

Other tests are superfluous.
However a calibration certificate can be supplied on individual units on request before despatch
In this instance the unit can be electronically tested with an Oxycal VM3COP40.17 ( a simplfied calibrated voltage generator) to 20.9% & 100%.
The complete instrument and sensor combination can also be calibrated in Dry Air & 100% Oxygen.

J.S.Lamb
03/10/12