# Certificate of Calibration

## Certificate number: 14291878691

## V1000 Foetal Heart Simulator
Part Number: 1410000
Serial number: PR01702A12

This certificate reports recorded values for the instrument specified above.

The output rates were measured in accordance with procedure: VM3COP55.01, using test equipment CE0185 - calibration traceable to UKAS calibration standards.

V1000 / BPM setting	Period values stated in milliseconds				Result / PASS/FAIL
	Target	Lower limit	Upper limit	Actual
60	1000.00	990.00	1010.00	993.62	PASS
90	666.70	660.03	673.37	665.96	PASS
120	500.00	495.00	505.00	500.29	PASS
150	400.00	396.00	404.00	400.33	PASS
180	333.30	329.97	336.63	332.79	PASS
210	285.70	282.84	288.25	285.97	PASS

Signed:	______________________

Name:	Phil Crossley

Date:	16th April 2015