Certificate of Calibration

Certificate number: 1425984817

V1000 Foetal Heart Simulator

Part Number: 1410000

Serial number: PR01883A18

This certificate reports recorded values for the instrument specified above.

The output rates were measured in accordance with procedure: VM3COP55.01, using test equipment CE0185 - calibration traceable to UKAS calibration standards.

	V1000 

BPM setting


	Period values stated in milliseconds
	Result

PASS/FAIL

	
	Target
	Lower limit
	Upper limit
	Actual
	

	60
	1000.00
	990.00
	1010.00
	992.87
	PASS

	90
	666.70
	660.03
	673.37
	665.47
	PASS

	120
	500.00
	495.00
	505.00
	499.91
	PASS

	150
	400.00
	396.00
	404.00
	400.03
	PASS

	180
	333.30
	329.97
	336.63
	332.54
	PASS

	210
	285.70
	282.84
	288.25
	285.75
	PASS


Signed:
 ______________________                                                                                         

Name:
     S Nixon

Date:
19th March 

