	CERTIFICATE  OF CALIBRATION
Issued by:                   Vandagraph  Ltd.
Date issued:               12 June 2015
Certificate No.:            1434116294 
Signed:                       ____________



Client                       Steve Sanders

Date of calibration. 
12 June 2015
Instrument 
Mixcheck


Serial number: 

284767

Procedure I.D.  

VM3COP40.56



Performance Status  
PASS
The equipment detailed herein has been calibrated  and now conforms to the manufacturer’s published specifications. The Q.A. of the above is in strict conformity to BS EN ISO 9001-2000 & ISO 13485-2003.  All measurements made are traceable to NAMAS.
	Equipment used to carry out calibration

	Fluke 77 DVM. CE055
	

	
	

	Microcal  Serial Number CE076
	

	Oxygen BOC Medical 99.5% License 07351
	

	Helium  BOC Medical 99.5% License 07351
	


	Setting
	Oxygen Reading
	Helium Reading

	100% Oxygen
	100%
	*1.8%

	Dry Air
	21.0%
	0.2%

	100% Helium
	0.2%
	100%

	Trimix
	32.3%
	37.4%


* 100% Oxygen will produce a small error of approximately 1.4% (helium) due to the difference in the Thermal conductivity of Air and Oxygen.
Vandagraph Cal Cert Mixcheck 267472.doc                            9th March 2010

