CERTIFICATE OF CALIBRATION / Issued by: Vandagraph Ltd. / Date issued: 13 August 2015 / Certificate No.: 1439473289 / Signed: ____________

Client Pure Diving
Date of calibration.	13 August 2015
Instrument	Mixcheck
Serial number:	270454
Procedure I.D.	VM3COP40.56
Performance Status	PASS

The equipment detailed herein has been calibrated and now conforms to the manufacturer’s published specifications. The Q.A. of the above is in strict conformity to BS EN ISO 9001-2000 & ISO 13485-2003. All measurements made are traceable to NAMAS.

Equipment used to carry out calibration
Fluke 77 DVM. CE055

Microcal Serial Number CE076
Oxygen BOC Medical 99.5% License 07351
Helium BOC Medical 99.5% License 07351

Setting	Oxygen Reading	Helium Reading
100% Oxygen	100%	*1.3%
Dry Air	21.3%	0
100% Helium	0.2%	99.9%
Trimix	32.5%	37.6%

* 100% Oxygen will produce a small error of approximately 1.4% (helium) due to the difference in the Thermal conductivity of Air and Oxygen.